## ReadMe
PERMEATION DEVICE WEIGHT LOSS WORKSHEET
What this Workbook is all About
This workbook enables the user to estimate the rate at which the weight of a permeation tube
decreases. A linear relationship between the tube's weight and elapsed time is established. If the
estimated weight at time zero is significantly different from the actual weight at time zero, then
then at least the earliest data pair should be removed from the analysis. Uncertainty of the slope
estimate (the rate of weight loss or drift) will be determined. The traceability protocol requires that
this estimate have a relative uncertainty of less than 1%.
How the Workbook is Organized
The workbook consists of several worksheets, which are displayed as tabs on the bottom of
the screen. Each worksheet performs a distinct function as described below.
ReadMe - describes the workbook and explains how to use the worksheets
Data - allows for user input of the calibration data (elapsed time and weight)
Regression - calculates the drift and its uncertainty
Calibration Curve - graphically displays the drift line together with its confidence bands
Residual - graphically displays the vertical difference between the observed and estimated weights.
Report - summarizes the assay results for a permeation device
Chart Data - includes the data used to create the curve and residual charts.
How the Worksheets are Setup
Each worksheet contains instructions that guide the user through the steps in using
the worksheet. The worksheets are also color coded to simplify use.
Shaded cells that are bordered in blue lines are cells whose contents you can change
(i.e., enter data). In other sheets you can change the following variables:
Sheet |  | Variable | Location | Current Value
Data |  | Unit of Time | H22 | m
Data |  | Unit of Weight | H24 | g
Report |  | Device ID | F5 | testdata
Chart Data |  | significance level | D2 | 1e-05
Derived values are colored red. These cells contain formulas that should not be changed. The
cells are protected to prevent alteration.
How to Use the Worksheets
Step 1: | Enter the elapsed times (all in the same units) and corresponding tube weights
in the Data worksheet. The worksheet will compute the total weight loss for
each observation.
Step 2: | Follow the steps of the Regression worksheet. Those steps involve reviewing the
regression line's coefficients. The intercept should be consistent with the initial weight
and the slope's uncertainty should be less than 1%.
Step 2: | Examine the Calibration Curve chart. You may need to adjust the chart's axis scaling.
The points should appear to fall virtually on top of the black line. The black line should be
very close to its confidence bands (colored red and blue). To adjust the "closeness" of
the confidence bands, change the significance level of cell D2 in the Chart Data
worksheet.
Step 3: | Examine the Residual worksheet. The residuals should appear to be random in both
magnitude and direction. If they appear to follow a regular pattern, then the simple linear
model is not appropriate. The device does not have a constant rate of weight loss. More
time may be required to establish and measure a linear relationship. Observations taken
before the linear relationship is established should be discarded and not used in the
statistical analysis.
Step 4: | Print the one-page report provided in the Report sheet. The report summarizes the assay
data and indicates the uncertainty of the estimate.

## Data
DATA ENTRY WORKSHEET
Elapsed
| Time | Weight
i | Xi | Yi | Enter the data in the blue bordered spaces. The first
0 | 0 | 4.3542 | column (X) is for the elapsed time. The time of the first
1 | 8641 | 4.3374 | entry should be zero. The second column (Y) is for the tube
2 | 18722 | 4.3168 | weights.
3 | 40322 | 4.2735
4 | 64802 | 4.2245 | n = number of weighings. This can't exceed 50.
5 | 74882 | 4.2038 | n = | 7
6 | 84962 | 4.1844 | No data entry is required for derived values, which are
colored red, such as n and i. These values are tabulated
automatically and their cells are protected from alteration.
Multiple weighings at a single point in time requires multiple
entries in each column. Reenter the time in column X and
enter the corresponding weight in column Y.
Enter the time and weight units in the spaces below:
|  |  |  | Unit of Time = | m
|  |  |  | Unit of Weight = | g
| 84962 | = Max (Xi)
| 4.1844 | = Min (Yi)

## Regression
WORKSHEET MODEL: LINEAR WITH NON-ZERO INTERCEPT
This sheet derives the regression equation in the form: y = b0 + b1 x + e. The intercept
and slope are estimated. The sheet determines whether the intercept (weight estimated for
time zero) is significantly different from the observed weight at time zero. It also estimates the
uncertainty in the slope estimate and compares this uncertainty with EPA's 1% limit. Reference:
Appendix A, EPA Traceability Protocol Assay and Certification of Gaseous Calibration
Standards, USEPA, 1997.
STEP 1
Review the estimates of the intercept (bo), slope of the regression line (b1), and their
confidence limits along with the estimates of variance-covariance matrix (V) and the
residual error variance (Var).
Derivation of the estimated intercept (b0) and slope (b1) of the regression line
X'X = | 7 | 292331 |  | Y'Y = | 127.6972 | 95% Confidence Limits
| 292331 | 19076198421 |  | df = | 5 | b0 lower limit = | 4.3539
(X'X)-1 = | 0.39679 | -6.0806e-06 |  | t(0.95, df) = | 2.5706 | b0 upper limit = | 4.3549
| -6.0806e-06 | 1.456e-10
det(X'X) = | 48075975386 |  |  | b0 = | 4.3544 | b1 lower limit = | -2.0155e-06
X'Y = | 29.8946 |  |  | b1 = | -2.0053e-06 | b1 upper limit = | -1.9952e-06
1234672.7769
Derivation of the error variance (Var) and variance-covariance matrix (V)
|  | b'X'Y = | 127.6972 |  | SS(model), 2df
|  | b'X'Y - sum(Y)2 / n = | 0.027619 |  | SS(regression) 1df
|  | (Y'Y - b'X'Y) = | 5.3819e-07 |  | SS(residual)
|  | Var = (Y'Y - b'X'Y) / df = | 1.0764e-07 |  | MS(residual), n-2 df
|  | V = Var * (X'X)-1 = | 4.271e-08 | -6.5451e-13
|  |  | -6.5451e-13 | 1.5672e-17
STEP 2
Examine the upper and lower limits for the intercept, b0 and compare with the initial weight.
| b0 lower limit = | 4.3539
| b0 upper limit = | 4.3549
| y0 = | 4.3542 | Initial weight is within the confidence limits for the intercept
If y0 is within the confidence limits for the intercept, proceed to STEP 3. Otherwise, consider
removing the first observed weight from the analysis. Re-enter the times and weights.
Remember that the first time (X0) should be zero. This will require adjustment of the other
elapsed times. After entering the data, return to STEP 1, above.
STEP 3
Examine the upper and lower limits for the slope, b1. The limits should differ from
the estimate by no more than +/- 1% of the estimated slope.
|  | (b1 upper - b1) / ¦b1¦ = | 0.5075% | Conclusion: Uncertainty is acceptable.
|  | (b1 lower - b1) / ¦b1¦ = | -0.5075%
If the uncertainty is unacceptable, consider collecting additional data. Also, view the Curve and
Residual plots. They may reveal a nonlinear relationship for a portion of the data. The initial
measurements may not align with subsequent measurements if the device was in the process of
stabilizing or equilibrating during those times. If this is the case, the initial points of the Residual
chart would appear to be outliers. The residuals of will the same sign (all positive or all
negative) and their magnitude will likely be greater than the magnitude of subsequent residuals.
If this is the case, consider removing the initial points from the computations and re-enter the
remaining times and weights with the times adjusted so the first entry has time zero.
If the uncertainty is acceptable, print the Report spreadsheet and include it with the
certification documentation.

## Report
ASSAY RESULTS FOR PERMEATION DEVICE
This sheet provides calibration information and assay results, including uncertainty estimates
for a permeation device identified as: |  |  |  | testdata
Reference: Appendix B, EPA Traceability Protocol Assay and Certification of Gaseous
Calibration Standards, USEPA, 1997.
Test Results
Intercept (b0), slope (b1), and their confidence limits
X'X = | 7 | 292331 |  | Y'Y = | 127.6972 | 95% Confidence Limits
| 292331 | 19076198421 |  | df = | 5 | b0 lower limit = | 4.3539
(X'X)-1 = | 0.39679 | -6.0806e-06 |  | t(0.95, df) = | 2.5706 | b0 upper limit = | 4.3549
| -6.0806e-06 | 1.456e-10
det(X'X) = | 48075975386 |  |  | b0 = | 4.3544 | b1 lower limit = | -2.0155e-06
X'Y = | 29.8946 |  |  | b1 = | -2.0053e-06 | b1 upper limit = | -1.9952e-06
1234672.7769
Error variance (Var) and variance-covariance matrix (V).
|  | b'X'Y = | 127.6972 |  | SS(model), 2df
|  | b'X'Y - sum(Y)2 / n = | 0.027619 |  | SS(regression) 1df
|  | (Y'Y - b'X'Y) = | 5.3819e-07 |  | SS(residual)
|  | Var = (Y'Y - b'X'Y) / df = | 1.0764e-07 |  | MS(residual), n-2 df
|  | V = Var * (X'X)-1 = | 4.271e-08 | -6.5451e-13
|  |  | -6.5451e-13 | 1.5672e-17
Upper and lower limits for the intercept, b0:
|  |  | b0 lower limit = | 4.3539
|  |  | b0 upper limit = | 4.3549
|  |  | y0 = | 4.3542
Upper and lower limits for the slope, b1:
|  |  | (b1 upper - b1) / ¦b1¦ = | 0.5075%
|  |  | (b1 lower - b1) / ¦b1¦ = | -0.5075%
Estimated rate of weight loss, b1 = |  |  |  | 2.0053e-06 | g/m

## Chart Data
|  | Significance Level = | 1e-05 | (This sig level is for charting purposes only.)
|  | 1 - Conf Level = | 99.999 |  | t = | 17.8969
[5 empty] | level = | 1e-05
5
[5 empty] | Title = | Case 2 Calibration Curve, with 99.999% Confidence Bands (t = 17.9) for Estimation (red) and Prediction (blue)
[5 empty] | Line | [5 empty] | Error Components
Chart Data | [6 empty] | Est. | Est. | Pred. | Pred. |  | Case 2
|  | X | X2 | Y | Y-hat | Residual | Lower | Upper | Lower | Upper | Var | x'Vx
0 | 1 | 0 | 0 | 4.3542 | 4.3544 | -0.00019746 | 4.3507 | 4.3581 | 4.3475 | 4.3613 | 1.0764e-07 | 4.271e-08
1 | 1 | 8641 | 74666881 | 4.3374 | 4.3371 | 0.00037471 | 4.3338 | 4.3403 | 4.3304 | 4.3438 | 1.0764e-07 | 3.2569e-08
2 | 1 | 18722 | 350513284 | 4.3168 | 4.3169 | -9.3412e-05 | 4.3141 | 4.3196 | 4.3104 | 4.3233 | 1.0764e-07 | 2.3696e-08
3 | 1 | 40322 | 1625863684 | 4.2735 | 4.2735 | -4.9979e-05 | 4.2713 | 4.2758 | 4.2673 | 4.2798 | 1.0764e-07 | 1.5409e-08
4 | 1 | 64802 | 4199299204 | 4.2245 | 4.2245 | 6.0845e-05 | 4.2217 | 4.2272 | 4.218 | 4.2309 | 1.0764e-07 | 2.3697e-08
5 | 1 | 74882 | 5607313924 | 4.2038 | 4.2042 | -0.00045929 | 4.201 | 4.2075 | 4.1975 | 4.2109 | 1.0764e-07 | 3.2569e-08
6 | 1 | 84962 | 7218541444 | 4.1844 | 4.184 | 0.00036458 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
| 1 | 84962 | 0 | 4.1844 | 4.184 | 0 | 4.1802 | 4.1878 | 4.177 | 4.191 | 1.0764e-07 | 4.4626e-08
Note on A14: Beginning in this row, all formulas include "if" statements so special values (-999 or zero) are entered as needed to properly feed into other worksheets. Chart axis ranges begin at zero, so -999 helps remove extraneous points from the charts. Sumproduct functions in "Case 3" require the zeros under X-squared. The Residual plots require the zeros in the residuals column.