## Test Data
CORRELATION TEST PM CEMS AND REFERENCE METHOD TEST DATA
| PM CEMS | PM concentration | Facility: | Plant A
| response | mg/acm | Location: | Anytown, USA
Run | x | y | Emission Unit: | Boiler No. 2
1 | 4.9 | 2.1 | Test Dates: | Apr - May 2002
2 | 7.6 | 2.9
3 | 7.8 | 5.1 |  | Value | Units
4 | 9.2 | 4.8 | Emission limit: | 20 | mg/acm
5 | 10.1 | 2.9
6 | 11.1 | 6.3
10 | 12 | 8.4
11 | 18.2 | 8.4
12 | 15.7 | 4.9
13 | 13.9 | 6.3
14 | 19.7 | 13
15 | 21 | 9.1
16 | 22.3 | 13.9
17 | 24.1 | 8.7
19 | 24.7 | 12.6
20 | 26.4 | 15.9
21 | 29.1 | 16.2
22 | 29.3 | 13.5
23 | 30.5 | 15.9
24 | 33.1 | 18.5
PS-11 Correlation Test
Calculation Spreadsheet
[6 empty] | Version 2-6 | 2004-10-25

## Linear Calcs
CALCULATIONS FOR LINEAR CORRELATION
Facility: | Plant A |  |  |  | Emission Unit: | Boiler No. 2
Location: | Anytown, USA |  |  |  | Test Dates: | Apr - May 2002
| PM CEMS | PM concentration
| response | mg/acm |  | CALCULATED STATISTICAL PARAMETERS
Run | x | y | (xi-x~)2 | (yi-y~)2 | (xi-x~)(yi-y~) | y^ | (y^i-yi)2
1 | 4.9 | 2.1 | 185.9132 | 54.3169 | 100.4899 | 2.0446 | 0.0030729
2 | 7.6 | 2.9 | 119.5742 | 43.1649 | 71.843 | 3.5149 | 0.37816
3 | 7.8 | 5.1 | 115.2402 | 19.0969 | 46.9119 | 3.6239 | 2.179
4 | 9.2 | 4.8 | 87.1422 | 21.8089 | 43.5945 | 4.3863 | 0.17116
5 | 10.1 | 2.9 | 71.1492 | 43.1649 | 55.4179 | 4.8764 | 3.9062
6 | 11.1 | 6.3 | 55.2792 | 10.0489 | 23.569 | 5.421 | 0.77264
10 | 12 | 8.4 | 42.7062 | 1.1449 | 6.9924 | 5.9111 | 6.1945
11 | 18.2 | 8.4 | 0.11223 | 1.1449 | 0.35845 | 9.2876 | 0.78777
12 | 15.7 | 4.9 | 8.0372 | 20.8849 | 12.9559 | 7.9261 | 9.1573
13 | 13.9 | 6.3 | 21.4832 | 10.0489 | 14.6929 | 6.9458 | 0.41711
14 | 19.7 | 13 | 1.3572 | 12.4609 | 4.1124 | 10.1044 | 8.3843
15 | 21 | 9.1 | 6.0762 | 0.1369 | -0.91205 | 10.8124 | 2.9323
16 | 22.3 | 13.9 | 14.1752 | 19.6249 | 16.679 | 11.5204 | 5.6627
17 | 24.1 | 8.7 | 30.9692 | 0.5929 | -4.2851 | 12.5006 | 14.4447
19 | 24.7 | 12.6 | 38.0072 | 9.7969 | 19.2965 | 12.8274 | 0.051699
20 | 26.4 | 15.9 | 61.8582 | 41.3449 | 50.572 | 13.7532 | 4.6089
21 | 29.1 | 16.2 | 111.6192 | 45.2929 | 71.1025 | 15.2236 | 0.95345
22 | 29.3 | 13.5 | 115.8852 | 16.2409 | 43.383 | 15.3325 | 3.3579
23 | 30.5 | 15.9 | 143.1612 | 41.3449 | 76.935 | 15.986 | 0.0073916
24 | 33.1 | 18.5 | 212.1392 | 81.5409 | 131.522 | 17.4019 | 1.2058
Sum | 370.7 | 189.4 | 1441.8855 | 493.202 | 785.231 |  | 65.576
Mean | 18.535 | 9.47 |  |  | y^ at mean x = | 9.47
PS-11 Correlation Test
Calculation Spreadsheet
Version 2-6 | 2004-10-25

## Linear Summary
RESULTS OF LINEAR CORRELATION
Facility: | Plant A |  |  | Emission Unit: | Boiler No. 2
Location: | Anytown, USA |  |  | Test Dates: | Apr - May 2002
|  |  | Correlation equation: |  | y = | -0.62391 | + | 0.54459 | x
Variable | Equation | Value
n | Number of data points = | 20 | Summary of Acceptance Criteria for PS-11
x~ | x~ = 1/n*(Sum of (xi)) = | 18.535 | Criterion |  | Actual | Allowable | Acceptable?
Sxx | Sxx = Sum((xi-x~)^2) = | 1441.8855 | Correlation coefficient |  | 0.92718 | > 0.85 | yes
y~ | y~ = 1/n*(Sum of (yi)) = | 9.47 | Confidence interval |  | 4.483% | < 10% | yes
Syy | Syy = Sum((yi-y~)^2) = | 493.202 | Tolerance interval |  | 15.55% | < 25% | yes
Sxy | Sxy = Sum((xi-x~)(yi-y~)) = | 785.231 | * Indicates correlation coefficient is undefined.
b0 | b0 = y~-b1x~ = | -0.62391
b1 | b1 = Sxy/Sxx = | 0.54459
SL | SL = SQRT(1/(n-2)(Sum(yi^-yi)^2)) = | 1.9087
y^mean | y^ at mean x value = | 9.47
tf | tf from table = | 2.1009
CI | CI = tf*SL*SQRT(1/n) = | 0.89667
EL | Emission Limit = | 20
CI% | CI% =CI/EL*100% = | 4.483%
n' | n' = n = | 20
vf | vf from table = | 1.3845
un' | un' from table = | 1.1765
kT | kt = un'*vf = | 1.6289
TI | TI = kt*SL = | 3.1091
TI% | TI% = TI/EL*100% = | 15.55%
Sy | Sy = SQRT(Syy/(n-1)) = | 5.0949 |  |  |  | PS-11 Correlation Test
r2 | r2 = 1-(SL^2/Sy^2) = | 0.85965 |  |  |  | Calculation Spreadsheet
r | r = SQRT((1-SL^2/Sy^2)) = | 0.92718 |  |  |  | Version 2-6 |  | 2004-10-25
* Indicates correlation coefficient is undefined.

## Polynomial Calcs
CALCULATIONS FOR POLYNOMIAL CORRELATION
Facility: | Plant A |  |  |  |  | Emission Unit: | Boiler No. 2
Location: | Anytown, USA |  |  |  |  | Test Dates: | Apr - May 2002
| PM CEMS | PM concentration
| response | mg/acm |  |  |  | CALCULATED STATISTICAL PARAMETERS
Run | x | y | x2 | x3 | x4 | xy | x2y | y^ | (y^-y)2 | delta | (yi-y~)2 | CI
1 | 4.9 | 2.1 | 24.01 | 117.649 | 576.4801 | 10.29 | 50.421 | 2.4929 | 0.15438 | 0.38195 | 54.3169 | 2.5415
2 | 7.6 | 2.9 | 57.76 | 438.976 | 3336.2176 | 22.04 | 167.504 | 3.7138 | 0.66231 | 0.17287 | 43.1649 | 1.7098
3 | 7.8 | 5.1 | 60.84 | 474.552 | 3701.5056 | 39.78 | 310.284 | 3.8064 | 1.6734 | 0.16357 | 19.0969 | 1.6632
4 | 9.2 | 4.8 | 84.64 | 778.688 | 7163.9296 | 44.16 | 406.272 | 4.4626 | 0.11386 | 0.11631 | 21.8089 | 1.4025
5 | 10.1 | 2.9 | 102.01 | 1030.301 | 10406.0401 | 29.29 | 295.829 | 4.892 | 3.968 | 0.099589 | 43.1649 | 1.2978
6 | 11.1 | 6.3 | 123.21 | 1367.631 | 15180.7041 | 69.93 | 776.223 | 5.3761 | 0.85361 | 0.090375 | 10.0489 | 1.2363
10 | 12 | 8.4 | 144 | 1728 | 20736 | 100.8 | 1209.6 | 5.8181 | 6.6664 | 0.088367 | 1.1449 | 1.2225
11 | 18.2 | 8.4 | 331.24 | 6028.568 | 109719.9376 | 152.88 | 2782.416 | 9.0242 | 0.38966 | 0.12011 | 1.1449 | 1.4252
12 | 15.7 | 4.9 | 246.49 | 3869.893 | 60757.3201 | 76.93 | 1207.801 | 7.6975 | 7.8259 | 0.10836 | 20.8849 | 1.3537
13 | 13.9 | 6.3 | 193.21 | 2685.619 | 37330.1041 | 87.57 | 1217.223 | 6.7706 | 0.22148 | 0.095907 | 10.0489 | 1.2735
14 | 19.7 | 13 | 388.09 | 7645.373 | 150613.8481 | 256.1 | 5045.17 | 9.8423 | 9.9711 | 0.12035 | 12.4609 | 1.4266
15 | 21 | 9.1 | 441 | 9261 | 194481 | 191.1 | 4013.1 | 10.5646 | 2.1452 | 0.11621 | 0.1369 | 1.4019
16 | 22.3 | 13.9 | 497.29 | 11089.567 | 247297.3441 | 309.97 | 6912.331 | 11.2994 | 6.7632 | 0.10915 | 19.6249 | 1.3586
17 | 24.1 | 8.7 | 580.81 | 13997.521 | 337340.2561 | 209.67 | 5053.047 | 12.3372 | 13.2291 | 0.098456 | 0.5929 | 1.2904
19 | 24.7 | 12.6 | 610.09 | 15069.223 | 372209.8081 | 311.22 | 7687.134 | 12.6884 | 0.007815 | 0.095864 | 9.7969 | 1.2733
20 | 26.4 | 15.9 | 696.96 | 18399.744 | 485753.2416 | 419.76 | 11081.664 | 13.6979 | 4.8494 | 0.09599 | 41.3449 | 1.2741
21 | 29.1 | 16.2 | 846.81 | 24642.171 | 717087.1761 | 471.42 | 13718.322 | 15.3447 | 0.7315 | 0.14224 | 45.2929 | 1.551
22 | 29.3 | 13.5 | 858.49 | 25153.757 | 737005.0801 | 395.55 | 11589.615 | 15.4688 | 3.8763 | 0.14915 | 16.2409 | 1.5882
23 | 30.5 | 15.9 | 930.25 | 28372.625 | 865365.0625 | 484.95 | 14790.975 | 16.2197 | 0.10221 | 0.20446 | 41.3449 | 1.8595
24 | 33.1 | 18.5 | 1095.61 | 36264.691 | 1200361.2721 | 612.35 | 20268.785 | 17.8828 | 0.3809 | 0.43072 | 81.5409 | 2.6989
Sum | 370.7 | 189.4 | 8312.81 | 208415.549 | 5576422.3277 | 4295.76 | 108583.716 |  | 64.5858 |  | 493.202
[9 empty] | Minimum> | 0.088367 |  | 1.2225
PS-11 Correlation Test
Calculation Spreadsheet
Version 2-6 | 2004-10-25

## Polynomial Summary
RESULTS OF POLYNOMIAL CORRELATION
Facility: | Plant A |  | Emission Unit: | Boiler No. 2
Location: | Anytown, USA |  | Test Dates: | Apr - May 2002
Variable | Equation | Value | Correlation equation: | y = | 0.41383 | + | 0.40633 | x | + | 0.0036688 | x2
n | Number of data points = | 20
S1 | S1 = Sum (xi) = | 370.7 | Summary of Acceptance Criteria for PS-11
S2 | S2 = Sum (xi2) = | 8312.81 | Criterion | Actual | Allowable |  | Acceptable?
S3 | S3 = Sum (xi3) = | 208415.549 | Correlation coefficient | 0.92393 | > 0.85 |  | yes
S4 | S4 = Sum (xi4) = | 5576422.3277 | Confidence interval | 6.112% | < 10% |  | yes
S5 | S5 = Sum (yi) = | 189.4 | Tolerance interval | 16.3% | < 25% |  | yes
S6 | S6 = Sum (xiyi) = | 4295.76 | * Indicates correlation coefficient is undefined.
S7 | S7 = Sum (xi2yi) = | 108583.716
det A | det A = nS2S4-S2S2S2+S1S3S2-S3S3n+S2S1S3-S4S1S1 = | 2121380546.3749 | Check for Correlation Curve Minimum/Maximum
b0 | b0 = (S5S2S4+S1S3S7+S2S6S3-S7S2S2-S3S3S5-S4S6S1)/det A = | 0.41383 | Correlation curve minimum point |  |  |  | -55.3758
b1 | b1 = (nS6S4+S5S3S2+S2S1S7-S2S6S2-S7S3n-S4S1S5)/det A = | 0.40633 | Minimum allowable x value |  |  |  | 4.9
b2 | b2 = (nS2S7+S1S6S2+S5S1S3-S2S2S5-S3S6n-S7S1S1)/det A = | 0.0036688 | Is correlation curve minimum < minimum x value? |  |  |  | yes
Sp | Sp = Sqrt((1/(n-3)Sum of (y^-Y)2) = | 1.9491 | Correlation curve maximum point |  |  |  | NA
D | D = n(S2S4-S32)+S1(S3S2-S1S4)+S2(S1S3-S22) = | 2121380546.375 | Extrapolation limit for x (125% of maximum x value) |  |  |  | NA
C0 | C0 = (S2S4-S32)/D = | 1.3758 | Is correlation curve maximum > extrapolation limit? |  |  |  | NA
C1 | C1 = (S3S2-S1S4)/D = | -0.15776
C2 | C2 = (S1S3-S22)/D = | 0.0038451
C3 | C3 = (nS4-S22)/D = | 0.019999
C4 | C4 = (S1S2-nS3)/D = | -0.00051229
C5 | C5 = (nS2-S12)/D = | 1.3594e-05
tf | tf from table = | 2.1098
EL | Emission Limit = | 20
CI | CI = tf*Sp *sqrt(delta-min) = | 1.2225
CI% | CI% = CI/EL*100= | 6.112%
vf | vf from table = | 1.4001
un' | un' from table = | 1.1947
n' | n' = 1/delta = | 11.3165
kT | kt = un'*vf = | 1.6727
TI | TI = kt*sp = | 3.2604
TI% | TI% = TI/EL *100% = | 16.3%
y~ | y~ = 1/n*(Sum of (yi)) = | 9.47
Sy | Sy = sqrt(Sum of ( yi-y~)^2/(n-1) = | 5.0949
r2 | r2 = 1-(Sp^2/Sy^2) = | 0.85364
r | r = sqrt((1-Sp^2/Sy^2)) = | 0.92393
Max-min? | b2 > 0? | minimum |  |  |  |  | PS-11 Correlation Test
xmax-min | y = -b1/2b2 = | -55.3758 |  |  |  |  | Calculation Spreadsheet
1.25xmax |  | 41.375 |  |  |  |  | Version 2-6 |  |  | 2004-10-25
* Indicates correlation coefficient is undefined.

## Logarithmic Calcs
CALCULATIONS FOR LOGARITHMIC CORRELATION
Facility: | Plant A |  |  |  | Emission Unit: | Boiler No. 2
Location: | Anytown, USA |  |  |  | Test Dates: | Apr - May 2002
| PM CEMS response |  | PM concentration
| Measured | Transformed | mg/acm |  | CALCULATED STATISTICAL PARAMETERS
Run | x | x' = ln (x) | y | (xi'-x'~)2 | (yi-y~)2 | (x'i-x'~)(yi-y~) | y^ | (y^i-yi)2
1 | 4.9 | 1.5892 | 2.1 | 1.4466 | 54.3169 | 8.8643 | -0.54331 | 6.9871
2 | 7.6 | 2.0281 | 2.9 | 0.58345 | 43.1649 | 5.0184 | 3.1108 | 0.04443
3 | 7.8 | 2.0541 | 5.1 | 0.54444 | 19.0969 | 3.2245 | 3.327 | 3.1434
4 | 9.2 | 2.2192 | 4.8 | 0.32808 | 21.8089 | 2.6749 | 4.7014 | 0.0097256
5 | 10.1 | 2.3125 | 2.9 | 0.22987 | 43.1649 | 3.15 | 5.4784 | 6.6481
6 | 11.1 | 2.4069 | 6.3 | 0.14826 | 10.0489 | 1.2206 | 6.2644 | 0.0012679
10 | 12 | 2.4849 | 8.4 | 0.094299 | 1.1449 | 0.32858 | 6.9134 | 2.2098
11 | 18.2 | 2.9014 | 8.4 | 0.011976 | 1.1449 | -0.11709 | 10.3811 | 3.9246
12 | 15.7 | 2.7537 | 4.9 | 0.001469 | 20.8849 | 0.17516 | 9.1509 | 18.0703
13 | 13.9 | 2.6319 | 6.3 | 0.025632 | 10.0489 | 0.50751 | 8.1371 | 3.375
14 | 19.7 | 2.9806 | 13 | 0.035581 | 12.4609 | 0.66587 | 11.0404 | 3.84
15 | 21 | 3.0445 | 9.1 | 0.063774 | 0.1369 | -0.093438 | 11.5724 | 6.1129
16 | 22.3 | 3.1046 | 13.9 | 0.097718 | 19.6249 | 1.3848 | 12.0725 | 3.3398
17 | 24.1 | 3.1822 | 8.7 | 0.15227 | 0.5929 | -0.30047 | 12.7187 | 16.1503
19 | 24.7 | 3.2068 | 12.6 | 0.17207 | 9.7969 | 1.2984 | 12.9235 | 0.10463
20 | 26.4 | 3.2734 | 15.9 | 0.23172 | 41.3449 | 3.0952 | 13.4776 | 5.868
21 | 29.1 | 3.3707 | 16.2 | 0.33495 | 45.2929 | 3.895 | 14.2883 | 3.6547
22 | 29.3 | 3.3776 | 13.5 | 0.34293 | 16.2409 | 2.36 | 14.3453 | 0.71454
23 | 30.5 | 3.4177 | 15.9 | 0.39155 | 41.3449 | 4.0235 | 14.6795 | 1.4897
24 | 33.1 | 3.4995 | 18.5 | 0.50062 | 81.5409 | 6.3891 | 15.3605 | 9.8562
| Sum | 55.8398 | 189.4 | 5.7373 | 493.202 | 47.7648 |  | 95.5444
| Mean | 2.792 | 9.47 |  |  | y^ at mean x = | 9.47
PS-11 Correlation Test
Calculation Spreadsheet
Version 2-6 | 2004-10-25

## Logarithmic Summary
RESULTS OF LOGARITHMIC CORRELATION
Facility: | Plant A |  |  | Emission Unit: | Boiler No. 2
Location: | Anytown, USA |  |  | Test Dates: | Apr - May 2002
|  |  | Correlation equation: | y = | -13.7742 | + | 8.3253 | Ln(x)
Variable | Equation | Value
n | Number of data points = | 20 | Summary of Acceptance Criteria for PS-11
x'~ | x'~ = 1/n*(Sum of (x'i)) = | 2.792 | Criterion | Actual |  | Allowable |  | Acceptable?
Sx'x' | Sx'x' = Sum((x'i-x'~)^2) = | 5.7373 | Correlation coefficient | 0.89192 |  | > 0.85 |  | yes
y~ | y~ = 1/n*(Sum of (yi)) = | 9.47 | Confidence interval | 5.412% |  | < 10% |  | yes
Syy | Syy = Sum((yi-y~)^2) = | 493.202 | Tolerance interval | 18.76% |  | < 25% |  | yes
Sx'y | Sx'y = Sum((x'i-x'~)(yi-y~)) = | 47.7648 | * Indicates correlation coefficient is undefined.
b0 | b0 = y~-b1x'~ = | -13.7742
b1 | b1 = Sx'y/Sx'x' = | 8.3253
SL | SL = SQRT(1/(n-2)(Sum(yi^-yi)^2)) = | 2.3039
y^mean | y^ at mean x value = | 9.47
tf | tf from table = | 2.1009
CI | CI = tf*SL*SQRT(1/n) = | 1.0823
EL | Emission Limit = | 20
CI% | CI% =CI/EL*100% = | 5.412%
n' | n' = n = | 20
vf | vf from table = | 1.3845
un' | un' from table = | 1.1765
kT | kt = un'*vf = | 1.6289
TI | TI = kt*SL = | 3.7529
TI% | TI% = TI/EL*100% = | 18.76%
Sy | Sy = SQRT(Syy/(n-1)) = | 5.0949 |  |  |  |  | PS-11 Correlation Test
r2 | r2 = 1-(SL^2/Sy^2) = | 0.79551 |  |  |  |  | Calculation Spreadsheet
r | r = SQRT((1-SL^2/Sy^2)) = | 0.89192 |  |  |  |  | Version 2-6 | 2004-10-25
* Indicates correlation coefficient is undefined.

## Exponential Calcs
CALCULATIONS FOR EXPONENTIAL CORRELATION
Facility: | Plant A |  |  |  |  | Emission Unit: | Boiler No. 2
Location: | Anytown, USA |  |  |  |  | Test Dates: | Apr - May 2002
PM concentration
| PM CEMS | mg/acm
| response | Measured | Transformed | CALCULATED STATISTICAL PARAMETERS
Run | x | y | y' = Ln(y) | (xi-x~)2 | (y'i-y'~)2 | (xi-x~)(y'i-y'~) | y'^ | (y'^i-y'i)2
1 | 4.9 | 2.1 | 0.74194 | 185.9132 | 1.7819 | 18.201 | 1.1635 | 0.17773
2 | 7.6 | 2.9 | 1.0647 | 119.5742 | 1.0244 | 11.0673 | 1.3444 | 0.078207
3 | 7.8 | 5.1 | 1.6292 | 115.2402 | 0.20032 | 4.8047 | 1.3578 | 0.073701
4 | 9.2 | 4.8 | 1.5686 | 87.1422 | 0.25826 | 4.744 | 1.4515 | 0.013708
5 | 10.1 | 2.9 | 1.0647 | 71.1492 | 1.0244 | 8.5371 | 1.5118 | 0.19991
6 | 11.1 | 6.3 | 1.8405 | 55.2792 | 0.05582 | 1.7566 | 1.5788 | 0.068512
10 | 12 | 8.4 | 2.1282 | 42.7062 | 0.0026439 | -0.33602 | 1.6391 | 0.23926
11 | 18.2 | 8.4 | 2.1282 | 0.11223 | 0.0026439 | -0.017225 | 2.0544 | 0.005455
12 | 15.7 | 4.9 | 1.5892 | 8.0372 | 0.23773 | 1.3823 | 1.8869 | 0.088616
13 | 13.9 | 6.3 | 1.8405 | 21.4832 | 0.05582 | 1.0951 | 1.7664 | 0.0055053
14 | 19.7 | 13 | 2.5649 | 1.3572 | 0.23828 | 0.56868 | 2.1548 | 0.16818
15 | 21 | 9.1 | 2.2083 | 6.0762 | 0.017282 | 0.32405 | 2.2419 | 0.0011322
16 | 22.3 | 13.9 | 2.6319 | 14.1752 | 0.30811 | 2.0899 | 2.329 | 0.091742
17 | 24.1 | 8.7 | 2.1633 | 30.9692 | 0.007484 | 0.48143 | 2.4496 | 0.081936
19 | 24.7 | 12.6 | 2.5337 | 38.0072 | 0.20874 | 2.8167 | 2.4898 | 0.0019308
20 | 26.4 | 15.9 | 2.7663 | 61.8582 | 0.47542 | 5.423 | 2.6036 | 0.026469
21 | 29.1 | 16.2 | 2.785 | 111.6192 | 0.50155 | 7.4821 | 2.7845 | 2.8584e-07
22 | 29.3 | 13.5 | 2.6027 | 115.8852 | 0.27655 | 5.6611 | 2.7979 | 0.038097
23 | 30.5 | 15.9 | 2.7663 | 143.1612 | 0.47542 | 8.2499 | 2.8783 | 0.012529
24 | 33.1 | 18.5 | 2.9178 | 212.1392 | 0.70721 | 12.2486 | 3.0524 | 0.018126
Sum |  | 189.4 | 41.5363 | 1441.8855 | 7.8599 | 96.5802 |  | 1.3907
Mean | 18.535 | 9.47 | 2.0768 |  |  | y^ at mean x = | 2.0768
PS-11 Correlation Test
Calculation Spreadsheet
Version 2-6 | 2004-10-25

## Exponential Summary
RESULTS OF EXPONENTIAL CORRELATION
Facility: | Plant A |  |  | Emission Unit: | Boiler No. 2
Location: | Anytown, USA |  |  | Test Dates: | Apr - May 2002
|  |  | Correlation equation: |  | y = | 2.3055 | e | 0.066982 | x
Variable | Equation | Value
n | Number of data points = | 20 | Summary of Acceptance Criteria for PS-11
x~ | x~ = 1/n*(Sum of (xi)) = | 18.535 | Criterion |  | Actual |  | Allowable |  | Acceptable?
Sxx | Sxx = Sum((xi-x~)^2) = | 1441.8855 | Correlation coefficient |  | 0.90179 |  | > 0.85 |  |  | yes
y'~ | y'~ = 1/n*(Sum of (y'i)) = | 2.0768 | Confidence interval |  | 5.224% |  | < 10% |  |  | yes
Sy'y' | Sy'y' = Sum((y'i-y'~)^2) = | 7.8599 | Tolerance interval |  | 18.69% |  | < 25% |  |  | yes
Sxy' | Sxy' = Sum((xi-x~)(y'i-y'~)) = | 96.5802 | * Indicates correlation coefficient is undefined.
b0' | b0' = y'~-b1x~ = | 0.8353
b0 | b0 = e^b0' = | 2.3055
b1 | b1 = Sxy'/Sxx = | 0.066982
SL | SL = SQRT(1/(n-2)(Sum(y'i^-y'i)^2)) = | 0.27796
y'^mean | y'^ at mean x value = | 2.0768
tf | tf from table = | 2.1009
CI' | CI' = tf*SL*SQRT(1/n) = | 0.13058
LCL' | LCL' = y~' - CI' | 1.9462
UCL' | UCL' = y~' + CI' | 2.2074
CI | CI = (e^UCL' - e^LCL')/2 | 1.0449
EL | Emission Limit = | 20
CI% | CI% =CI/EL*100% = | 5.224%
n' | n' = n = | 20
vf | vf from table = | 1.3845
un' | un' from table = | 1.1765
kT | kT = un'*vf = | 1.6289
TI' | TI' = kt*SL = | 0.45278
LTL' | LTL' = y~' - TI' | 1.624
UTL' | UTL' = y~' + TI' | 2.5296
TI | TI = (e^UTL' - e^LTL')/2 | 3.7374
TI% | TI% = TI/EL*100% = | 18.69%
Sy' | Sy' = SQRT(Sy'y'/(n-1)) = | 0.64318 |  |  |  | PS-11 Correlation Test
r2 | r2 = 1-(SL^2/Sy'^2) = | 0.81323 |  |  |  | Calculation Spreadsheet
r | r = SQRT((1-SL^2/Sy'^2)) = | 0.90179 |  |  |  | Version 2-6 |  | 2004-10-25
* Indicates correlation coefficient is undefined.

## Power Calcs
CALCULATIONS FOR POWER CORRELATION
Facility: | Plant A |  |  |  |  | Emission Unit: | Boiler No. 2
Location: | Anytown, USA |  |  |  |  | Test Dates: | Apr - May 2002
PM concentration
| PM CEMS response |  | mg/acm
| Measured | Transformed | Measured | Transformed | CALCULATED STATISTICAL PARAMETERS
Run | x | x' = Ln(x) | y | y' = Ln(y) | (x'i-x'~)2 | (y'i-y'~)2 | (x'i-x'~)(y'i-y'~) | y'^ | (y'I^-y'i)2
1 | 4.9 | 1.5892 | 2.1 | 0.74194 | 1.4466 | 1.7819 | 1.6055 | 0.7727 | 0.00094627
2 | 7.6 | 2.0281 | 2.9 | 1.0647 | 0.58345 | 1.0244 | 0.77308 | 1.2486 | 0.033816
3 | 7.8 | 2.0541 | 5.1 | 1.6292 | 0.54444 | 0.20032 | 0.33025 | 1.2768 | 0.12424
4 | 9.2 | 2.2192 | 4.8 | 1.5686 | 0.32808 | 0.25826 | 0.29109 | 1.4558 | 0.012737
5 | 10.1 | 2.3125 | 2.9 | 1.0647 | 0.22987 | 1.0244 | 0.48526 | 1.557 | 0.2423
6 | 11.1 | 2.4069 | 6.3 | 1.8405 | 0.14826 | 0.05582 | 0.090971 | 1.6593 | 0.032844
10 | 12 | 2.4849 | 8.4 | 2.1282 | 0.094299 | 0.0026439 | -0.01579 | 1.7439 | 0.14775
11 | 18.2 | 2.9014 | 8.4 | 2.1282 | 0.011976 | 0.0026439 | 0.005627 | 2.1955 | 0.0045208
12 | 15.7 | 2.7537 | 4.9 | 1.5892 | 0.001469 | 0.23773 | 0.018688 | 2.0353 | 0.19893
13 | 13.9 | 2.6319 | 6.3 | 1.8405 | 0.025632 | 0.05582 | 0.037826 | 1.9032 | 0.0039277
14 | 19.7 | 2.9806 | 13 | 2.5649 | 0.035581 | 0.23828 | 0.092077 | 2.2813 | 0.080434
15 | 21 | 3.0445 | 9.1 | 2.2083 | 0.063774 | 0.017282 | 0.033199 | 2.3506 | 0.020265
16 | 22.3 | 3.1046 | 13.9 | 2.6319 | 0.097718 | 0.30811 | 0.17352 | 2.4158 | 0.046714
17 | 24.1 | 3.1822 | 8.7 | 2.1633 | 0.15227 | 0.007484 | 0.033758 | 2.4999 | 0.1133
19 | 24.7 | 3.2068 | 12.6 | 2.5337 | 0.17207 | 0.20874 | 0.18952 | 2.5266 | 5.0561e-05
20 | 26.4 | 3.2734 | 15.9 | 2.7663 | 0.23172 | 0.47542 | 0.33191 | 2.5988 | 0.028077
21 | 29.1 | 3.3707 | 16.2 | 2.785 | 0.33495 | 0.50155 | 0.40987 | 2.7043 | 0.0065084
22 | 29.3 | 3.3776 | 13.5 | 2.6027 | 0.34293 | 0.27655 | 0.30795 | 2.7118 | 0.011897
23 | 30.5 | 3.4177 | 15.9 | 2.7663 | 0.39155 | 0.47542 | 0.43145 | 2.7553 | 0.00012175
24 | 33.1 | 3.4995 | 18.5 | 2.9178 | 0.50062 | 0.70721 | 0.59502 | 2.844 | 0.0054442
Sum |  |  | 189.4 | 41.5363 | 5.7373 | 7.8599 | 6.2208 |  | 1.1148
Mean |  | 2.792 |  | 2.0768 |  |  | y^ at mean x = | 2.0768
PS-11 Correlation Test
Calculation Spreadsheet
Version 2-6 | 2004-10-25

## Power Summary
RESULTS OF POWER CORRELATION
Facility: | Plant A |  |  | Emission Unit: | Boiler No. 2
Location: | Anytown, USA |  |  | Test Dates: | Apr - May 2002
|  |  | Correlation equation: |  | y = | 0.38656 | x | 1.0843
Variable | Equation | Value
n | Number of data points = | 20 | Summary of Acceptance Criteria for PS-11
x'~ | x'~ = 1/n*(Sum of (x'i)) = | 2.792 | Criterion |  | Actual |  | Allowable |  | Acceptable?
Sx'x' | Sx'x' = Sum((x'i-x'~)^2) = | 5.7373 | Correlation coefficient |  | 0.92211 |  | > 0.85 |  | yes
y'~ | y'~ = 1/n*(Sum of (y'i)) = | 2.0768 | Confidence interval |  | 4.675% |  | < 10% |  | yes
Sy'y' | Sy'y' = Sum((y'i-y'~)^2) = | 7.8599 | Tolerance interval |  | 16.62% |  | < 25% |  | yes
Sx'y' | Sx'y' = Sum((x'i-x'~)(y'i-y'~)) = | 6.2208 | * Indicates correlation coefficient is undefined.
b0' | b0' = y'~-b1x~ = | -0.95047
b0 | b0 = e^b0' = | 0.38656
b1 | b1 = Sx'y'/Sx'x' = | 1.0843
SL | SL = SQRT(1/(n-2)(Sum(y'i^-y'i)^2)) = | 0.24887
y'^mean | y'^ at mean x value = | 2.0768
tf | tf from table = | 2.1009
CI' | CI = tf*SL*SQRT(1/n) = | 0.11691
LCL' | LCL' = y~' - CI' | 1.9599
UCL' | UCL' = y~' + CI' | 2.1937
CI | CI = (e^UCL' - e^LCL')/2 | 0.93497
EL | Emission Limit = | 20
CI% | CI% =CI/EL*100% = | 4.675%
n' | n' = n = | 20
vf | vf from table = | 1.3845
un' | un' from table = | 1.1765
kT | kT = un'*vf = | 1.6289
TI' | TI' = kTSL = | 0.40538
LTL' | LTL' = y~' - TI' | 1.6714
UTL' | UTL' = y~' + TI' | 2.4822
TI | TI = (e^UTL' - e^LTL')/2 | 3.3239
TI% | TI% = TI/EL*100% = | 16.62%
Sy' | Sy' = SQRT(Sy'y'/(n-1)) = | 0.64318 | [5 empty] | PS-11 Correlation Test
r2 | r2 = 1-(SL^2/Sy'^2) = | 0.85028 | [5 empty] | Calculation Spreadsheet
r | r = SQRT((1-SL^2/Sy'^2)) = | 0.92211 | [5 empty] | Version 2-6 | 2004-10-25
* Indicates correlation coefficient is undefined.

## Best Model
SELECTION OF BEST MODEL
Facility: | Plant A |  |  |  |  | Emission Unit: | Boiler No. 2
Location: | Anytown, USA |  |  |  |  | Test Dates: | Apr - May 2002
Model | Correlation coefficient (a) | > 0.85 ? | Confidence interval half range percentage | < 10% ? | Tolerance interval half range percentage | < 25% ? | Min/max within allowable range? | Does model meet all criteria?
Linear | 0.92718 | Yes | 4.483% | Yes | 15.55% | Yes | (b) | Yes | Linear
Polynomial | 0.92393 | Yes | 6.112% | Yes | 16.3% | Yes | Yes | Yes | Polynomial
Logarithmic | 0.89192 | Yes | 5.412% | Yes | 18.76% | Yes | (b) | Yes | Logarithmic
Exponential | 0.90179 | Yes | 5.224% | Yes | 18.69% | Yes | (b) | Yes | Exponential
Power | 0.92211 | Yes | 4.675% | Yes | 16.62% | Yes | (b) | Yes | Power
| 0.1 | [6 empty] | Yes | None acceptable
(a) * indicates correlation coefficient is undefined; model does not satisfy criterion.
(b) Not applicable; criterion applies only to polynomial model.
[7 empty] | Best model: | Linear
PS-11 Correlation Test
Calculation Spreadsheet
Version 2-6 | 2004-10-25

## Predicted Values
PREDICTED PM CONCENTRATIONS
Facility: | Plant A |  |  | Emission Unit: | Boiler No. 2
Location: | Anytown, USA |  |  | Test Dates: | Apr - May 2002
PM CEMs | Predicted PM concentrations (y), in units of : |  |  | mg/acm
response (x) | Linear | Polynomial | Logarithmic | Exponential | Power
4 | 1.5544 | 2.0978 | -2.2329 | 3.0139 | 1.7379
5 | 2.099 | 2.5372 | -0.37512 | 3.2227 | 2.2136
6 | 2.6436 | 2.9839 | 1.1428 | 3.4459 | 2.6974
7 | 3.1882 | 3.4379 | 2.4261 | 3.6847 | 3.1881
8 | 3.7328 | 3.8992 | 3.5378 | 3.9399 | 3.6848
9 | 4.2774 | 4.3679 | 4.5184 | 4.2129 | 4.1868
10 | 4.822 | 4.844 | 5.3956 | 4.5047 | 4.6934
11 | 5.3665 | 5.3273 | 6.189 | 4.8168 | 5.2044
12 | 5.9111 | 5.8181 | 6.9134 | 5.1505 | 5.7193
13 | 6.4557 | 6.3161 | 7.5798 | 5.5073 | 6.2379
14 | 7.0003 | 6.8215 | 8.1968 | 5.8888 | 6.7598
15 | 7.5449 | 7.3342 | 8.7712 | 6.2967 | 7.2849
16 | 8.0895 | 7.8543 | 9.3085 | 6.733 | 7.8129
17 | 8.6341 | 8.3817 | 9.8132 | 7.1994 | 8.3437
18 | 9.1786 | 8.9164 | 10.2891 | 7.6981 | 8.8772
19 | 9.7232 | 9.4585 | 10.7392 | 8.2314 | 9.4132
20 | 10.2678 | 10.0079 | 11.1662 | 8.8017 | 9.9515
21 | 10.8124 | 10.5646 | 11.5724 | 9.4114 | 10.4922
22 | 11.357 | 11.1287 | 11.9597 | 10.0634 | 11.035
23 | 11.9016 | 11.7001 | 12.3298 | 10.7606 | 11.5799
24 | 12.4462 | 12.2789 | 12.6841 | 11.506 | 12.1268
25 | 12.9908 | 12.865 | 13.024 | 12.3031 | 12.6756
26 | 13.5353 | 13.4584 | 13.3505 | 13.1554 | 13.2262
27 | 14.0799 | 14.0592 | 13.6647 | 14.0668 | 13.7787
28 | 14.6245 | 14.6673 | 13.9675 | 15.0413 | 14.3329
29 | 15.1691 | 15.2828 | 14.2596 | 16.0833 | 14.8887
30 | 15.7137 | 15.9056 | 14.5419 | 17.1974 | 15.4462
31 | 16.2583 | 16.5357 | 14.8149 | 18.3888 | 16.0052
32 | 16.8029 | 17.1731 | 15.0792 | 19.6627 | 16.5658
33 | 17.3474 | 17.8179 | 15.3354 | 21.0249 | 17.1278 | PS-11 Correlation Test
34 | 17.892 | 18.4701 | 15.5839 | 22.4814 | 17.6913 | Calculation Spreadsheet
35 | 18.4366 | 19.1296 | 15.8252 | 24.0388 | 18.2562 | Version 2-6 | 2004-10-25

## Stats Table
VALUES FOR STATISTICAL PARAMETERS
df | tf | vf | un' (75)
3 | 3.1824 | 2.92 | 1.2661
4 | 2.7764 | 2.3724 | 1.2467
5 | 2.5706 | 2.0893 | 1.2329
6 | 2.4469 | 1.9154 | 1.2225
7 | 2.3646 | 1.7972 | 1.2145
8 | 2.306 | 1.711 | 1.208
9 | 2.2622 | 1.6452 | 1.2028
10 | 2.2281 | 1.5931 | 1.1984
11 | 2.201 | 1.5506 | 1.1947
12 | 2.1788 | 1.5153 | 1.1915
13 | 2.1604 | 1.4854 | 1.1888
14 | 2.1448 | 1.4597 | 1.1864
15 | 2.1314 | 1.4373 | 1.1842
16 | 2.1199 | 1.4176 | 1.1824
17 | 2.1098 | 1.4001 | 1.1807
18 | 2.1009 | 1.3845 | 1.1792
19 | 2.093 | 1.3704 | 1.1778
20 | 2.086 | 1.3576 | 1.1765
21 | 2.0796 | 1.346 | 1.1754
22 | 2.0739 | 1.3353 | 1.1743
23 | 2.0687 | 1.3255 | 1.1734
24 | 2.0639 | 1.3165 | 1.1725
25 | 2.0595 | 1.308 | 1.1717
26 | 2.0555 | 1.3002 | 1.1709
27 | 2.0518 | 1.2929 | 1.1702
28 | 2.0484 | 1.2861 | 1.1695
29 | 2.0452 | 1.2797 | 1.1689
30 | 2.0423 | 1.2737 | 1.1683
31 | 2.0395 | 1.268 | 1.1678
32 | 2.0369 | 1.2626 | 1.1673
33 | 2.0345 | 1.2576 | 1.1668
34 | 2.0322 | 1.2528 | 1.1663
35 | 2.0301 | 1.2482 | 1.1659
36 | 2.0281 | 1.2438 | 1.1655
37 | 2.0262 | 1.2397 | 1.1651
38 | 2.0244 | 1.2358 | 1.1647
39 | 2.0227 | 1.232 | 1.1644
40 | 2.0211 | 1.2284 | 1.1641
41 | 2.0195 | 1.2249 | 1.1637
42 | 2.0181 | 1.2216 | 1.1634
43 | 2.0167 | 1.2184 | 1.1631
44 | 2.0154 | 1.2154 | 1.1629
45 | 2.0141 | 1.2124 | 1.1626
46 | 2.0129 | 1.2096 | 1.1623
47 | 2.0117 | 1.2069 | 1.1621
48 | 2.0106 | 1.2043 | 1.1619
49 | 2.0096 | 1.2017 | 1.1616
50 | 2.0086 | 1.1993 | 1.1614
51 | 2.0076 | 1.1969 | 1.1612
52 | 2.0066 | 1.1946 | 1.161
53 | 2.0057 | 1.1924 | 1.1608
54 | 2.0049 | 1.1903 | 1.1606
55 | 2.004 | 1.1882 | 1.1604
56 | 2.0032 | 1.1862 | 1.1603
57 | 2.0025 | 1.1842 | 1.1601
58 | 2.0017 | 1.1823 | 1.1599
59 | 2.001 | 1.1805 | 1.1598
60 | 2.0003 | 1.1787 | 1.1596
61 | 1.9996 | 1.1769 | 1.1595
62 | 1.999 | 1.1752 | 1.1593
63 | 1.9983 | 1.1736 | 1.1592
64 | 1.9977 | 1.172 | 1.1591
65 | 1.9971 | 1.1704 | 1.1589
66 | 1.9966 | 1.1689 | 1.1588
67 | 1.996 | 1.1674 | 1.1587
68 | 1.9955 | 1.166 | 1.1586
69 | 1.9949 | 1.1645 | 1.1585
70 | 1.9944 | 1.1632 | 1.1583
71 | 1.9939 | 1.1618 | 1.1582
72 | 1.9935 | 1.1605 | 1.1581
73 | 1.993 | 1.1592 | 1.158
74 | 1.9925 | 1.1579 | 1.1579
75 | 1.9921 | 1.1567 | 1.1578
76 | 1.9917 | 1.1555 | 1.1577
77 | 1.9913 | 1.1543 | 1.1576
78 | 1.9908 | 1.1532 | 1.1575
79 | 1.9905 | 1.1521 | 1.1575
80 | 1.9901 | 1.151 | 1.1574
81 | 1.9897 | 1.1499 | 1.1573
82 | 1.9893 | 1.1488 | 1.1572
83 | 1.989 | 1.1478 | 1.1571
84 | 1.9886 | 1.1468 | 1.157
85 | 1.9883 | 1.1458 | 1.157
86 | 1.9879 | 1.1448 | 1.1569
87 | 1.9876 | 1.1438 | 1.1568
88 | 1.9873 | 1.1429 | 1.1567
89 | 1.987 | 1.1419 | 1.1567
90 | 1.9867 | 1.141 | 1.1566
91 | 1.9864 | 1.1401 | 1.1565
92 | 1.9861 | 1.1393 | 1.1565
93 | 1.9858 | 1.1384 | 1.1564
94 | 1.9855 | 1.1376 | 1.1563
95 | 1.9853 | 1.1367 | 1.1563
96 | 1.985 | 1.1359 | 1.1562
97 | 1.9847 | 1.1351 | 1.1562
98 | 1.9845 | 1.1343 | 1.1561
99 | 1.9842 | 1.1336 | 1.156
100 | 1.984 | 1.1328 | 1.156

## Version
PS-11 Correlation Test
Calculation Spreadsheet
Version 2-6
2004-10-25