## ReadMe
Instructions for Certification Workbook
This workbook is designed to calculate certification values and uncertainties based on data from one
or more "bracketing procedure" comparisons between NIST-traceable "reference" gas generator(s) and a
"candidate" gas generator. Certification is based on the "point-to-point" method of propagating traceability,
where generators are certified at specific setpoints that correspond to the setpoints of the NIST
standard. The generator output at each setpoint is a separate NIST-traceable quantity, independent
of the other setpoint outputs.
The main inputs to this workbook are the bracketing procedure results summaries, prepared using
a companion workbook that is specific to the equipment used during the bracketing procedures (i.e., the
candidate generator, reference generator, and mercury detector). The bracketing summaries are copied
directly into this workbook, which performs all averaging and uncertainty calculations. The output is a
certification table which includes all necessary parameters to use the "candidate" gas generator certified here
(1) as a "user" generator for calibrating emissions measurement instrumentation, or (2) as a "field-reference"
gas generator to certify other gas generators using these same workbooks.
The workbook is not structurally limited in the number of setpoints or bracketing procedure tables it
can accommodate. The prototype was designed for 8 setpoints, but will accept less data by simply
displaying blank rows. If 8 is not enough, more rows can easily be added to the uncertainty and certification
tables by copying.
The layers that perform calculations ("Uncertainty" and "Certification") have protection enabled, to
prevent users from accidentally overwriting the formulae. The only unprotected cell in either layer is the
pull-down list where the user selects the method for calculating the reproducibility uncertainty component.
The "Bracketing_Results" layer is unprotected in columns A-J, where the input data tables are pasted.
1. Workbook Organization
The workbook consists of four worksheet layers , which are displayed as tabs at the bottom of the
of the screen. Detailed descriptions of each layer are included below the step-by-step instructions.
Read Me (this sheet) - describes the workbook, explaining how to use the worksheets.
Bracketing Results - allows the user to paste the bracketing results summary data tables. These
tables are stacked vertically in columns A-J, with whatever space the user wants to be between them.
Uncertainty - calculates the overall combined uncertainty of the candidate gas concentrations,
based on all of the component uncertainties (e.g., reference gas, bracketing, reproducibility).
Certification - Tabulates the certification and uncertainty data.
2. Step by Step
Step 1 -- Import or paste all applicable bracketing summaries into the Bracketing Results layer,
ensuring that the headers are included and all formulae are converted to values.
Step 2 -- If there has been an increase or decrease in the number of bracketing summaries and/or the
number of reference generators used, execute the Ctrl-J macro to resize the workbook ranges.
Step 3 -- Select a "Reproducibility" approach from the pull-down menu in cell G13 of the Uncertainty
layer (e.g., "Type-B" for a single bracketing, "Bound on Bias" for 2 to 5, "Statistical" for more).
Step 4 -- Press F9 to recalculate the workbook
Step 5 -- Check the header information and data tables of the Uncertainty and Certification layers to
make sure everything has calculated properly before saving the workbook with a new name.
2. Bracketing Results
Other than the bracketing summary tables, the only other "content" of this layer is in columns K and L,
where there are calculations of the number of occupied rows, a data field for the reference ID, and
some calculations related to the number of reference gas generators used.
This layer contains 10 named ranges: one that contains a calculation of the number of occupied rows
in the worksheet, 8 for the data columns, and one that is used by a macro to populate the reference-
specific columns of the Uncertainty layer. The named range specifications are as follows:
| Range Name |  | Column | Range Specification
| Brack_Layer_Rows |  |  | ='Bracketing Results'!$L$1
| Bracketing_K |  | F | =OFFSET('Bracketing Results'!$F$2,0,0,Brack_Layer_Rows)
| Bracketing_Ratio |  | G | =OFFSET('Bracketing Results'!$G$2,0,0,Brack_Layer_Rows)
| Bracketing_S2 |  | D | =OFFSET('Bracketing Results'!$D$2,0,0,Brack_Layer_Rows)
| Bracketing_Uncertainty |  | E | =OFFSET('Bracketing Results'!$E$2,0,0,Brack_Layer_Rows)
| Ref_Cert_Uncertainty |  | C | =OFFSET('Bracketing Results'!$C$2,0,0,Brack_Layer_Rows)
| Reference_Cert |  | B | =OFFSET('Bracketing Results'!$B$2,0,0,Brack_Layer_Rows)
| Reference_ID |  | L | =OFFSET('Bracketing Results'!$L$2,0,0,Brack_Layer_Rows)
| Setpoint |  | A | =OFFSET('Bracketing Results'!$A$2,0,0,Brack_Layer_Rows)
| Unique_IDs |  |  | =OFFSET('Bracketing Results'!$L$12,0,0,252)
3. Uncertainty
The uncertainty calculations are described in detail in the traceability protocol. This layer receives
all of its data input from the bracketing results summaries. There are three unprotected cells in this
layer: the "Candidate Location" cell, a drop-down menu cell for selecting the method for calculating
reproducibility uncertainty component, and the last column of the reproducibility uncertainty table (i.e.,
the "Type-B Default" column). Different calculations can be used to calculate this reproducibility
uncertainty term, depending on the number of bracketing procedures that are performed (J). If J is
less than 5, the statistical approach probably should not be used. The Bound on Bias approach is
appropriate when the number of bracketing procedures is between 2 and 5. When a candidate
generator (a field unit) is being certified based on one bracketing procedure, the Type-B Default
approach should be used. Appropriate values for these defaults (which may be equipment- and/or
concentration-specific) are specified in the traceability protocol.
| Selection |  | Calculation
| Statistical |  | Uses "3-level nested experiment design" calculations, with this reproducibility uncertainty term being "level 3"
| Bound on Bias |  | Assumes a uniform distribution among bracketing results, and calculates uncertainty based on the Max/Min "bounds" of the range
| Type-B Default |  | A default value is inserted in lieu of a calculated uncertainty component
Because this workbook is designed to accommodate more than one reference gas, the number of
columns in the Uncertainty layer may vary. The workbook includes a macro, accessed with the Ctrl-J
keystroke combination, that will adjust the dimensions of the Uncertainty layer (and populate the
reference-specific uncertainty columns). The prototype workbook is configured for a single reference
gas generator, since this is the most common certification arrangement. The Ctrl-J macro needs to be
executed only when one of the "dimensions" (i.e., number of bracketing summaries and/or reference
generators) is changed.
The output of the this layer is the rightmost column: the combined uncertainty of the candidate gas
concentration. This layer includes 6 named ranges, all of which are used by the Ctrl-J macro to
populate the reference-specific cells.
| Range Name |  | Range Specification
| ColHead_RefID |  | =OFFSET(Uncertainty!$C$13,0,1,1,COUNTA(Uncertainty!$6:$6)-1)
| Heading_RefID |  | =OFFSET(Uncertainty!$C$6,0,1,1,COUNTA(Uncertainty!$6:$6)-1)
| Heading_RefSN |  | =OFFSET(Uncertainty!$C$7,0,1,1,COUNTA(Uncertainty!$6:$6)-1)
| Ref_Cert_Table |  | =OFFSET(Uncertainty!$C$42,0,1,COUNTA(Uncertainty!$A:$A)-3,COUNTA(Uncertainty!$6:$6)-1)
| Ref_Cert_u_Table |  | =OFFSET(Uncertainty!$C$16,0,1,COUNTA(Uncertainty!$A:$A)-3,COUNTA(Uncertainty!$6:$6)-1)
| Ref_Km1_Table |  | =OFFSET(Uncertainty!$C$30,0,1,COUNTA(Uncertainty!$A:$A)-3,COUNTA(Uncertainty!$6:$6)-1)
4. Certification
This layer tabulates the certification and uncertainty data. The layout is, by design, identical to the
reference table input to the bracketing spreadsheet (allowing for the possibility that the "candidate"
generator may become to "reference" generator for a subsequent certification). This table does not
include information about the reference gas generator(s). If this is required for reporting purposes,
the tables in the Bracketing Results layer should be included.

## Bracketing Results
Bracketing Procedure Results Summary | [8 empty] | Reference | Reference
>""
Operator: |  | Unspecified |  |  | Candidate ID: |  | DL57TB |  | DO NOT ERASE OR
Company/Lab: |  | Unspecified |  |  | Candidate S/N: |  | Unspecified |  | ALTER ANY CELLS
Start Date/Time: |  | 2008-06-17 12:24 |  |  | Reference ID: |  | PO24RI |  | IN THESE COLUMNS
End Date/Time: |  | 2008-06-17 15:00 |  |  | Reference S/N: |  | 0803027443 |  | Rows
48
Reference Generator Certification Parameters |  |  | Bracketing Parameters |  |  |  | Candidate Generator Parameters |  | References
|  |  | Uncertainty Terms |  | Ratio Measurements |  |  |  | 1
|  |  | StdErr Among Brackets S2 | Uncertainty of Mean Ratio u(Rmean) |  |  |  |  | 1
Setpoint ug/scm | Certified ug/scm | Uncertainty 1s ug/scm |  |  | No. of Brackets K | Average Ratio Rmean | Setpoint ug/scm | Bracketing (Ccand)j ug/scm | Reference
8.1 | 8.06 | 0.057 | 0.0067514 | 0.0061758 | 3 | 1.0782 | 8.1 | 8.69 | PO24RI | PO24RI
5.7 | 5.35 | 0.05 | 0.007203 | 0.0062459 | 3 | 1.1035 | 5.7 | 5.9037 |  | PO24RI
2.7 | 2.44 | 0.027 | 0.0018367 | 0.010811 | 3 | 1.1547 | 2.7 | 2.8174 |  | PO24RI
Operator: |  | Unspecified |  |  | Candidate ID: |  | DL57TB
Company/Lab: |  | Unspecified |  |  | Candidate S/N: |  | Unspecified
Start Date/Time: |  | 2008-09-09 09:10 |  |  | Reference ID: |  | PO24RI
End Date/Time: |  | 2008-09-09 12:28 |  |  | Reference S/N: |  | 0803027443
Reference Generator Certification Parameters |  |  | Bracketing Parameters |  |  |  | Candidate Generator Parameters
|  |  | Uncertainty Terms |  | Ratio Measurements
|  |  | StdErr Among Brackets S2 | Uncertainty of Mean Ratio u(Rmean)
Setpoint ug/scm | Certified ug/scm | Uncertainty 1s ug/scm |  |  | No. of Brackets K | Average Ratio Rmean | Setpoint ug/scm | Bracketing (Ccand)j ug/scm
8.1 | 8.06 | 0.057 | 0.010955 | 0.011139 | 3 | 1.1336 | 8.1 | 9.1366 |  | PO24RI
5.7 | 5.35 | 0.05 | 0.0025628 | 0.008644 | 3 | 1.1647 | 5.7 | 6.2309 |  | PO24RI
2.7 | 2.44 | 0.027 | 0.0024015 | 0.010219 | 3 | 1.195 | 2.7 | 2.9157 |  | PO24RI
Operator: |  | Unspecified |  |  | Candidate ID: |  | DL57TB
Company/Lab: |  | Unspecified |  |  | Candidate S/N: |  | Unspecified
Start Date/Time: |  | 2008-11-17 13:11 |  |  | Reference ID: |  | PO24RI
End Date/Time: |  | 2008-11-17 16:29 |  |  | Reference S/N: |  | 0803027443
Reference Generator Certification Parameters |  |  | Bracketing Parameters |  |  |  | Candidate Generator Parameters
|  |  | Uncertainty Terms |  | Ratio Measurements
|  |  | StdErr Among Brackets S2 | Uncertainty of Mean Ratio u(Rmean)
Setpoint ug/scm | Certified ug/scm | Uncertainty 1s ug/scm |  |  | No. of Brackets K | Average Ratio Rmean | Setpoint ug/scm | Bracketing (Ccand)j ug/scm
8.1 | 8.06 | 0.057 | 0.0036016 | 0.0066992 | 3 | 1.1087 | 8.1 | 8.9358 |  | PO24RI
5.7 | 5.35 | 0.05 | 0.0026569 | 0.0061938 | 3 | 1.1351 | 5.7 | 6.0726 |  | PO24RI
2.7 | 2.44 | 0.027 | 0.0030941 | 0.010698 | 3 | 1.177 | 2.7 | 2.8718 |  | PO24RI
Note on A11: Ed Brown: The "nominal" setpoint, used for identification, not for calculation
Note on A30: Ed Brown: The "nominal" setpoint, used for identification, not for calculation
Note on A49: Ed Brown: The "nominal" setpoint, used for identification, not for calculation
Note on H11: Ed Brown: The "nominal" setpoint, used for identification, not for calculation
Note on H30: Ed Brown: The "nominal" setpoint, used for identification, not for calculation
Note on H49: Ed Brown: The "nominal" setpoint, used for identification, not for calculation

## Uncertainty
Uncertainty Calculation
|  | Candidate Location: | Unspecified
|  | Candidate Device: | DL57TB
|  | Candidate S/N: | Unspecified
|  | Reference ID: | PO24RI
|  | Reference S/N: | 0803027443 | #N/A
|  | Certification Dates: | 17-Jun / 17-Nov-2008
| Uncertainty Calculations | [6 empty] | Calculation Area
| From Bracketing |  | From Reference Gas |  |  | Reproducibility | Combined | Do Not Erase
| No. of | Combined | PO24RI | Combined |  | Bound on Bias | Overall | Statistical | Bound on Bias | Type-B Default
Setpoint | Sets | uComparison | uRef-Cert | uReference |  | uReproducibility | uCandidate
ug/scm | J | ug/scm | ug/scm | ug/scm |  | ug/scm | ug/scm
2.7 | 2 | 0.018149 | 0.027 |  | 0.03172 | 0.028385 | 0.046274 | 0.049119 | 0.028385 | 0.014333
5.7 | 2 | 0.028527 | 0.05 |  | 0.056704 | 0.094458 | 0.1138 | 0.16318 | 0.094458 | 0.030337
8.1 | 2 | 0.051327 | 0.057 |  | 0.063035 | 0.12893 | 0.15241 | 0.22129 | 0.12893 | 0.044567
0 | [7 empty] | #VALUE! | 0 | #VALUE!
0 | [7 empty] | #VALUE! | 0 | #VALUE!
0 | [7 empty] | #VALUE! | 0 | #VALUE!
0 | [7 empty] | #VALUE! | 0 | #VALUE!
0 | [7 empty] | #VALUE! | 0 | #VALUE!
Calculation Area
Do Not Erase
|  | Setpoint, ug/scm | SK-1
|  | 2.7 | 5
|  | 5.7 | 5
|  | 8.1 | 5
Calculation Area
Do Not Erase
|  | Setpoint, ug/scm | Ref-Cert
|  | 2.7 | 2.44
|  | 5.7 | 5.35
|  | 8.1 | 8.06
Note on A14: Ed Brown: The "nominal" setpoint, used for identification, not for calculation

## Certification
Certification Report
|  | Candidate Location: | Unspecified
|  | Candidate Device: | DL57TB
|  | Candidate S/N: | Unspecified |  |  | Do Not Erase or Modify These Cells
|  | Reference ID: | PO24RI |  |  | PO24RI | PO24RI | PO24RI | PO24RI | PO24RI | PO24RI | PO24RI
|  | Reference S/N: | 0803027443 |  |  | 0803027443 | 0803027443 | 0803027443 | 0803027443 | 0803027443 | 0803027443 | 0803027443
|  | Certification Date(s): | 17-Jun / 17-Nov-2008
Candidate Generator Certification
| Certification Values |  | Certification Uncertainties
| Setpoint | Certified | 1s | 2s | Relative
| ug/scm | ug/scm | ug/scm | ug/scm | 2s, %
| 2.7 | 2.8666 | 0.046274 | 0.092548 | 3.229%
| 5.7 | 6.0673 | 0.1138 | 0.22761 | 3.751%
| 8.1 | 8.9133 | 0.15241 | 0.30483 | 3.42%
0
0
0
0
0
Note on B13: Ed Brown: The "nominal" setpoint, used for identification, not for calculation