## Read Me
Instructions for Bracketing Workbook
This workbook is designed to process the data from a single "bracketing procedure" comparison
between a NIST-traceable "reference" gas generator and "candidate" generator going through
certification. The comparisons are based on the "point-to-point" method of propagating traceability,
where generators are certified at specific setpoints that correspond to the setpoints of the NIST
standard. The generator output at each setpoint is a separate NIST-traceable quantity, independent
of the other setpoint outputs.
The main inputs to this workbook are the reference generator certification and the data files from
the bracketing procedure. The latter input files are recorded during the bracketing by the candidate
generator, reference generator, and mercury detector software. The outputs of this workbook are, for
each setpoint, a comparison ratio and several uncertainty components related to the comparisons.
It is important to note that this workbook does not calculate certified values for the candidate generator
A companion "certification workbook" incorporates the results of this workbook (and possibly others
comparing the same two generators) to calculate the certified output values at each setpoint.
The workbook is not structurally limited in the number of setpoints or "brackets per setpoint" it
can accommodate. This prototype was designed for a maximum of 7 setpoints and 6 brackets per setpoint,
but the calculations are not disrupted by use of smaller datasets. Experienced Excel users may customize
the "dimensions" of the workbook to their specific application.
All layers that perform calculations have protection enabled, to prevent users from accidentally
overwriting the formulae. User input cells are unprotected, and use a blue display font. Some calculation
cells, mainly those using functions that lookup values in the raw data layers, use a dark blue font to indicate
that the user may overwrite the formulae or values in these cells. No password is required to remove the
protection from any layer in this workbook.
1. Workbook Organization
The workbook consists of several worksheet layers , which are displayed as tabs at the bottom
of the screen. Detailed descriptions of each layer are included below the step-by-step instructions.
Read Me (this sheet) - describes the workbook, explaining how to use the worksheets.
Test_Info - allows the user to input the basic information about the bracketing procedure, the
reference gas generator, and the detector linearity check.
Cand_Raw & Ref_Raw - allows the user to paste the raw data file from both the "Candidate"
and "Reference" gas generators. This version of the workbook compares two Thermo 81i
generators, where the "display" value is recorded in the 4th column (Excel column "D").
Detector_Raw - allows the user to paste the raw data file from the detector, for this version a
Thermo 80i, where the mercury measurement is recorded in the 6th column (Excel column "F").
Det_Parse - parses the 80i data by allowing the user to structure it in a way that allows other
layers to lookup timestamps for individual measurement intervals.
Point-1, Point-2, etc. - These are the "setpoint" layers, which may have different names but are
always behind the Det_Parse layer. Each of these layers contains the formulae necessary to report
the average comparison ratio and related uncertainty components for that setpoint.
Results - Re-assembles the data from other layers into a table suitable for incorporating into the
companion certification workbook.
2. Step by Step
Step 1 -- Fill in the input cells (blue font) of the Test_Info layer.
Step 2 -- Populate the Cand_Raw and Ref_Raw layers with 81i data (optional).
Step 3 -- Populate the Detector_Raw layer with 80i data.
Step 4 -- On the Det_Parse layer, click on the "Recalculate" button to populate columns A-B.
Step 5a -- Populate columns C-E of the Det_Parse layer based on measurement intervals OR
Step 5b -- Fill in the "Detector Measurement Intervals" tables (columns M & O of the setpoint layers)
and execute the Ctrl-P macro (once on each setpoint layer) to populate the Det_Parse layer
Step 6 -- Ensure that all header information on the setpoint layers is accurate.
Step 7 -- Inspect the "Detector Measurement Intervals" tables and correct any inaccuracies. Most
table entries can be made using the Ctrl-R macro, provided the Det_Parse layer is populated
correctly, the cell pointer is in the cell to be filled, and all earlier timestamp entries are valid.
Step 8 -- Fill in the names of the spreadsheet layers in column K of the Results layer and press F9 to
recalculate all workbook formulae.
Step 9 -- Inspect all layers before saving the workbook with a new name.
3. Test Info
The blue font cells are the input ranges. The top-center box contains general test information. Any of
these cells can be linked to other workbooks, or to "_Raw" layers within this workbook (referencing
calculation layers may risk creating a "circular reference" problem). It is common practice, for
example, to use a simple pointing formula (e.g. "=Detector_Raw!A2") to populate the "Date/Time"
cells from one of the "_Raw" layers.
The bottom-center box receives information about the reference generator certifications. This table
may be populated by file links or, if the layout is the same, pasting the table directly from the source
file. The values in columns D and E are optional, included only to match the customary certification
output format. To the right of the reference generator certification data is the detector linearity data.
Note that the response units are not specified, but these linearity check responses need to be
reported in the same units as used in the "Detector_Raw" layer.
4. Cand_Raw & Ref_Raw |  |  |  | (These layers are not protected)
In order to input the raw data from each of the two Thermo 81i calibrators, the user must create a raw
data file using the iPort software. If the software was not configured to load records in real time, the
records can be loaded after the bracketing procedure is complete, provided this is done before the 81i
buffers are overwritten by newer data. This will create a *.dat file with the 1-minute raw data (create a
long records file rather than the short records). The file must be opened in Microsoft Excel, whose
"Text Import Wizard" will prompt the user to either open as "Fixed Width" or "Delimited". Click on the
"Delimited" tab and click on the "Next" button. At the next prompt, click the "Space" box and then click
"Finish" to complete the import. The file should open up in a new spreadsheet. Highlight all the data
and click "Copy". Select the "Cand_Raw" or "Ref_Raw" sheet (as appropriate) and "right-click" on cell
A1. "Paste" the new data into this sheet, ensuring that the "Time" is in Column A and '"conc" is in
Column D. The "Time" column must be altered to include the date in Excel's internal representation.
The workbook includes a macro for this purpose, executed using the Ctrl-T keystroke combination.
5. Detector_Raw |  |  | (This layer is not protected)
In order to input the raw data from the Thermo 80i detector, the user must create a raw data file using
the iPort software. If the software was not configured to load records in real time, the records can be
loaded after the bracketing procedure is complete, provided this is done before the 81i buffers are
overwritten by newer data. The software will create a *.dat file with the 1-minute raw data (create a
long records file rather than the short records). The file must be opened in Microsoft Excel, whose
"Text Import Wizard" will prompt the user to either open as "Fixed Width" or "Delimited". Click on the
"Delimited" tab and click on the "Next" button. At the next prompt, click the "Space" box and then click
"Finish" to complete the import. The file should open up in a new spreadsheet. Highlight all the data
and click "Copy". Select the "Detector_Raw" sheet and "right-click" on cell A1. "Paste" the new data
into this sheet, ensuring that the "Time" is in Column A and '"HgT" is in Column F. The "Time" column
column must be altered to include the date in Excel's internal representation. The workbook includes a
macro for this purpose, executed using the Ctrl-T keystroke combination.
This layer contains two named ranges for the data columns, defined as follows:
| Range Name |  | Column |  | Example Range Specification
| Detector_Time |  | Time/Date |  | =OFFSET(Detector_Raw!$A$1,1,0,COUNTA(Detector_Raw!$A:$A)-1)
| Detector_Table |  | Entire Table |  | =OFFSET(Detector_Raw!$A$1,1,0,COUNTA(Detector_Raw!$A:$A)-1,22)
The example specifications essentially define named ranges whose heights extend to the last
occupied row of the layer.
6. Det_Parse
This layer is used to process the data from the "Detector_Raw" layer into a form more useable by
the setpoint layers (i.e., Point-1, Point-2, etc.). The formulae used in this layer are as follows:
| =Test_Info!$G$8 |  | used in cell A2 (the first data row of the "Date & Time" column)
$A4=IF($A3<Test_Info!$G$9,MIN(IF(Detector_Time>$A3,Detector_Time,"")),NA())
This function populates the column of time/date stamps, based on the start and end times specified
in the Test_Info table.
B3=VLOOKUP(A3,Detector_Table,6)
This function populates the B column with detector response values matching the time stamps in A.
In the prototype spreadsheet, rows 5 through the bottom of the table have been converted into values
(i.e., no formulae) to reduce the workbook's ongoing computational complexity. To re-calculate or re-
populate this layer, the formulae in row 4 must be copied down beyond the end of the data, the
workbook recalculated, and the non-numeric rows deleted from the bottom of the layer. The workbook
includes a macro, executed by clicking the "Recalculate" button, that performs all of these steps and
converts rows 5 through the bottom into values.
Columns C-E must be populated by the user, by copying response values over into the appropriate
column based on the measurement intervals of the bracketing procedure. The prototype workbook
contains several calculations in the setpoint layers that make use of the entries in columns C-E, all of
which assume that these columns contain only "stable" response data.
As an alternative to populating the columns, the user may choose to enter the interval start and stop
times directly into the tables in the setpoint layers. The workbook includes a macro, executed with
the Ctrl-P keystroke combination, that will automatically populate the C-E columns of the Det_Parse
layer based on the current layer's interval tables. This macro must be executed for each setpoint layer
after the interval start and stop times are entered in that layer.
The Det_Parse layer contains six named ranges for the data columns, defined as follows:
| Range Name |  | Column |  | Range Specification
| D_Time |  | A |  | =OFFSET('Det_Parse'!$A$2,1,0,COUNTA('Det_Parse'!$A:$A)-1)
| D_Time1 |  | A |  | =OFFSET('Det_Parse'!$A$2,0,0,COUNTA('Det_Parse'!$A:$A)-1)
| D_Conc |  | B |  | =OFFSET('Det_Parse'!$B$2,1,0,COUNTA('Det_Parse'!$A:$A)-1)
| D_Ref |  | C |  | =OFFSET('Det_Parse'!$C$2,1,0,COUNTA('Det_Parse'!$A:$A)-1)
| D_Cand |  | D |  | =OFFSET('Det_Parse'!$D$2,1,0,COUNTA('Det_Parse'!$A:$A)-1)
| D_Zero |  | E |  | =OFFSET('Det_Parse'!$E$2,1,0,COUNTA('Det_Parse'!$A:$A)-1)
7. Setpoints: Point-1, Point-2, etc.
These layers perform, for each setpoint, all calculations related to bracketing ratio and uncertainty
components. The "input" area of this layer is to the right, where the interval start and stop times are
tabulated. The prototype spreadsheet has this table populated with formulae to calculate the times
based on the user's entries in columns C-E of the "Det_Parse" layer (notwithstanding the option to
enter the times directly in these layers and populate the Det_Parse columns using the Ctrl-P macro,
as mentioned above). It is important to inspect this table carefully, overwriting the formulae where
necessary to have the table accurately reflect the first and last record of each measurement interval.
Often, once the first "valid" timestamp is entered near the top of the table, the formulae will calculate
the remaining timestamps. Missing or overwritten formulae in this layer can be replaced by placing the
cell pointer in the affected cell and pressing the Ctrl-R keystroke combination.
Once all timestamps are entered correctly, most of the calculations should proceed to completion
with a simple press of the {Calc} (F9) key. If, while investigating abnormal results, it is found that a
particular measurement interval needs to be "dropped" from the set, do so by simply overwriting
text comments (or simply a space) into the associated timestamp cells
8. Results
This layer is a simple lookup table that brings all of the data from the other layers into a format
suitable for passing along to the Certification Workbook. The only required input is the names of the
setpoint layers (e.g., Point-1, Point-2, etc.) in the table to the right.

## Test_Info
Operator: |  | Unspecified
Company/Lab: |  | Unspecified
Candidate ID: |  | DL57TB
Candidate S/N: |  | Unspecified
Reference ID: |  | PO24RI
Reference S/N: |  | 0803027443
Start Date/Time: |  | 2008-06-17 12:24
End Date/Time: |  | 2008-06-17 15:00
Barometric Pressure: |  | 29.17 | in Hg
Reference Generator Certification |  |  |  |  | Detector Linearity
Certification Values |  | Certification Uncertainties |  |  | Injected ug/scm | Response | Linearity Uncertainty Factor
Setpoint | Certified | 1s | 2s | Relative
ug/scm | ug/scm | ug/scm | ug/scm | 2s, % | 0 | -0.0031243
2.7 | 2.44 | 0.027 | 0.054 | 2.21% | 2.44 | 2.5228 | Reserved
5.7 | 5.35 | 0.05 | 0.1 | 1.87% | 5.35 | 5.5633 | Reserved
8.1 | 8.06 | 0.057 | 0.114 | 1.41% | 8.06 | 8.0628 | Reserved
Note on A15: Ed Brown: The "nominal" setpoint, used for identification, not for calculation

## Cand_Raw
;; | Model | 81I- | (prog | iSeries | 81i | 01.00.00.078)
;;
Time | Date | Flags | conc | syssp | hgflo | dlflo | ctemp | pres
39616.3535 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.3054 | 8755.5 | 13.1563
39616.3542 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.3176 | 8765.26 | 13.0763
39616.3549 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.2932 | 8755.5 | 13.0847
39616.3556 | 2008-06-17 | -------- | 8.1127 | 1.9672 | 10.2932 | 8760.38 | 13.2613
39616.3562 | 2008-06-17 | -------- | 8.1117 | 1.9672 | 10.2932 | 8770.14 | 13.1016
39616.3569 | 2008-06-17 | -------- | 8.1137 | 1.9672 | 10.3176 | 8765.26 | 13.1942
39616.3576 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.2932 | 8760.38 | 13.015
39616.3583 | 2008-06-17 | -------- | 8.1129 | 1.9672 | 10.2687 | 8750.62 | 13.1563
39616.359 | 2008-06-17 | -------- | 8.1135 | 1.9672 | 10.2932 | 8770.14 | 13.1164
39616.3597 | 2008-06-17 | -------- | 8.1105 | 1.9672 | 10.3054 | 8745.75 | 13.1479
39616.3604 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.2932 | 8760.38 | 13.0594
39616.3611 | 2008-06-17 | -------- | 8.1125 | 1.9672 | 10.3054 | 8745.75 | 13.1921
39616.3618 | 2008-06-17 | -------- | 8.1126 | 1.9672 | 10.2932 | 8755.5 | 13.1269
39616.3625 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.3054 | 8745.75 | 13.1542
39616.3632 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.2932 | 8750.62 | 13.1795
39616.3639 | 2008-06-17 | -------- | 8.1137 | 1.9672 | 10.3176 | 8755.5 | 13.1185
39616.3646 | 2008-06-17 | -------- | 8.1126 | 1.9672 | 10.3176 | 8755.5 | 13.2089
39616.3653 | 2008-06-17 | -------- | 8.1135 | 1.9672 | 10.3054 | 8755.5 | 13.0995
39616.366 | 2008-06-17 | -------- | 8.1142 | 1.9672 | 10.3054 | 8760.38 | 13.2256
39616.3667 | 2008-06-17 | -------- | 8.1142 | 1.9672 | 10.3054 | 8765.26 | 13.0425
39616.3674 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.3054 | 8750.62 | 13.1058
39616.3681 | 2008-06-17 | -------- | 8.1141 | 1.9672 | 10.3054 | 8750.62 | 13.255
39616.3688 | 2008-06-17 | -------- | 8.1146 | 1.9672 | 10.2809 | 8760.38 | 13.1795
39616.3694 | 2008-06-17 | -------- | 8.1132 | 1.9672 | 10.2932 | 8750.62 | 13.0319
39616.3701 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.3176 | 8745.75 | 13.2005
39616.3708 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.3176 | 8750.62 | 13.0995
39616.3715 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.2809 | 8745.75 | 13.1058
39616.3722 | 2008-06-17 | -------- | 8.1133 | 1.9672 | 10.2932 | 8750.62 | 13.1521
39616.3729 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.3054 | 8750.62 | 13.0868
39616.3736 | 2008-06-17 | -------- | 3.9905 | 1.9672 | 10.3176 | 9941.42 | 13.1984
39616.3743 | 2008-06-17 | -------- | 0 | 1.9672 | 10.2932 | 9951.18 | 13.0636
39616.375 | 2008-06-17 | -------- | 5.3971 | 1.9672 | 10.2932 | 8750.62 | 13.1648
39616.3757 | 2008-06-17 | -------- | 8.1133 | 1.9672 | 10.2932 | 8755.5 | 13.1542
39616.3764 | 2008-06-17 | -------- | 8.1126 | 1.9672 | 10.2809 | 8750.62 | 13.2319
39616.3771 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.3054 | 8755.5 | 13.1079
39616.3778 | 2008-06-17 | -------- | 8.1129 | 1.9672 | 10.3054 | 8745.75 | 13.0847
39616.3785 | 2008-06-17 | -------- | 8.1155 | 1.9672 | 10.2932 | 8755.5 | 13.2026
39616.3792 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.2809 | 8755.5 | 13.1984
39616.3799 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.2932 | 8760.38 | 13.0509
39616.3806 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.2809 | 8750.62 | 13.1584
39616.3812 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.3176 | 8755.5 | 13.1711
39616.3819 | 2008-06-17 | -------- | 8.112 | 1.9672 | 10.3176 | 8750.62 | 13.15
39616.3826 | 2008-06-17 | -------- | 8.1133 | 1.9672 | 10.2809 | 8750.62 | 13.2361
39616.3833 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.2809 | 8750.62 | 13.0657
39616.384 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.2932 | 8745.75 | 13.2193
39616.3847 | 2008-06-17 | -------- | 8.1126 | 1.9672 | 10.2932 | 8750.62 | 13.0805
39616.3854 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.2932 | 8750.62 | 13.2382
39616.3861 | 2008-06-17 | -------- | 8.1132 | 1.9672 | 10.3054 | 8750.62 | 13.0974
39616.3868 | 2008-06-17 | -------- | 8.1146 | 1.9672 | 10.3054 | 8755.5 | 13.1669
39616.3875 | 2008-06-17 | -------- | 8.1145 | 1.9672 | 10.2809 | 8755.5 | 13.0974
39616.3882 | 2008-06-17 | -------- | 8.1128 | 1.9672 | 10.2932 | 8750.62 | 13.129
39616.3889 | 2008-06-17 | -------- | 8.1129 | 1.9672 | 10.3054 | 8755.5 | 13.1858
39616.3896 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.2932 | 8745.75 | 13.0277
39616.3903 | 2008-06-17 | -------- | 8.1129 | 1.9672 | 10.2932 | 8745.75 | 13.0425
39616.391 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.3176 | 8745.75 | 13.1963
39616.3917 | 2008-06-17 | -------- | 8.1124 | 1.9672 | 10.2932 | 8740.87 | 13.0425
39616.3924 | 2008-06-17 | -------- | 8.1136 | 1.9672 | 10.3054 | 8745.75 | 13.1648
39616.3931 | 2008-06-17 | -------- | 8.1127 | 1.9672 | 10.3054 | 8750.62 | 13.2172
39616.3938 | 2008-06-17 | -------- | 8.1135 | 1.9672 | 10.3054 | 8750.62 | 13.07
39616.3944 | 2008-06-17 | -------- | 8.1127 | 1.9672 | 10.2687 | 8755.5 | 13.1563
39616.3951 | 2008-06-17 | -------- | 8.1135 | 1.9672 | 10.2932 | 8750.62 | 13.1584
39616.3958 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.3054 | 8765.26 | 13.0467
39616.3965 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.2932 | 8750.62 | 13.1605
39616.3972 | 2008-06-17 | -------- | 8.1138 | 1.9672 | 10.3054 | 8750.62 | 13.2298
39616.3979 | 2008-06-17 | -------- | 8.1127 | 1.9672 | 10.2809 | 8750.62 | 13.0552
39616.3986 | 2008-06-17 | -------- | 8.1114 | 1.9672 | 10.3054 | 8750.62 | 13.1416
39616.3993 | 2008-06-17 | -------- | 8.1133 | 1.9672 | 10.2932 | 8750.62 | 13.1795
39616.4 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.2932 | 8755.5 | 13.0594
39616.4007 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.2932 | 8740.87 | 13.1732
39616.4014 | 2008-06-17 | -------- | 8.1132 | 1.9672 | 10.2932 | 8750.62 | 13.0847
39616.4021 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.2809 | 8745.75 | 13.1858
39616.4028 | 2008-06-17 | -------- | 8.1128 | 1.9672 | 10.3054 | 8750.62 | 13.0826
39616.4035 | 2008-06-17 | -------- | 8.1131 | 1.9672 | 10.3054 | 8745.75 | 13.1942
39616.4042 | 2008-06-17 | -------- | 8.1129 | 1.9672 | 10.3054 | 8750.62 | 13.0678
39616.4049 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.3054 | 8740.87 | 13.1753
39616.4056 | 2008-06-17 | -------- | 8.1112 | 1.9672 | 10.3176 | 8745.75 | 13.1984
39616.4062 | 2008-06-17 | -------- | 8.1106 | 1.9672 | 10.3054 | 8745.75 | 13.1185
39616.4069 | 2008-06-17 | -------- | 8.1108 | 1.9672 | 10.2809 | 8740.87 | 13.1332
39616.4076 | 2008-06-17 | -------- | 8.1107 | 1.9672 | 10.2932 | 8750.62 | 13.1121
39616.4083 | 2008-06-17 | -------- | 8.1095 | 1.9672 | 10.2932 | 8745.75 | 13.0298
39616.409 | 2008-06-17 | -------- | 8.1107 | 1.9672 | 10.3054 | 8750.62 | 13.2005
39616.4097 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.2809 | 8750.62 | 13.0721
39616.4104 | 2008-06-17 | -------- | 8.1105 | 1.9672 | 10.3054 | 8745.75 | 13.1648
39616.4111 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.2932 | 8750.62 | 13.15
39616.4118 | 2008-06-17 | -------- | 8.112 | 1.9672 | 10.3054 | 8745.75 | 13.1563
39616.4125 | 2008-06-17 | -------- | 8.1132 | 1.9672 | 10.2932 | 8755.5 | 13.1605
39616.4132 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.3054 | 8755.5 | 13.0383
39616.4139 | 2008-06-17 | -------- | 8.1108 | 1.9672 | 10.2932 | 8745.75 | 13.1584
39616.4146 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.3054 | 8745.75 | 13.2068
39616.4153 | 2008-06-17 | -------- | 8.0997 | 1.9672 | 10.3054 | 8750.62 | 13.1395
39616.416 | 2008-06-17 | -------- | 8.1099 | 1.9672 | 10.2932 | 8760.38 | 13.2277
39616.4167 | 2008-06-17 | -------- | 8.1096 | 1.9672 | 10.2687 | 8745.75 | 13.1921
39616.4174 | 2008-06-17 | -------- | 8.1105 | 1.9672 | 10.2932 | 8750.62 | 13.1185
39616.4181 | 2008-06-17 | -------- | 8.1097 | 1.9672 | 10.2809 | 8740.87 | 13.1269
39616.4187 | 2008-06-17 | -------- | 8.1101 | 1.9672 | 10.2932 | 8740.87 | 13.2089
39616.4194 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.3054 | 8750.62 | 13.0298
39616.4201 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.2809 | 8745.75 | 13.2005
39616.4208 | 2008-06-17 | -------- | 8.1098 | 1.9672 | 10.2932 | 8745.75 | 13.1079
39616.4215 | 2008-06-17 | -------- | 8.1127 | 1.9672 | 10.3054 | 8750.62 | 13.1732
39616.4222 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.3054 | 8750.62 | 13.0657
39616.4229 | 2008-06-17 | -------- | 8.1131 | 1.9672 | 10.2809 | 8745.75 | 13.1648
39616.4236 | 2008-06-17 | -------- | 8.1129 | 1.9672 | 10.2932 | 8750.62 | 13.211
39616.4243 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.2932 | 8750.62 | 13.1353
39616.425 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.2932 | 8745.75 | 13.1732
39616.4257 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.2932 | 8750.62 | 13.0615
39616.4264 | 2008-06-17 | -------- | 8.1103 | 1.9672 | 10.2932 | 8745.75 | 13.129
39616.4271 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.2932 | 8745.75 | 13.1732
39616.4278 | 2008-06-17 | -------- | 8.1111 | 1.9672 | 10.2932 | 8750.62 | 13.07
39616.4285 | 2008-06-17 | -------- | 8.1112 | 1.9672 | 10.3054 | 8745.75 | 13.1058
39616.4292 | 2008-06-17 | -------- | 8.1101 | 1.9672 | 10.3176 | 8745.75 | 13.1858
39616.4299 | 2008-06-17 | -------- | 8.1116 | 1.9672 | 10.2809 | 8750.62 | 13.2529
39616.4306 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.3176 | 8755.5 | 13.0467
39616.4313 | 2008-06-17 | -------- | 8.1126 | 1.9672 | 10.3054 | 8745.75 | 13.1837
39616.4319 | 2008-06-17 | -------- | 8.1127 | 1.9672 | 10.3176 | 8755.5 | 13.0192
39616.4326 | 2008-06-17 | -------- | 8.1122 | 1.9672 | 10.2687 | 8740.87 | 13.1206
39616.4333 | 2008-06-17 | -------- | 8.1126 | 1.9672 | 10.3054 | 8750.62 | 13.1879
39616.434 | 2008-06-17 | -------- | 8.1129 | 1.9672 | 10.3054 | 8750.62 | 13.015
39616.4347 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.3054 | 8745.75 | 13.129
39616.4354 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.3054 | 8755.5 | 13.1563
39616.4361 | 2008-06-17 | -------- | 8.1107 | 1.9672 | 10.2932 | 8755.5 | 13.0742
39616.4368 | 2008-06-17 | -------- | 8.1109 | 1.9672 | 10.2932 | 8740.87 | 13.1353
39616.4375 | 2008-06-17 | -------- | 8.1108 | 1.9672 | 10.3054 | 8740.87 | 13.2131
39616.4382 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.3176 | 8750.62 | 13.0953
39616.4389 | 2008-06-17 | -------- | 8.1114 | 1.9672 | 10.2932 | 8740.87 | 13.1479
39616.4396 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.2809 | 8750.62 | 13.1774
39616.4403 | 2008-06-17 | -------- | 8.1101 | 1.9672 | 10.2932 | 8740.87 | 13.1879
39616.441 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.3298 | 8750.62 | 13.0065
39616.4417 | 2008-06-17 | -------- | 8.112 | 1.9672 | 10.2932 | 8745.75 | 13.1248
39616.4424 | 2008-06-17 | -------- | 8.1116 | 1.9672 | 10.2932 | 8750.62 | 13.2214
39616.4431 | 2008-06-17 | -------- | 8.1109 | 1.9672 | 10.3054 | 8750.62 | 13.0932
39616.4437 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.2932 | 8745.75 | 13.1227
39616.4444 | 2008-06-17 | -------- | 8.1137 | 1.9672 | 10.3054 | 8745.75 | 13.2298
39616.4451 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.2809 | 8750.62 | 13.1037
39616.4458 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.2809 | 8745.75 | 13.0678
39616.4465 | 2008-06-17 | -------- | 8.1101 | 1.9672 | 10.3054 | 8740.87 | 13.2005
39616.4472 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.3054 | 8750.62 | 13.0678
39616.4479 | 2008-06-17 | -------- | 8.1108 | 1.9672 | 10.2932 | 8745.75 | 13.1669
39616.4486 | 2008-06-17 | -------- | 8.1106 | 1.9672 | 10.2687 | 8750.62 | 13.1332
39616.4493 | 2008-06-17 | -------- | 8.1089 | 1.9672 | 10.2809 | 8740.87 | 13.1584
39616.45 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.2809 | 8750.62 | 13.0594
39616.4507 | 2008-06-17 | -------- | 8.1103 | 1.9672 | 10.3054 | 8740.87 | 13.1206
39616.4514 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.3054 | 8750.62 | 13.1837
39616.4521 | 2008-06-17 | -------- | 8.1117 | 1.9672 | 10.2809 | 8750.62 | 13.0784
39616.4528 | 2008-06-17 | -------- | 8.112 | 1.9672 | 10.3176 | 8745.75 | 13.1332
39616.4535 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.3176 | 8760.38 | 13.0383
39616.4542 | 2008-06-17 | -------- | 8.1124 | 1.9672 | 10.3176 | 8750.62 | 13.1816
39616.4549 | 2008-06-17 | -------- | 8.1132 | 1.9672 | 10.2932 | 8750.62 | 13.0573
39616.4556 | 2008-06-17 | -------- | 8.1129 | 1.9672 | 10.2932 | 8760.38 | 13.1332
39616.4563 | 2008-06-17 | -------- | 8.1112 | 1.9672 | 10.3054 | 8750.62 | 13.2089
39616.4569 | 2008-06-17 | -------- | 8.1114 | 1.9672 | 10.2932 | 8750.62 | 13.0256
39616.4576 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.2932 | 8745.75 | 13.1921
39616.4583 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.3176 | 8750.62 | 13.0889
39616.459 | 2008-06-17 | -------- | 8.1102 | 1.9672 | 10.2809 | 8740.87 | 13.0636
39616.4597 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.2932 | 8745.75 | 13.1458
39616.4604 | 2008-06-17 | -------- | 8.1099 | 1.9672 | 10.2932 | 8740.87 | 13.0552
39616.4611 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.2932 | 8740.87 | 13.1837
39616.4618 | 2008-06-17 | -------- | 8.1122 | 1.9672 | 10.3054 | 8750.62 | 13.1332
39616.4625 | 2008-06-17 | -------- | 8.1101 | 1.9672 | 10.2932 | 8745.75 | 13.034
39616.4632 | 2008-06-17 | -------- | 8.1126 | 1.9672 | 10.3054 | 8745.75 | 13.2047
39616.4639 | 2008-06-17 | -------- | 8.1117 | 1.9672 | 10.2809 | 8755.5 | 13.0784
39616.4646 | 2008-06-17 | -------- | 8.112 | 1.9672 | 10.3176 | 8745.75 | 13.0742
39616.4653 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.2932 | 8750.62 | 13.1774
39616.466 | 2008-06-17 | -------- | 8.1132 | 1.9672 | 10.3054 | 8750.62 | 13.0784
39616.4667 | 2008-06-17 | -------- | 8.1136 | 1.9672 | 10.2809 | 8745.75 | 13.1605
39616.4674 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.3176 | 8755.5 | 13.1753
39616.4681 | 2008-06-17 | -------- | 8.1105 | 1.9672 | 10.3054 | 8750.62 | 13.0911
39616.4688 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.2809 | 8750.62 | 13.1963
39616.4694 | 2008-06-17 | -------- | 8.11 | 1.9672 | 10.2932 | 8745.75 | 13.0847
39616.4701 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.3054 | 8745.75 | 13.0235
39616.4708 | 2008-06-17 | -------- | 8.1091 | 1.9672 | 10.3054 | 8740.87 | 13.1732
39616.4715 | 2008-06-17 | -------- | 8.1103 | 1.9672 | 10.2809 | 8745.75 | 13.2026
39616.4722 | 2008-06-17 | -------- | 8.1107 | 1.9672 | 10.3054 | 8750.62 | 13.2277
39616.4729 | 2008-06-17 | -------- | 8.1116 | 1.9672 | 10.3176 | 8750.62 | 13.0362
39616.4736 | 2008-06-17 | -------- | 8.1109 | 1.9672 | 10.2932 | 8745.75 | 13.0531
39616.4743 | 2008-06-17 | -------- | 8.1103 | 1.9672 | 10.3054 | 8740.87 | 13.1121
39616.475 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.3176 | 8745.75 | 13.19
39616.4757 | 2008-06-17 | -------- | 8.1127 | 1.9672 | 10.3054 | 8755.5 | 13.034
39616.4764 | 2008-06-17 | -------- | 8.1108 | 1.9672 | 10.2932 | 8740.87 | 13.1248
39616.4771 | 2008-06-17 | -------- | 8.1137 | 1.9672 | 10.3176 | 8755.5 | 13.0953
39616.4778 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.2809 | 8750.62 | 13.0932
39616.4785 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.2932 | 8745.75 | 13.2068
39616.4792 | 2008-06-17 | -------- | 8.1109 | 1.9672 | 10.2565 | 8755.5 | 13.0657
39616.4799 | 2008-06-17 | -------- | 8.1106 | 1.9672 | 10.3054 | 8740.87 | 13.1753
39616.4806 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.3054 | 8745.75 | 13.2214
39616.4812 | 2008-06-17 | -------- | 8.1098 | 1.9672 | 10.2809 | 8755.5 | 13.0615
39616.4819 | 2008-06-17 | -------- | 8.1083 | 1.9672 | 10.2809 | 8735.99 | 13.0615
39616.4826 | 2008-06-17 | -------- | 8.1092 | 1.9672 | 10.2687 | 8740.87 | 13.1669
39616.4833 | 2008-06-17 | -------- | 8.1106 | 1.9672 | 10.3054 | 8745.75 | 13.2214
39616.484 | 2008-06-17 | -------- | 8.1106 | 1.9672 | 10.2687 | 8750.62 | 13.0425
39616.4847 | 2008-06-17 | -------- | 8.1097 | 1.9672 | 10.3054 | 8745.75 | 13.11
39616.4854 | 2008-06-17 | -------- | 8.1114 | 1.9672 | 10.2687 | 8740.87 | 13.1774
39616.4861 | 2008-06-17 | -------- | 8.112 | 1.9672 | 10.3054 | 8750.62 | 13.2172
39616.4868 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.2932 | 8755.5 | 13.0531
39616.4875 | 2008-06-17 | -------- | 8.1122 | 1.9672 | 10.3054 | 8750.62 | 13.1921
39616.4882 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.2809 | 8745.75 | 13.1521
39616.4889 | 2008-06-17 | -------- | 8.1117 | 1.9672 | 10.2932 | 8745.75 | 13.1037
39616.4896 | 2008-06-17 | -------- | 8.112 | 1.9672 | 10.2809 | 8755.5 | 13.1837
39616.4903 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.2809 | 8750.62 | 13.11
39616.491 | 2008-06-17 | -------- | 8.1116 | 1.9672 | 10.2932 | 8745.75 | 13.19
39616.4917 | 2008-06-17 | -------- | 8.1122 | 1.9672 | 10.2809 | 8750.62 | 13.1332
39616.4924 | 2008-06-17 | -------- | 8.1103 | 1.9672 | 10.2932 | 8745.75 | 13.1037
39616.4931 | 2008-06-17 | -------- | 8.11 | 1.9672 | 10.3054 | 8745.75 | 13.2026
39616.4938 | 2008-06-17 | -------- | 8.1105 | 1.9672 | 10.3054 | 8745.75 | 13.0636
39616.4944 | 2008-06-17 | -------- | 8.1097 | 1.9672 | 10.2809 | 8740.87 | 13.0065
39616.4951 | 2008-06-17 | -------- | 8.1097 | 1.9672 | 10.2932 | 8740.87 | 13.0911
39616.4958 | 2008-06-17 | -------- | 8.1097 | 1.9672 | 10.3054 | 8745.75 | 13.1837
39616.4965 | 2008-06-17 | -------- | 8.1091 | 1.9672 | 10.3054 | 8750.62 | 13.2068
39616.4972 | 2008-06-17 | -------- | 8.1108 | 1.9672 | 10.3054 | 8745.75 | 13.0784
39616.4979 | 2008-06-17 | -------- | 8.1109 | 1.9672 | 10.2932 | 8745.75 | 13.0678
39616.4986 | 2008-06-17 | -------- | 8.1128 | 1.9672 | 10.3054 | 8750.62 | 13.2319
39616.4993 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.2932 | 8755.5 | 13.0488
39616.5 | 2008-06-17 | -------- | 8.1109 | 1.9672 | 10.2932 | 8745.75 | 13.1921
39616.5007 | 2008-06-17 | -------- | 8.1137 | 1.9672 | 10.3176 | 8755.5 | 13.1458
39616.5014 | 2008-06-17 | -------- | 8.1116 | 1.9672 | 10.3054 | 8760.38 | 13.0615
39616.5021 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.2932 | 8740.87 | 13.1648
39616.5028 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.3176 | 8755.5 | 13.2089
39616.5035 | 2008-06-17 | -------- | 8.1122 | 1.9672 | 10.2932 | 8750.62 | 13.0847
39616.5042 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.2932 | 8750.62 | 13.2131
39616.5049 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.3054 | 8745.75 | 13.1837
39616.5056 | 2008-06-17 | -------- | 8.1102 | 1.9672 | 10.2809 | 8750.62 | 13.0446
39616.5062 | 2008-06-17 | -------- | 8.1102 | 1.9672 | 10.2932 | 8745.75 | 13.1374
39616.5069 | 2008-06-17 | -------- | 8.11 | 1.9672 | 10.3054 | 8740.87 | 13.2152
39616.5076 | 2008-06-17 | -------- | 8.1111 | 1.9672 | 10.2809 | 8755.5 | 13.0467
39616.5083 | 2008-06-17 | -------- | 8.1098 | 1.9672 | 10.2932 | 8740.87 | 13.0573
39616.509 | 2008-06-17 | -------- | 8.1101 | 1.9672 | 10.3054 | 8745.75 | 13.2131
39616.5097 | 2008-06-17 | -------- | 8.1127 | 1.9672 | 10.2932 | 8750.62 | 13.0657
39616.5104 | 2008-06-17 | -------- | 8.1111 | 1.9672 | 10.2932 | 8740.87 | 13.0742
39616.5111 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.2932 | 8745.75 | 13.2131
39616.5118 | 2008-06-17 | -------- | 8.1083 | 1.9672 | 10.2932 | 8750.62 | 13.1248
39616.5125 | 2008-06-17 | -------- | 8.1111 | 1.9672 | 10.3054 | 8750.62 | 13.0953
39616.5132 | 2008-06-17 | -------- | 8.1132 | 1.9672 | 10.3054 | 8750.62 | 13.1921
39616.5139 | 2008-06-17 | -------- | 8.1125 | 1.9672 | 10.2687 | 8755.5 | 13.0171
39616.5146 | 2008-06-17 | -------- | 8.1116 | 1.9672 | 10.2932 | 8750.62 | 13.1605
39616.5153 | 2008-06-17 | -------- | 6.3849 | 1.9672 | 10.2932 | 9951.18 | 13.2529
39616.516 | 2008-06-17 | -------- | 0 | 1.9672 | 10.3054 | 9941.42 | 13.11
39616.5167 | 2008-06-17 | -------- | 0 | 1.9672 | 10.2809 | 9941.42 | 13.0488
39616.5174 | 2008-06-17 | -------- | 0 | 1.9672 | 10.2687 | 9926.77 | 13.1037
39616.5181 | 2008-06-17 | -------- | 0 | 1.9672 | 10.3176 | 9941.42 | 13.2152
39616.5188 | 2008-06-17 | -------- | 0 | 1.9672 | 10.3054 | 9936.54 | 13.1416
39616.5194 | 2008-06-17 | -------- | 0 | 1.9672 | 10.2932 | 9936.54 | 13.0995
39616.5201 | 2008-06-17 | -------- | 6.3266 | 1.9672 | 10.3054 | 8750.62 | 13.1016
39616.5208 | 2008-06-17 | -------- | 8.1086 | 1.9672 | 10.2687 | 8750.62 | 13.1563
39616.5215 | 2008-06-17 | -------- | 8.1093 | 1.9672 | 10.2932 | 8755.5 | 13.1542
39616.5222 | 2008-06-17 | -------- | 8.1081 | 1.9672 | 10.3054 | 8760.38 | 13.1269
39616.5229 | 2008-06-17 | -------- | 8.1075 | 1.9672 | 10.2809 | 8750.62 | 13.1479
39616.5236 | 2008-06-17 | -------- | 8.1101 | 1.9672 | 10.3054 | 8770.14 | 13.0171
39616.5243 | 2008-06-17 | -------- | 8.1091 | 1.9672 | 10.2932 | 8750.62 | 13.19
39616.525 | 2008-06-17 | -------- | 8.1104 | 1.9672 | 10.2932 | 8760.38 | 13.2068
39616.5257 | 2008-06-17 | -------- | 8.1094 | 1.9672 | 10.2932 | 8760.38 | 13.0383
39616.5264 | 2008-06-17 | -------- | 8.1108 | 1.9672 | 10.2809 | 8750.62 | 13.1164
39616.5271 | 2008-06-17 | -------- | 8.1101 | 1.9672 | 10.2932 | 8765.26 | 13.1711
39616.5278 | 2008-06-17 | -------- | 8.1114 | 1.9672 | 10.2932 | 8755.5 | 13.0319
39616.5285 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.3054 | 8755.5 | 13.1584
39616.5292 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.3054 | 8750.62 | 13.1816
39616.5299 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.2932 | 8760.38 | 13.0129
39616.5306 | 2008-06-17 | -------- | 8.1104 | 1.9672 | 10.2932 | 8750.62 | 13.1521
39616.5312 | 2008-06-17 | -------- | 8.1109 | 1.9672 | 10.2932 | 8755.5 | 13.1984
39616.5319 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.3054 | 8755.5 | 13.0383
39616.5326 | 2008-06-17 | -------- | 8.1084 | 1.9672 | 10.2932 | 8745.75 | 13.0531
39616.5333 | 2008-06-17 | -------- | 8.1092 | 1.9672 | 10.2932 | 8750.62 | 13.1837
39616.534 | 2008-06-17 | -------- | 8.1106 | 1.9672 | 10.2809 | 8755.5 | 13.1816
39616.5347 | 2008-06-17 | -------- | 8.1097 | 1.9672 | 10.2809 | 8750.62 | 13.0678
39616.5354 | 2008-06-17 | -------- | 8.1084 | 1.9672 | 10.2809 | 8750.62 | 13.0425
39616.5361 | 2008-06-17 | -------- | 8.1076 | 1.9672 | 10.2932 | 8755.5 | 12.9861
39616.5368 | 2008-06-17 | -------- | 8.1103 | 1.9672 | 10.3054 | 8750.62 | 13.0404
39616.5375 | 2008-06-17 | -------- | 8.1089 | 1.9672 | 10.2932 | 8740.87 | 13.1311
39616.5382 | 2008-06-17 | -------- | 8.1101 | 1.9672 | 10.2687 | 8755.5 | 13.1879
39616.5389 | 2008-06-17 | -------- | 8.1109 | 1.9672 | 10.3054 | 8760.38 | 13.034
39616.5396 | 2008-06-17 | -------- | 8.1111 | 1.9672 | 10.2932 | 8745.75 | 13.0065
39616.5403 | 2008-06-17 | -------- | 8.1107 | 1.9672 | 10.2932 | 8745.75 | 13.2131
39616.541 | 2008-06-17 | -------- | 8.1114 | 1.9672 | 10.3054 | 8755.5 | 13.1879
39616.5417 | 2008-06-17 | -------- | 8.1116 | 1.9672 | 10.2932 | 8760.38 | 13.0847
39616.5424 | 2008-06-17 | -------- | 8.1088 | 1.9672 | 10.3054 | 8745.75 | 13.1037
39616.5431 | 2008-06-17 | -------- | 8.1108 | 1.9672 | 10.2809 | 8755.5 | 13.1816
39616.5437 | 2008-06-17 | -------- | 8.1108 | 1.9672 | 10.2809 | 8750.62 | 13.1963
39616.5444 | 2008-06-17 | -------- | 8.1091 | 1.9672 | 10.2809 | 8755.5 | 13.0721
39616.5451 | 2008-06-17 | -------- | 8.1096 | 1.9672 | 10.2932 | 8755.5 | 13.0277
39616.5458 | 2008-06-17 | -------- | 8.1085 | 1.9672 | 10.2687 | 8750.62 | 13.07
39616.5465 | 2008-06-17 | -------- | 8.1091 | 1.9672 | 10.2932 | 8755.5 | 13.1732
39616.5472 | 2008-06-17 | -------- | 8.1098 | 1.9672 | 10.2809 | 8750.62 | 13.2068
39616.5479 | 2008-06-17 | -------- | 8.1082 | 1.9672 | 10.2932 | 8750.62 | 13.2047
39616.5486 | 2008-06-17 | -------- | 8.1094 | 1.9672 | 10.3054 | 8750.62 | 13.1795
39616.5493 | 2008-06-17 | -------- | 8.1103 | 1.9672 | 10.2809 | 8755.5 | 13.1248
39616.55 | 2008-06-17 | -------- | 8.1103 | 1.9672 | 10.2809 | 8750.62 | 13.0615
39616.5507 | 2008-06-17 | -------- | 8.1106 | 1.9672 | 10.3054 | 8745.75 | 13.0425
39616.5514 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.2932 | 8750.62 | 13.0277
39616.5521 | 2008-06-17 | -------- | 6.0524 | 1.9672 | 8.2777 | 9721.75 | 13.1669
39616.5528 | 2008-06-17 | -------- | 5.7299 | 1.9672 | 8.2777 | 9736.39 | 13.1016
39616.5535 | 2008-06-17 | -------- | 5.7286 | 1.9672 | 8.2777 | 9731.51 | 13.1037
39616.5542 | 2008-06-17 | -------- | 5.729 | 1.9672 | 8.241 | 9726.63 | 13.1584
39616.5549 | 2008-06-17 | -------- | 5.7288 | 1.9672 | 8.2899 | 9731.51 | 13.1395
39616.5556 | 2008-06-17 | -------- | 5.7289 | 1.9672 | 8.2899 | 9731.51 | 13.0277
39616.5563 | 2008-06-17 | -------- | 5.7271 | 1.9672 | 8.2655 | 9726.63 | 13.0192
39616.5569 | 2008-06-17 | -------- | 5.7267 | 1.9672 | 8.2899 | 9736.39 | 13.1437
39616.5576 | 2008-06-17 | -------- | 5.7268 | 1.9672 | 8.2655 | 9731.51 | 13.129
39616.5583 | 2008-06-17 | -------- | 5.7268 | 1.9672 | 8.2533 | 9736.39 | 13.0721
39616.559 | 2008-06-17 | -------- | 5.7258 | 1.9672 | 8.2777 | 9731.51 | 13.1164
39616.5597 | 2008-06-17 | -------- | 5.7268 | 1.9672 | 8.2777 | 9736.39 | 13.1521
39616.5604 | 2008-06-17 | -------- | 5.7263 | 1.9672 | 8.2777 | 9736.39 | 13.1753
39616.5611 | 2008-06-17 | -------- | 5.727 | 1.9672 | 8.2655 | 9736.39 | 13.0319
39616.5618 | 2008-06-17 | -------- | 5.727 | 1.9672 | 8.2655 | 9721.75 | 13.034
39616.5625 | 2008-06-17 | -------- | 5.7277 | 1.9672 | 8.2777 | 9731.51 | 13.0657
39616.5632 | 2008-06-17 | -------- | 5.7284 | 1.9672 | 8.2777 | 9736.39 | 13.0446
39616.5639 | 2008-06-17 | -------- | 5.728 | 1.9672 | 8.2655 | 9721.75 | 13.0826
39616.5646 | 2008-06-17 | -------- | 5.7281 | 1.9672 | 8.2777 | 9736.39 | 13.169
39616.5653 | 2008-06-17 | -------- | 5.7276 | 1.9672 | 8.2899 | 9731.51 | 13.11
39616.566 | 2008-06-17 | -------- | 5.7272 | 1.9672 | 8.2777 | 9736.39 | 13.1963
39616.5667 | 2008-06-17 | -------- | 5.7289 | 1.9672 | 8.2899 | 9741.27 | 13.0192
39616.5674 | 2008-06-17 | -------- | 5.7272 | 1.9672 | 8.2655 | 9731.51 | 13.07
39616.5681 | 2008-06-17 | -------- | 5.7273 | 1.9672 | 8.2533 | 9736.39 | 13.1732
39616.5687 | 2008-06-17 | -------- | 5.7257 | 1.9672 | 8.2777 | 9736.39 | 13.0594
39616.5694 | 2008-06-17 | -------- | 5.7263 | 1.9672 | 8.2777 | 9721.75 | 13.0889
39616.5701 | 2008-06-17 | -------- | 5.7262 | 1.9672 | 8.2899 | 9736.39 | 13.2193
39616.5708 | 2008-06-17 | -------- | 5.727 | 1.9672 | 8.2533 | 9736.39 | 13.1016
39616.5715 | 2008-06-17 | -------- | 5.7268 | 1.9672 | 8.2655 | 9726.63 | 13.0826
39616.5722 | 2008-06-17 | -------- | 5.7256 | 1.9672 | 8.2899 | 9726.63 | 13.0298
39616.5729 | 2008-06-17 | -------- | 5.7252 | 1.9672 | 8.2777 | 9721.75 | 13.1079
39616.5736 | 2008-06-17 | -------- | 5.7265 | 1.9672 | 8.2777 | 9731.51 | 13.11
39616.5743 | 2008-06-17 | -------- | 5.7269 | 1.9672 | 8.2655 | 9736.39 | 13.1416
39616.575 | 2008-06-17 | -------- | 5.7248 | 1.9672 | 8.2655 | 9731.51 | 13.1795
39616.5757 | 2008-06-17 | -------- | 5.7274 | 1.9672 | 8.2655 | 9736.39 | 13.0615
39616.5764 | 2008-06-17 | -------- | 5.7277 | 1.9672 | 8.2777 | 9741.27 | 13.0509
39616.5771 | 2008-06-17 | -------- | 5.7279 | 1.9672 | 8.2655 | 9731.51 | 13.07
39616.5778 | 2008-06-17 | -------- | 5.7271 | 1.9672 | 8.2777 | 9736.39 | 13.0784
39616.5785 | 2008-06-17 | -------- | 5.7276 | 1.9672 | 8.2533 | 9726.63 | 13.1185
39616.5792 | 2008-06-17 | -------- | 5.7274 | 1.9672 | 8.2777 | 9736.39 | 13.1963
39616.5799 | 2008-06-17 | -------- | 5.7277 | 1.9672 | 8.2655 | 9736.39 | 13.1563
39616.5806 | 2008-06-17 | -------- | 5.7267 | 1.9672 | 8.2777 | 9736.39 | 13.1416
39616.5813 | 2008-06-17 | L------- | 5.2222 | 1.9672 | 5.737 | 13024.5 | 13.1584
39616.5819 | 2008-06-17 | L------- | 2.7194 | 1.9672 | 5.7003 | 13014.7 | 13.0974
39616.5826 | 2008-06-17 | L------- | 2.7179 | 1.9672 | 5.7003 | 13009.8 | 13.0171
39616.5833 | 2008-06-17 | L------- | 2.7173 | 1.9672 | 5.7003 | 13000 | 13.0847
39616.584 | 2008-06-17 | L------- | 2.7173 | 1.9672 | 5.7003 | 13004.9 | 13.1563
39616.5847 | 2008-06-17 | L------- | 2.7176 | 1.9672 | 5.7125 | 13009.8 | 13.1837
39616.5854 | 2008-06-17 | L------- | 2.717 | 1.9672 | 5.7125 | 13009.8 | 13.1227
39616.5861 | 2008-06-17 | L------- | 2.7174 | 1.9672 | 5.6881 | 13019.6 | 13.0721
39616.5868 | 2008-06-17 | L------- | 2.7176 | 1.9672 | 5.737 | 13019.6 | 13.0362
39616.5875 | 2008-06-17 | L------- | 2.7162 | 1.9672 | 5.737 | 13014.7 | 13.0002
39616.5882 | 2008-06-17 | L------- | 2.717 | 1.9672 | 5.6759 | 13009.8 | 13.0108
39616.5889 | 2008-06-17 | L------- | 2.7166 | 1.9672 | 5.7125 | 13014.7 | 13.0552
39616.5896 | 2008-06-17 | L------- | 2.7176 | 1.9672 | 5.737 | 13004.9 | 13.1669
39616.5903 | 2008-06-17 | L------- | 2.7175 | 1.9672 | 5.7248 | 13014.7 | 13.2654
39616.591 | 2008-06-17 | L------- | 2.7184 | 1.9672 | 5.7614 | 13024.5 | 13.0425
39616.5917 | 2008-06-17 | L------- | 2.7178 | 1.9672 | 5.7125 | 13014.7 | 13.0721
39616.5924 | 2008-06-17 | L------- | 2.7178 | 1.9672 | 5.7248 | 13014.7 | 13.1584
39616.5931 | 2008-06-17 | L------- | 2.7183 | 1.9672 | 5.737 | 13039.1 | 13.1248
39616.5938 | 2008-06-17 | L------- | 2.7184 | 1.9672 | 5.7125 | 13024.5 | 13.11
39616.5944 | 2008-06-17 | L------- | 2.7179 | 1.9672 | 5.7003 | 13029.3 | 13.211
39616.5951 | 2008-06-17 | L------- | 2.7186 | 1.9672 | 5.7125 | 13024.5 | 13.1669
39616.5958 | 2008-06-17 | L------- | 2.7175 | 1.9672 | 5.6881 | 13024.5 | 13.0129
39616.5965 | 2008-06-17 | L------- | 2.7179 | 1.9672 | 5.7125 | 13014.7 | 13.0995
39616.5972 | 2008-06-17 | L------- | 2.7174 | 1.9672 | 5.737 | 13024.5 | 13.1984
39616.5979 | 2008-06-17 | L------- | 2.7171 | 1.9672 | 5.7003 | 13014.7 | 13.1732
39616.5986 | 2008-06-17 | L------- | 2.7168 | 1.9672 | 5.7858 | 13024.5 | 13.1079
39616.5993 | 2008-06-17 | L------- | 2.7175 | 1.9672 | 5.7248 | 13014.7 | 12.9821
39616.6 | 2008-06-17 | L------- | 2.7164 | 1.9672 | 5.7248 | 13004.9 | 13.0446
39616.6007 | 2008-06-17 | L------- | 2.7176 | 1.9672 | 5.7614 | 13014.7 | 13.0192
39616.6014 | 2008-06-17 | L------- | 2.7174 | 1.9672 | 5.737 | 13019.6 | 13.1816
39616.6021 | 2008-06-17 | L------- | 2.7173 | 1.9672 | 5.737 | 13014.7 | 13.1332
39616.6028 | 2008-06-17 | L------- | 2.717 | 1.9672 | 5.6881 | 13024.5 | 13.1753
39616.6035 | 2008-06-17 | L------- | 2.718 | 1.9672 | 5.737 | 13029.3 | 13.0488
39616.6042 | 2008-06-17 | L----L-- | 1.6042 | 1.9672 | 5.7248 | 16522.7 | 13.0362
39616.6049 | 2008-06-17 | -----L-- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1248
39616.6056 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9633.88 | 13.0995
39616.6062 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1269
39616.6069 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.2277
39616.6076 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1311
39616.6083 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0615
39616.609 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.034
39616.6097 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0911
39616.6104 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0974
39616.6111 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0261 | 9629 | 13.0509
39616.6118 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0826
39616.6125 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1479
39616.6132 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1563
39616.6139 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1711
39616.6146 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1206
39616.6153 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0531
39616.616 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0615
39616.6167 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9619.24 | 13.1774
39616.6174 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0362
39616.6181 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1037
39616.6188 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0974
39616.6194 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.169
39616.6201 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.1542
39616.6208 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1521
39616.6215 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 12.9881
39616.6222 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9619.24 | 13.0678
39616.6229 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0953
39616.6236 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1353
39616.6243 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0805
39616.625 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.169
39616.6257 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0531
39616.6264 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 12.9841
39616.6271 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1942
39616.6278 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1837
39616.6285 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.1648
39616.6292 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1332
39616.6299 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0974
39616.6306 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0784
39616.6312 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1269
39616.6319 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1669
39616.6326 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.169
39616.6333 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 13.0742
39616.634 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.2172
39616.6347 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.1879
39616.6354 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0213
39616.6361 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0742
39616.6368 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1605
39616.6375 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.2277
39616.6382 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1837
39616.6389 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1984
39616.6396 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.169
39616.6403 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9604.6 | 13.1584
39616.641 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1248
39616.6417 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1269
39616.6424 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0319
39616.6431 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0509
39616.6438 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 12.9982
39616.6444 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0932
39616.6451 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.2026
39616.6458 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9633.88 | 13.0636
39616.6465 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9629 | 13.1458
39616.6472 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1374
39616.6479 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9638.77 | 13.1963
39616.6486 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.1879
39616.6493 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.234
39616.65 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1164
39616.6507 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.2424
39616.6514 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0953
39616.6521 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9633.88 | 13.1584
39616.6528 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9633.88 | 13.0615
39616.6535 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0678
39616.6542 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.1942
39616.6549 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9633.88 | 13.0362
39616.6556 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.1795
39616.6562 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1669
39616.6569 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9633.88 | 13.0594
39616.6576 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1353
39616.6583 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9633.88 | 13.1353
39616.659 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1016
39616.6597 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0235
39616.6604 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.0446
39616.6611 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1858
39616.6618 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.1816
39616.6625 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1711
39616.6632 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0847
39616.6639 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0235
39616.6646 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 12.9921
39616.6653 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0826
39616.666 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1584
39616.6667 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.2026
39616.6674 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9633.88 | 13.1248
39616.6681 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.19
39616.6687 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9633.88 | 13.1374
39616.6694 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0467
39616.6701 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.0932
39616.6708 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.129
39616.6715 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1458
39616.6722 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 13.1248
39616.6729 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1269
39616.6736 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0509
39616.6743 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1795
39616.675 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.1774
39616.6757 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0889
39616.6764 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1079
39616.6771 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1648
39616.6778 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1942
39616.6785 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9633.88 | 13.0911
39616.6792 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.07
39616.6799 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1079
39616.6806 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.2131
39616.6813 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.1437
39616.6819 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9629 | 13.1605
39616.6826 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0087
39616.6833 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.0319
39616.684 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0847
39616.6847 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1605
39616.6854 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.2172
39616.6861 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0721
39616.6868 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0721
39616.6875 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0932
39616.6882 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.2319
39616.6889 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9629 | 13.0932
39616.6896 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9633.88 | 13.0531
39616.6903 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.0932
39616.691 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1416
39616.6917 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1711
39616.6924 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.2298
39616.6931 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1079
39616.6937 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0087
39616.6944 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1837
39616.6951 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.1711
39616.6958 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9614.36 | 13.2047
39616.6965 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9638.77 | 12.9861
39616.6972 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0847
39616.6979 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0467
39616.6986 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.11
39616.6993 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1016
39616.7 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1942
39616.7007 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1963
39616.7014 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1816
39616.7021 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1584
39616.7028 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0065
39616.7035 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.07
39616.7042 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1458
39616.7049 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.19
39616.7056 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9624.12 | 13.1395
39616.7063 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9624.12 | 13.1164
39616.7069 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.2235
39616.7076 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9638.77 | 13.0383
39616.7083 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.169
39616.709 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9633.88 | 13.0594
39616.7097 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.034
39616.7104 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1521
39616.7111 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1816
39616.7118 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0488
39616.7125 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.129
39616.7132 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.1942
39616.7139 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.2298
39616.7146 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1669
39616.7153 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1395
39616.716 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 13.0192
39616.7167 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0932
39616.7174 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 13.1774
39616.7181 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0256
39616.7188 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0657
39616.7194 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9619.24 | 13.1374
39616.7201 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.1984
39616.7208 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.0847
39616.7215 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.2068
39616.7222 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9633.88 | 13.0868
39616.7229 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0552
39616.7236 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.0404
39616.7243 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0868
39616.725 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9633.88 | 13.1711
39616.7257 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9624.12 | 13.1479
39616.7264 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0995
39616.7271 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9633.88 | 13.1921
39616.7278 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9629 | 13.0636
39616.7285 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.2256
39616.7292 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 13.0805
39616.7299 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0889
39616.7306 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0594
39616.7312 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1753
39616.7319 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9633.88 | 13.07
39616.7326 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.1711
39616.7333 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9633.88 | 13.1037
39616.734 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0826
39616.7347 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.1795
39616.7354 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9633.88 | 13.034
39616.7361 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.1058
39616.7368 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.1374
39616.7375 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.2172
39616.7382 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1942
39616.7389 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9619.24 | 13.0742
39616.7396 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0784
39616.7403 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1584
39616.741 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 13.1879
39616.7417 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 12.974
39616.7424 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1563
39616.7431 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1984
39616.7438 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9633.88 | 13.0805
39616.7444 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1142
39616.7451 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9633.88 | 13.0383
39616.7458 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.129
39616.7465 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.2068
39616.7472 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0404
39616.7479 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9609.48 | 13.0573
39616.7486 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1627
39616.7493 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0889
39616.75 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1795
39616.7507 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9633.88 | 13.0826
39616.7514 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0974
39616.7521 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0932
39616.7528 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1837
39616.7535 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0805
39616.7542 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9629 | 13.1437
39616.7549 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9633.88 | 13.1521
39616.7556 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9633.88 | 13.15
39616.7562 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 13.1584
39616.7569 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0742
39616.7576 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1732
39616.7583 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1353
39616.759 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0383
39616.7597 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.015
39616.7604 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.034
39616.7611 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0995
39616.7618 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.1332
39616.7625 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 13.0531
39616.7632 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1669
39616.7639 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1269
39616.7646 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9624.12 | 13.0509
39616.7653 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.169
39616.766 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.1479
39616.7667 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0615
39616.7674 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1858
39616.7681 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9624.12 | 13.0784
39616.7688 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0868
39616.7694 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9604.6 | 13.1037
39616.7701 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1584
39616.7708 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1795
39616.7715 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0763
39616.7722 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.2319
39616.7729 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1458
39616.7736 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.0531
39616.7743 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.19
39616.775 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1563
39616.7757 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9619.24 | 13.0509
39616.7764 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0678
39616.7771 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.2214
39616.7778 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0932
39616.7785 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0847
39616.7792 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1605
39616.7799 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1753
39616.7806 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.1206
39616.7812 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0678
39616.7819 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.2005
39616.7826 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.1963
39616.7833 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0425
39616.784 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1984
39616.7847 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.07
39616.7854 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0993 | 9619.24 | 13.1753
39616.7861 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.2319
39616.7868 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9638.77 | 13.0784
39616.7875 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.2214
39616.7882 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.1016
39616.7889 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0742
39616.7896 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1479
39616.7903 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.2319
39616.791 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0404
39616.7917 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0615
39616.7924 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0261 | 9609.48 | 13.1353
39616.7931 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.2424
39616.7937 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1627
39616.7944 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0467
39616.7951 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.2047
39616.7958 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0826
39616.7965 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0531
39616.7972 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1984
39616.7979 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1437
39616.7986 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.1037
39616.7993 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1984
39616.8 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0531
39616.8007 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0784
39616.8014 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0319
39616.8021 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0087
39616.8028 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0974
39616.8035 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1774
39616.8042 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9614.36 | 13.1542
39616.8049 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0615
39616.8056 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0784
39616.8063 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1795
39616.8069 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1584
39616.8076 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.169
39616.8083 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1837
39616.809 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.034
39616.8097 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1648
39616.8104 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1563
39616.8111 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9614.36 | 13.1479
39616.8118 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0383
39616.8125 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.0784
39616.8132 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1711
39616.8139 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0995
39616.8146 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1437
39616.8153 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.2319
39616.816 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0721
39616.8167 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0911
39616.8174 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.2319
39616.8181 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 13.0826
39616.8187 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1437
39616.8194 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1648
39616.8201 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0636
39616.8208 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.2152
39616.8215 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1437
39616.8222 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0531
39616.8229 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.1648
39616.8236 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.15
39616.8243 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0362
39616.825 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0993 | 9614.36 | 13.1816
39616.8257 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1627
39616.8264 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0404
39616.8271 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1732
39616.8278 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 13.0995
39616.8285 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0065
39616.8292 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0953
39616.8299 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9629 | 13.2068
39616.8306 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.0573
39616.8313 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1332
39616.8319 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1774
39616.8326 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1879
39616.8333 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0362
39616.834 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.11
39616.8347 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1942
39616.8354 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1374
39616.8361 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0993 | 9614.36 | 13.034
39616.8368 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1605
39616.8375 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0256
39616.8382 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0446
39616.8389 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.2068
39616.8396 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0721
39616.8403 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.129
39616.841 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9624.12 | 13.1921
39616.8417 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9633.88 | 13.1332
39616.8424 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9633.88 | 13.0594
39616.8431 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0425
39616.8438 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1648
39616.8444 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.07
39616.8451 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0826
39616.8458 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0721
39616.8465 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1648
39616.8472 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.234
39616.8479 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0889
39616.8486 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1185
39616.8493 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1795
39616.85 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0911
39616.8507 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.11
39616.8514 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0805
39616.8521 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1332
39616.8528 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.2026
39616.8535 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.0319
39616.8542 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0847
39616.8549 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.2298
39616.8556 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1542
39616.8562 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0319
39616.8569 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0552
39616.8576 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1206
39616.8583 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.169
39616.859 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0552
39616.8597 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0932
39616.8604 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9619.24 | 13.2089
39616.8611 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0763
39616.8618 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0889
39616.8625 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.2382
39616.8632 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.2193
39616.8639 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.0002
39616.8646 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.1458
39616.8653 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0993 | 9614.36 | 13.1984
39616.866 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0383
39616.8667 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9599.72 | 13.0573
39616.8674 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0509
39616.8681 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1206
39616.8688 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9619.24 | 13.2068
39616.8694 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0826
39616.8701 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.2193
39616.8708 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.2131
39616.8715 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1079
39616.8722 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.2319
39616.8729 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.0362
39616.8736 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.2277
39616.8743 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0847
39616.875 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1016
39616.8757 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1416
39616.8764 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.0383
39616.8771 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0509
39616.8778 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.1648
39616.8785 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0261 | 9604.6 | 13.2131
39616.8792 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0974
39616.8799 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.2152
39616.8806 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0615
39616.8812 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1774
39616.8819 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0615
39616.8826 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1164
39616.8833 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9609.48 | 13.1984
39616.884 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1269
39616.8847 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0889
39616.8854 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0932
39616.8861 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0911
39616.8868 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.1648
39616.8875 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1963
39616.8882 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0192
39616.8889 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9594.84 | 13.0615
39616.8896 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.1037
39616.8903 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.211
39616.891 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0509
39616.8917 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1521
39616.8924 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9619.24 | 13.11
39616.8931 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.1164
39616.8938 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.2005
39616.8944 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.2235
39616.8951 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0383
39616.8958 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1037
39616.8965 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1458
39616.8972 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.1311
39616.8979 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.1942
39616.8986 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0425
39616.8993 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.0742
39616.9 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.2403
39616.9007 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0509
39616.9014 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0868
39616.9021 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.2277
39616.9028 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0509
39616.9035 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9609.48 | 13.1584
39616.9042 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1521
39616.9049 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0615
39616.9056 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9604.6 | 13.1395
39616.9062 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1563
39616.9069 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0636
39616.9076 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1605
39616.9083 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0087
39616.909 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9599.72 | 13.129
39616.9097 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.2068
39616.9104 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.1332
39616.9111 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.0488
39616.9118 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1437
39616.9125 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.2172
39616.9132 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0868
39616.9139 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0868
39616.9146 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0993 | 9619.24 | 13.1837
39616.9153 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9633.88 | 13.015
39616.916 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1458
39616.9167 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.2214
39616.9174 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1648
39616.9181 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1753
39616.9187 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0721
39616.9194 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1774
39616.9201 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.15
39616.9208 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 12.974
39616.9215 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9594.84 | 13.0911
39616.9222 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0953
39616.9229 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9619.24 | 13.1248
39616.9236 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0425
39616.9243 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1858
39616.925 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1669
39616.9257 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0467
39616.9264 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0235
39616.9271 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1185
39616.9278 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1416
39616.9285 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0277
39616.9292 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1563
39616.9299 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.2445
39616.9306 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0993 | 9619.24 | 13.1648
39616.9313 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.0425
39616.9319 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9594.84 | 13.1374
39616.9326 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.2131
39616.9333 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9619.24 | 13.0784
39616.934 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.0531
39616.9347 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.2382
39616.9354 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0171
39616.9361 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0383
39616.9368 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1669
39616.9375 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1795
39616.9382 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0678
39616.9389 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9619.24 | 13.1605
39616.9396 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.15
39616.9403 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.0868
39616.941 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0805
39616.9417 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1479
39616.9424 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.2361
39616.9431 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0953
39616.9437 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1016
39616.9444 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.2193
39616.9451 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1374
39616.9458 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1479
39616.9465 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0805
39616.9472 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.2193
39616.9479 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0531
39616.9486 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.2633
39616.9493 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0404
39616.95 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.0213
39616.9507 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.0636
39616.9514 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.2005
39616.9521 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9619.24 | 13.07
39616.9528 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9614.36 | 13.2047
39616.9535 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.1711
39616.9542 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0213
39616.9549 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.1206
39616.9556 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.169
39616.9563 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9624.12 | 13.1858
39616.9569 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0573
39616.9576 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1353
39616.9583 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9614.36 | 13.2403
39616.959 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0763
39616.9597 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.0889
39616.9604 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0636
39616.9611 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1584
39616.9618 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1816
39616.9625 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1121
39616.9632 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0993 | 9624.12 | 13.1627
39616.9639 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1269
39616.9646 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.1669
39616.9653 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0953
39616.966 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.2068
39616.9667 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0594
39616.9674 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.2005
39616.9681 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0721
39616.9688 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.1142
39616.9694 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.2214
39616.9701 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1227
39616.9708 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.169
39616.9715 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1563
39616.9722 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.0847
39616.9729 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9614.36 | 13.1942
39616.9736 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0404
39616.9743 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.1121
39616.975 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1921
39616.9757 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.169
39616.9764 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.129
39616.9771 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.2172
39616.9778 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.0615
39616.9785 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9614.36 | 13.1416
39616.9792 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1711
39616.9799 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0636
39616.9806 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.2068
39616.9812 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1248
39616.9819 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0171
39616.9826 | 2008-06-17 | -------- | 0 | 1.9672 | 11.1116 | 9609.48 | 13.0488
39616.9833 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9604.6 | 13.1816
39616.984 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9619.24 | 13.1227
39616.9847 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.0404
39616.9854 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0995
39616.9861 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9614.36 | 13.1479
39616.9868 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.1016
39616.9875 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1164
39616.9882 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1753
39616.9889 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0657
39616.9896 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1458
39616.9903 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.211
39616.991 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.0425
39616.9917 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.2005
39616.9924 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0871 | 9619.24 | 13.0404
39616.9931 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0573
39616.9938 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0261 | 9614.36 | 13.2047
39616.9944 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0446
39616.9951 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1563
39616.9958 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0383 | 9619.24 | 13.1374
39616.9965 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1521
39616.9972 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.169
39616.9979 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0657
39616.9986 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.234
39616.9993 | 2008-06-17 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.034
39617 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1563
39617.0007 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.2068
39617.0014 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0847
39617.0021 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.2068
39617.0028 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1079
39617.0035 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1879
39617.0042 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1921
39617.0049 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0277
39617.0056 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.1227
39617.0062 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.2256
39617.0069 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0467
39617.0076 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1837
39617.0083 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0404
39617.009 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.0383
39617.0097 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1732
39617.0104 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.0911
39617.0111 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1269
39617.0118 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0657
39617.0125 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.129
39617.0132 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9619.24 | 13.2361
39617.0139 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0742
39617.0146 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.1437
39617.0153 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1942
39617.016 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1563
39617.0167 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0319
39617.0174 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 12.9921
39617.0181 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.1711
39617.0188 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1837
39617.0194 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0911
39617.0201 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1816
39617.0208 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0805
39617.0215 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1479
39617.0222 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0763
39617.0229 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1584
39617.0236 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9633.88 | 13.0868
39617.0243 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.11
39617.025 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.2298
39617.0257 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9633.88 | 13.0235
39617.0264 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.2298
39617.0271 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0446
39617.0278 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1963
39617.0285 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.07
39617.0292 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1984
39617.0299 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0488
39617.0306 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0138 | 9599.72 | 13.0678
39617.0312 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1669
39617.0319 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9614.36 | 13.2193
39617.0326 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0953
39617.0333 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9619.24 | 13.0805
39617.034 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.2256
39617.0347 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1248
39617.0354 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.2152
39617.0361 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1353
39617.0368 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0889
39617.0375 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.2633
39617.0382 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0467
39617.0389 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.2089
39617.0396 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1416
39617.0403 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0277
39617.041 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0805
39617.0417 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.1711
39617.0424 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.034
39617.0431 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1921
39617.0437 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.15
39617.0444 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0425
39617.0451 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0721
39617.0458 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.15
39617.0465 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.19
39617.0472 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.1584
39617.0479 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0573
39617.0486 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1142
39617.0493 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.2235
39617.05 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.1795
39617.0507 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9633.88 | 13.0108
39617.0514 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1837
39617.0521 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.1079
39617.0528 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.0763
39617.0535 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.2319
39617.0542 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.2298
39617.0549 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0277
39617.0556 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.0298
39617.0563 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.2047
39617.0569 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9614.36 | 13.1353
39617.0576 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0932
39617.0583 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.1605
39617.059 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1753
39617.0597 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1311
39617.0604 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1206
39617.0611 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.1395
39617.0618 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.0192
39617.0625 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1963
39617.0632 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.2089
39617.0639 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0868
39617.0646 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.2005
39617.0653 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0261 | 9614.36 | 13.129
39617.066 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1521
39617.0667 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1206
39617.0674 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1206
39617.0681 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1816
39617.0687 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1816
39617.0694 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0192
39617.0701 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1016
39617.0708 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1521
39617.0715 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.07
39617.0722 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1921
39617.0729 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.211
39617.0736 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0847
39617.0743 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.2172
39617.075 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0256
39617.0757 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0826
39617.0764 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1164
39617.0771 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1648
39617.0778 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0192
39617.0785 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.0826
39617.0792 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1185
39617.0799 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.2068
39617.0806 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0763
39617.0813 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0425
39617.0819 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1837
39617.0826 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1227
39617.0833 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.0467
39617.084 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1732
39617.0847 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.2361
39617.0854 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.0594
39617.0861 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1248
39617.0868 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1605
39617.0875 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0467
39617.0882 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.0636
39617.0889 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 12.9881
39617.0896 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.0784
39617.0903 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1037
39617.091 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1479
39617.0917 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1227
39617.0924 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.129
39617.0931 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.19
39617.0938 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1521
39617.0944 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0404
39617.0951 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0974
39617.0958 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0721
39617.0965 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1858
39617.0972 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9633.88 | 13.1542
39617.0979 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0615
39617.0986 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1942
39617.0993 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.234
39617.1 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0261 | 9619.24 | 13.0678
39617.1007 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0826
39617.1014 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1795
39617.1021 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.2026
39617.1028 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.1984
39617.1035 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0995
39617.1042 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0889
39617.1049 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1732
39617.1056 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1921
39617.1062 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1479
39617.1069 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0488
39617.1076 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1711
39617.1083 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 13.0298
39617.109 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0993 | 9614.36 | 13.2089
39617.1097 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.2319
39617.1104 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9629 | 13.0425
39617.1111 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.2214
39617.1118 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0573
39617.1125 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0826
39617.1132 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1942
39617.1139 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1437
39617.1146 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.0742
39617.1153 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.0932
39617.116 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1963
39617.1167 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0573
39617.1174 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0636
39617.1181 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.2235
39617.1188 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.15
39617.1194 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9619.24 | 13.0425
39617.1201 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.11
39617.1208 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.2152
39617.1215 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0953
39617.1222 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.1774
39617.1229 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0615
39617.1236 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.2193
39617.1243 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1037
39617.125 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1079
39617.1257 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1058
39617.1264 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1774
39617.1271 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0805
39617.1278 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1963
39617.1285 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0277
39617.1292 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.0805
39617.1299 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.2005
39617.1306 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.0615
39617.1312 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1774
39617.1319 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9629 | 13.0847
39617.1326 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9619.24 | 13.1858
39617.1333 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1627
39617.134 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0256
39617.1347 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1984
39617.1354 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9638.77 | 13.0256
39617.1361 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1227
39617.1368 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1837
39617.1375 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.169
39617.1382 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.07
39617.1389 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1311
39617.1396 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.19
39617.1403 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1479
39617.141 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.15
39617.1417 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.1311
39617.1424 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1269
39617.1431 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.2068
39617.1438 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.1648
39617.1444 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9619.24 | 13.0763
39617.1451 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.2131
39617.1458 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.034
39617.1465 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0256
39617.1472 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1669
39617.1479 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.2256
39617.1486 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9629 | 13.0594
39617.1493 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.2047
39617.15 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1479
39617.1507 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0784
39617.1514 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1774
39617.1521 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0277
39617.1528 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0213
39617.1535 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9609.48 | 13.1542
39617.1542 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1795
39617.1549 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0467
39617.1556 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.1016
39617.1562 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.19
39617.1569 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 13.0108
39617.1576 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0826
39617.1583 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1669
39617.159 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1879
39617.1597 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 13.0213
39617.1604 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1206
39617.1611 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.2403
39617.1618 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9633.88 | 13.0721
39617.1625 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.1584
39617.1632 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.2047
39617.1639 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9629 | 12.9941
39617.1646 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.0995
39617.1653 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.1732
39617.166 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0721
39617.1667 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0594
39617.1674 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.2193
39617.1681 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1311
39617.1687 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.2026
39617.1694 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1669
39617.1701 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0298
39617.1708 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0826
39617.1715 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9604.6 | 13.1837
39617.1722 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0404
39617.1729 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9604.6 | 13.1858
39617.1736 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9619.24 | 13.1353
39617.1743 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0108
39617.175 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0467
39617.1757 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.1142
39617.1764 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0974
39617.1771 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0742
39617.1778 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9624.12 | 13.2193
39617.1785 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1332
39617.1792 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1269
39617.1799 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9599.72 | 13.1669
39617.1806 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1837
39617.1813 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 12.9679
39617.1819 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1521
39617.1826 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0636
39617.1833 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0129
39617.184 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9594.84 | 13.0805
39617.1847 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.0995
39617.1854 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9624.12 | 13.1437
39617.1861 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0425
39617.1868 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.1942
39617.1875 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1248
39617.1882 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1206
39617.1889 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1121
39617.1896 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0678
39617.1903 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9619.24 | 13.1942
39617.191 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9624.12 | 13.07
39617.1917 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1627
39617.1924 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1627
39617.1931 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0763
39617.1937 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.2089
39617.1944 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1142
39617.1951 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0362
39617.1958 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9599.72 | 13.1206
39617.1965 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1921
39617.1972 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0889
39617.1979 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.2005
39617.1986 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0868
39617.1993 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9594.84 | 13.1185
39617.2 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.2068
39617.2007 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0129
39617.2014 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.0383
39617.2021 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.1269
39617.2028 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1079
39617.2035 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0383
39617.2042 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.1542
39617.2049 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0742
39617.2056 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0171
39617.2063 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9604.6 | 13.1795
39617.2069 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.1942
39617.2076 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9604.6 | 13.0235
39617.2083 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.0404
39617.209 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9604.6 | 13.1627
39617.2097 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.2047
39617.2104 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.0404
39617.2111 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9589.96 | 13.1416
39617.2118 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1521
39617.2125 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.2131
39617.2132 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9604.6 | 13.1542
39617.2139 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9614.36 | 13.1921
39617.2146 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1879
39617.2153 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1627
39617.216 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0657
39617.2167 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.2005
39617.2174 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0868
39617.2181 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.234
39617.2188 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.0594
39617.2194 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.0678
39617.2201 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1437
39617.2208 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9624.12 | 13.0615
39617.2215 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.1332
39617.2222 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.1837
39617.2229 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0509
39617.2236 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.1058
39617.2243 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.2298
39617.225 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.0446
39617.2257 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0509
39617.2264 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.0594
39617.2271 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1774
39617.2278 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.169
39617.2285 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0678
39617.2292 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.0657
39617.2299 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.19
39617.2306 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0002
39617.2312 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1837
39617.2319 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1627
39617.2326 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0657
39617.2333 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.0615
39617.234 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1858
39617.2347 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1542
39617.2354 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.1227
39617.2361 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.1837
39617.2368 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1627
39617.2375 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9604.6 | 13.0446
39617.2382 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.1711
39617.2389 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.1858
39617.2396 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.0805
39617.2403 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9604.6 | 13.2047
39617.241 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.1837
39617.2417 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0805
39617.2424 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0868
39617.2431 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1437
39617.2438 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0932
39617.2444 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1795
39617.2451 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1206
39617.2458 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0509
39617.2465 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.255
39617.2472 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9624.12 | 13.0277
39617.2479 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9614.36 | 13.2068
39617.2486 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.0129
39617.2493 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.1563
39617.25 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0995
39617.2507 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.0932
39617.2514 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.211
39617.2521 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.0488
39617.2528 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.1921
39617.2535 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0868
39617.2542 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0615
39617.2549 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9599.72 | 13.1711
39617.2556 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1142
39617.2562 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.07
39617.2569 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.1627
39617.2576 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.19
39617.2583 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0763
39617.259 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.2235
39617.2597 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.1269
39617.2604 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.07
39617.2611 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9609.48 | 13.1563
39617.2618 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.2214
39617.2625 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9609.48 | 13.0953
39617.2632 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0256
39617.2639 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.1627
39617.2646 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.1374
39617.2653 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.0636
39617.266 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9599.72 | 13.1837
39617.2667 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.2319
39617.2674 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0911
39617.2681 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.129
39617.2688 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.0932
39617.2694 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1332
39617.2701 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1921
39617.2708 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0404
39617.2715 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.19
39617.2722 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9629 | 13.0425
39617.2729 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.1058
39617.2736 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1942
39617.2743 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0911
39617.275 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0261 | 9614.36 | 13.19
39617.2757 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0235
39617.2764 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.1984
39617.2771 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0784
39617.2778 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1479
39617.2785 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.2319
39617.2792 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9614.36 | 13.0362
39617.2799 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9599.72 | 13.0383
39617.2806 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.2005
39617.2812 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0868
39617.2819 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0826
39617.2826 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.1753
39617.2833 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9619.24 | 13.2131
39617.284 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9619.24 | 13.0932
39617.2847 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.034
39617.2854 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0889
39617.2861 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.2131
39617.2868 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9614.36 | 13.0235
39617.2875 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9594.84 | 13.1248
39617.2882 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.2214
39617.2889 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.1648
39617.2896 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0847
39617.2903 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9599.72 | 13.0467
39617.291 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1374
39617.2917 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.2131
39617.2924 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.1079
39617.2931 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0383 | 9609.48 | 13.0953
39617.2937 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1774
39617.2944 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.0826
39617.2951 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.1921
39617.2958 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0678
39617.2965 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.2047
39617.2972 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1374
39617.2979 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.0615
39617.2986 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.1521
39617.2993 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.1858
39617.3 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0826
39617.3007 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.11
39617.3014 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.1542
39617.3021 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.1879
39617.3028 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0256
39617.3035 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.2026
39617.3042 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9619.24 | 13.015
39617.3049 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.2068
39617.3056 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0192
39617.3063 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9594.84 | 13.169
39617.3069 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.2319
39617.3076 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9609.48 | 13.0108
39617.3083 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.1458
39617.309 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.1942
39617.3097 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.0826
39617.3104 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0871 | 9599.72 | 13.1311
39617.3111 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1669
39617.3118 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.0087
39617.3125 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.0362
39617.3132 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9599.72 | 13.19
39617.3139 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0742
39617.3146 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1837
39617.3153 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.0847
39617.316 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 12.9901
39617.3167 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0889
39617.3174 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0749 | 9599.72 | 13.0742
39617.3181 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0505 | 9604.6 | 13.2089
39617.3187 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0594
39617.3194 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0261 | 9589.96 | 13.0615
39617.3201 | 2008-06-18 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.2172
39617.3208 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.11
39617.3215 | 39617 | -------- | 0 | 1.9672 | 11.0383 | 9599.72 | 13.0678
39617.3222 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1795
39617.3229 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.034
39617.3236 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1605
39617.3243 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.2424
39617.325 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0319
39617.3257 | 39617 | -------- | 0 | 1.9672 | 11.0261 | 9599.72 | 13.0763
39617.3264 | 39617 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.2193
39617.3271 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.1753
39617.3278 | 39617 | -------- | 0 | 1.9672 | 11.0383 | 9604.6 | 13.0044
39617.3285 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9589.96 | 13.11
39617.3292 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.1521
39617.3299 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9594.84 | 13.2152
39617.3306 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9609.48 | 13.0362
39617.3313 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.0847
39617.3319 | 39617 | -------- | 0 | 1.9672 | 11.0749 | 9614.36 | 13.1605
39617.3326 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9609.48 | 13.0404
39617.3333 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.0911
39617.334 | 39617 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.211
39617.3347 | 39617 | -------- | 0 | 1.9672 | 11.0749 | 9594.84 | 13.1858
39617.3354 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.1942
39617.3361 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.0108
39617.3368 | 39617 | -------- | 0 | 1.9672 | 11.0261 | 9585.08 | 13.0636
39617.3375 | 39617 | -------- | 0 | 1.9672 | 11.0383 | 9594.84 | 13.0425
39617.3382 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9589.96 | 13.0213
39617.3389 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9589.96 | 13.1669
39617.3396 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.2005
39617.3403 | 39617 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.1437
39617.341 | 39617 | -------- | 0 | 1.9672 | 11.0871 | 9599.72 | 13.0277
39617.3417 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9614.36 | 13.0509
39617.3424 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.2235
39617.3431 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9594.84 | 13.211
39617.3438 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9604.6 | 13.1037
39617.3444 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.1142
39617.3451 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9614.36 | 13.1774
39617.3458 | 39617 | -------- | 0 | 1.9672 | 11.0749 | 9599.72 | 13.0573
39617.3465 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9599.72 | 13.1795
39617.3472 | 39617 | -------- | 0 | 1.9672 | 11.0749 | 9594.84 | 13.1563
39617.3479 | 39617 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.0847
39617.3486 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9585.08 | 13.1837
39617.3493 | 39617 | -------- | 0 | 1.9672 | 11.0749 | 9604.6 | 13.1437
39617.35 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9594.84 | 13.0805
39617.3507 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9594.84 | 13.0721
39617.3514 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9589.96 | 13.1627
39617.3521 | 39617 | -------- | 0 | 1.9672 | 11.0505 | 9589.96 | 13.1458
39617.3528 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9594.84 | 13.169
39617.3535 | 39617 | -------- | 0 | 1.9672 | 11.0627 | 9599.72 | 13.2005

## Ref_Raw
;; | Model | 81I- | (prog | iSeries | 81i | 01.00.01.079)
;;
Time | Date | Flags | conc | syssp | hgflo | dlflo | ctemp
39616.3639 | 2008-06-17 | ----LL-- | 0 | 1.9672 | 0 | 0 | 13.0298
39616.3646 | 2008-06-17 | ----L?-- | 0 | 1.9672 | 0 | 0 | 13.015
39616.3653 | 2008-06-17 | ----LL-- | 0 | 1.9672 | 0 | 0 | 12.9276
39616.366 | 2008-06-17 | ----LL-- | 0 | 1.9672 | 0 | 0 | 12.9134
39616.3667 | 2008-06-17 | ----LL-- | 0 | 1.9672 | 0 | 0 | 12.8243
39616.3674 | 2008-06-17 | ----LL-- | 0 | 1.9672 | 0 | 0 | 12.7817
39616.3681 | 2008-06-17 | ----LL-- | 0 | 1.9672 | 0 | 0 | 12.7226
39616.3688 | 2008-06-17 | ----LL-- | 0 | 1.9672 | 0 | 0 | 12.8142
39616.3694 | 2008-06-17 | ----LL-- | 0 | 1.9672 | 0 | 0 | 12.9841
39616.3701 | 2008-06-17 | ----LL-- | 0 | 1.9672 | 0 | 0 | 13.0847
39616.3708 | 2008-06-17 | ?--H??-- | 5.1838 | 1.9672 | 10.9136 | 9645.85 | 13.0383
39616.3715 | 2008-06-17 | -------- | 8.1154 | 1.9672 | 10.9284 | 9656.58 | 13.0277
39616.3722 | 2008-06-17 | -------- | 8.1109 | 1.9672 | 10.8988 | 9656.58 | 13.0383
39616.3729 | 2008-06-17 | -------- | 8.1112 | 1.9672 | 10.9284 | 9645.85 | 12.9195
39616.3736 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.9136 | 9645.85 | 12.8223
39616.3743 | 2008-06-17 | -------- | 8.1142 | 1.9672 | 10.9136 | 9651.21 | 12.7613
39616.375 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.9284 | 9651.21 | 12.7736
39616.3757 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.9431 | 9651.21 | 12.7491
39616.3764 | 2008-06-17 | -------- | 8.1127 | 1.9672 | 10.9136 | 9645.85 | 12.7206
39616.3771 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.8988 | 9635.12 | 12.8162
39616.3778 | 2008-06-17 | -------- | 8.1092 | 1.9672 | 10.9136 | 9629.75 | 13.0889
39616.3785 | 2008-06-17 | -------- | 8.1112 | 1.9672 | 10.9136 | 9629.75 | 13.0425
39616.3792 | 2008-06-17 | -------- | 8.1103 | 1.9672 | 10.9284 | 9629.75 | 12.8203
39616.3799 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.9136 | 9635.12 | 12.7471
39616.3847 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.9284 | 9640.48 | 12.8608
39616.3854 | 2008-06-17 | -------- | 8.115 | 1.9672 | 10.9284 | 9645.85 | 12.7959
39616.3861 | 2008-06-17 | -------- | 8.1147 | 1.9672 | 10.9431 | 9640.48 | 12.7593
39616.3868 | 2008-06-17 | -------- | 8.1157 | 1.9672 | 10.9136 | 9645.85 | 12.7328
39616.3875 | 2008-06-17 | -------- | 8.1141 | 1.9672 | 10.9136 | 9635.12 | 12.8
39616.3882 | 2008-06-17 | -------- | 8.1132 | 1.9672 | 10.9136 | 9629.75 | 13.0129
39616.3889 | 2008-06-17 | -------- | 8.1111 | 1.9672 | 10.9136 | 9613.65 | 13.0298
39616.3896 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.9431 | 9629.75 | 12.9357
39616.3903 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.9136 | 9635.12 | 12.743
39616.391 | 2008-06-17 | -------- | 8.1126 | 1.9672 | 10.9136 | 9640.48 | 12.7451
39616.3917 | 2008-06-17 | -------- | 8.1117 | 1.9672 | 10.9284 | 9640.48 | 12.7084
39616.3924 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.9136 | 9635.12 | 12.7451
39616.3931 | 2008-06-17 | -------- | 8.1141 | 1.9672 | 10.9284 | 9635.12 | 13.0213
39616.3938 | 2008-06-17 | -------- | 8.1126 | 1.9672 | 10.9284 | 9635.12 | 13.1121
39616.3944 | 2008-06-17 | -------- | 8.1139 | 1.9672 | 10.9284 | 9645.85 | 13.0678
39616.3951 | 2008-06-17 | -------- | 8.1126 | 1.9672 | 10.9284 | 9640.48 | 13.07
39616.3958 | 2008-06-17 | -------- | 8.112 | 1.9672 | 10.9136 | 9640.48 | 13.0953
39616.3965 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.9284 | 9640.48 | 13.0826
39616.3972 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.9284 | 9629.75 | 13.0721
39616.3979 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.9431 | 9624.39 | 13.0002
39616.3986 | 2008-06-17 | -------- | 8.1114 | 1.9672 | 10.9431 | 9640.48 | 12.8548
39616.3993 | 2008-06-17 | -------- | 8.1135 | 1.9672 | 10.9284 | 9635.12 | 12.7634
39616.4 | 2008-06-17 | -------- | 8.1135 | 1.9672 | 10.8988 | 9651.21 | 12.7369
39616.4007 | 2008-06-17 | -------- | 8.1131 | 1.9672 | 10.9136 | 9640.48 | 12.7654
39616.4014 | 2008-06-17 | -------- | 8.1151 | 1.9672 | 10.9284 | 9651.21 | 12.7736
39616.4021 | 2008-06-17 | -------- | 8.1141 | 1.9672 | 10.9284 | 9645.85 | 12.7471
39616.4028 | 2008-06-17 | -------- | 8.1146 | 1.9672 | 10.9136 | 9651.21 | 12.7878
39616.4035 | 2008-06-17 | -------- | 8.1144 | 1.9672 | 10.8988 | 9651.21 | 12.7512
39616.4042 | 2008-06-17 | -------- | 8.1133 | 1.9672 | 10.9136 | 9645.85 | 12.7349
39616.4049 | 2008-06-17 | -------- | 8.1131 | 1.9672 | 10.9136 | 9629.75 | 12.8426
39616.4056 | 2008-06-17 | -------- | 8.1129 | 1.9672 | 10.9284 | 9635.12 | 12.7858
39616.4062 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.8988 | 9629.75 | 12.8872
39616.4069 | 2008-06-17 | -------- | 8.1098 | 1.9672 | 10.9136 | 9629.75 | 13.1227
39616.4076 | 2008-06-17 | -------- | 8.1102 | 1.9672 | 10.9431 | 9613.65 | 12.9861
39616.4083 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.9136 | 9629.75 | 12.9397
39616.409 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.9136 | 9635.12 | 12.8507
39616.4097 | 2008-06-17 | -------- | 8.1136 | 1.9672 | 10.9284 | 9635.12 | 12.7979
39616.4104 | 2008-06-17 | -------- | 8.1132 | 1.9672 | 10.9284 | 9640.48 | 12.8426
39616.4111 | 2008-06-17 | -------- | 8.1135 | 1.9672 | 10.9284 | 9645.85 | 12.8284
39616.4118 | 2008-06-17 | -------- | 8.114 | 1.9672 | 10.9136 | 9635.12 | 12.8142
39616.4125 | 2008-06-17 | -------- | 8.1155 | 1.9672 | 10.9136 | 9645.85 | 12.8932
39616.4132 | 2008-06-17 | -------- | 8.1141 | 1.9672 | 10.9284 | 9635.12 | 12.8527
39616.4139 | 2008-06-17 | -------- | 8.1149 | 1.9672 | 10.9431 | 9640.48 | 12.8527
39616.4153 | 2008-06-17 | L-L-?L-- | 0.0032925 | 1.9608 | 10.9136 | 20017.3 | 11.9672
39616.416 | 2008-06-17 | ?---?L-- | 12.1472 | 1.9672 | 10.9136 | 9463.28 | 12.9841
39616.4167 | 2008-06-17 | -------- | 8.1194 | 1.9672 | 10.8988 | 9613.65 | 13.0171
39616.4174 | 2008-06-17 | -------- | 8.1138 | 1.9672 | 10.9136 | 9629.75 | 12.8345
39616.4181 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.9284 | 9635.12 | 12.7226
39616.4187 | 2008-06-17 | -------- | 8.1147 | 1.9672 | 10.9284 | 9624.39 | 12.8183
39616.4194 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.9284 | 9624.39 | 13.0425
39616.4201 | 2008-06-17 | -------- | 8.114 | 1.9672 | 10.9136 | 9640.48 | 13.1227
39616.4208 | 2008-06-17 | -------- | 8.1131 | 1.9672 | 10.9431 | 9640.48 | 13.1142
39616.4215 | 2008-06-17 | -------- | 8.1137 | 1.9672 | 10.9431 | 9640.48 | 13.0573
39616.4222 | 2008-06-17 | -------- | 8.1135 | 1.9672 | 10.8988 | 9651.21 | 12.9276
39616.4229 | 2008-06-17 | -------- | 8.1124 | 1.9672 | 10.9284 | 9645.85 | 12.9538
39616.4236 | 2008-06-17 | -------- | 8.1135 | 1.9672 | 10.9136 | 9640.48 | 12.8203
39616.4243 | 2008-06-17 | -------- | 8.1135 | 1.9672 | 10.9284 | 9645.85 | 12.8284
39616.425 | 2008-06-17 | -------- | 8.1128 | 1.9672 | 10.9136 | 9640.48 | 12.7512
39616.4257 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.9284 | 9640.48 | 12.739
39616.4264 | 2008-06-17 | -------- | 8.1098 | 1.9672 | 10.8841 | 9624.39 | 13.015
39616.4271 | 2008-06-17 | -------- | 8.1131 | 1.9672 | 10.8988 | 9624.39 | 13.1332
39616.4278 | 2008-06-17 | -------- | 8.112 | 1.9672 | 10.9136 | 9635.12 | 13.0383
39616.4285 | 2008-06-17 | -------- | 8.1122 | 1.9672 | 10.8988 | 9624.39 | 12.8892
39616.4292 | 2008-06-17 | -------- | 8.1139 | 1.9672 | 10.9284 | 9640.48 | 12.9175
39616.4299 | 2008-06-17 | -------- | 8.1139 | 1.9672 | 10.9284 | 9645.85 | 12.9175
39616.4306 | 2008-06-17 | -------- | 8.1129 | 1.9672 | 10.9136 | 9645.85 | 12.8284
39616.4313 | 2008-06-17 | -------- | 8.1153 | 1.9672 | 10.9284 | 9640.48 | 12.8203
39616.4319 | 2008-06-17 | -------- | 8.1143 | 1.9672 | 10.9284 | 9645.85 | 12.9013
39616.4326 | 2008-06-17 | -------- | 8.1152 | 1.9672 | 10.9136 | 9645.85 | 12.9114
39616.4333 | 2008-06-17 | -------- | 8.1139 | 1.9672 | 10.8841 | 9640.48 | 12.8791
39616.434 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.8988 | 9651.21 | 12.7817
39616.4347 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.9284 | 9635.12 | 12.7837
39616.4354 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.9284 | 9624.39 | 12.7247
39616.4361 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.9284 | 9619.02 | 13.0213
39616.4368 | 2008-06-17 | -------- | 8.1114 | 1.9672 | 10.9431 | 9608.29 | 12.9659
39616.4375 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.9136 | 9635.12 | 12.7776
39616.4382 | 2008-06-17 | -------- | 8.1112 | 1.9672 | 10.9284 | 9635.12 | 12.7634
39616.4389 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.9136 | 9645.85 | 12.7491
39616.4396 | 2008-06-17 | -------- | 8.1105 | 1.9672 | 10.9284 | 9608.29 | 13.1079
39616.4403 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.9284 | 9629.75 | 12.9518
39616.441 | 2008-06-17 | -------- | 8.1137 | 1.9672 | 10.9136 | 9645.85 | 12.974
39616.4417 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.9136 | 9645.85 | 12.9235
39616.4424 | 2008-06-17 | -------- | 8.1128 | 1.9672 | 10.8988 | 9645.85 | 12.8527
39616.4431 | 2008-06-17 | -------- | 8.1132 | 1.9672 | 10.9284 | 9645.85 | 12.7675
39616.4437 | 2008-06-17 | -------- | 8.1141 | 1.9672 | 10.9284 | 9645.85 | 12.7613
39616.4444 | 2008-06-17 | -------- | 8.1148 | 1.9672 | 10.9284 | 9645.85 | 12.7308
39616.4451 | 2008-06-17 | -------- | 8.1132 | 1.9672 | 10.9136 | 9645.85 | 12.7573
39616.4458 | 2008-06-17 | -------- | 8.1136 | 1.9672 | 10.9284 | 9635.12 | 12.7736
39616.4465 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.9284 | 9624.39 | 12.9114
39616.4472 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.8988 | 9629.75 | 13.0446
39616.4479 | 2008-06-17 | -------- | 8.1103 | 1.9672 | 10.9579 | 9619.02 | 13.0678
39616.4486 | 2008-06-17 | -------- | 8.1122 | 1.9672 | 10.9284 | 9619.02 | 13.0256
39616.4493 | 2008-06-17 | -------- | 8.1122 | 1.9672 | 10.9136 | 9619.02 | 12.804
39616.45 | 2008-06-17 | -------- | 8.1122 | 1.9672 | 10.9284 | 9635.12 | 12.7451
39616.4507 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.9284 | 9613.65 | 13.0932
39616.4514 | 2008-06-17 | -------- | 8.1136 | 1.9672 | 10.9431 | 9629.75 | 13.0531
39616.4521 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.9284 | 9635.12 | 13.015
39616.4528 | 2008-06-17 | -------- | 8.1142 | 1.9672 | 10.9284 | 9635.12 | 12.9155
39616.4535 | 2008-06-17 | -------- | 8.1132 | 1.9672 | 10.9136 | 9635.12 | 12.8406
39616.4542 | 2008-06-17 | -------- | 8.1129 | 1.9672 | 10.9136 | 9645.85 | 12.7552
39616.4549 | 2008-06-17 | -------- | 8.1144 | 1.9672 | 10.9136 | 9635.12 | 12.7491
39616.4556 | 2008-06-17 | -------- | 8.1129 | 1.9672 | 10.9136 | 9645.85 | 12.7512
39616.4563 | 2008-06-17 | -------- | 8.1128 | 1.9672 | 10.8988 | 9640.48 | 12.7349
39616.4569 | 2008-06-17 | -------- | 8.1126 | 1.9672 | 10.9136 | 9640.48 | 12.7532
39616.4576 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.9136 | 9635.12 | 12.7552
39616.4583 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.8988 | 9629.75 | 12.9821
39616.459 | 2008-06-17 | -------- | 8.1106 | 1.9672 | 10.9284 | 9624.39 | 13.0594
39616.4597 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.8988 | 9624.39 | 13.0425
39616.4604 | 2008-06-17 | -------- | 8.1106 | 1.9672 | 10.8841 | 9629.75 | 12.8446
39616.4611 | 2008-06-17 | -------- | 8.1103 | 1.9672 | 10.9136 | 9629.75 | 12.7084
39616.4618 | 2008-06-17 | -------- | 8.1112 | 1.9672 | 10.9431 | 9619.02 | 12.9377
39616.4625 | 2008-06-17 | -------- | 8.112 | 1.9672 | 10.9136 | 9629.75 | 13.0467
39616.4632 | 2008-06-17 | -------- | 8.1131 | 1.9672 | 10.8988 | 9629.75 | 13.0742
39616.4639 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.9284 | 9640.48 | 12.8831
39616.4646 | 2008-06-17 | -------- | 8.1148 | 1.9672 | 10.9136 | 9651.21 | 12.7613
39616.4653 | 2008-06-17 | -------- | 8.1136 | 1.9672 | 10.9136 | 9640.48 | 12.7979
39616.466 | 2008-06-17 | -------- | 8.1142 | 1.9672 | 10.9284 | 9640.48 | 12.7736
39616.4667 | 2008-06-17 | -------- | 8.1128 | 1.9672 | 10.9136 | 9635.12 | 12.7736
39616.4674 | 2008-06-17 | -------- | 8.1124 | 1.9672 | 10.9284 | 9645.85 | 12.8851
39616.4681 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.9136 | 9629.75 | 13.0087
39616.4688 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.9136 | 9656.58 | 12.8629
39616.4694 | 2008-06-17 | -------- | 8.1097 | 1.9672 | 10.9284 | 9624.39 | 13.0531
39616.4701 | 2008-06-17 | -------- | 8.1117 | 1.9672 | 10.8988 | 9624.39 | 13.0678
39616.4708 | 2008-06-17 | -------- | 8.1098 | 1.9672 | 10.8988 | 9629.75 | 12.8365
39616.4715 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.9136 | 9645.85 | 12.69
39616.4722 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.9136 | 9624.39 | 12.9134
39616.4729 | 2008-06-17 | -------- | 8.1112 | 1.9672 | 10.9136 | 9629.75 | 13.1058
39616.4736 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.9284 | 9629.75 | 13.0235
39616.4743 | 2008-06-17 | -------- | 8.1136 | 1.9672 | 10.8841 | 9635.12 | 12.974
39616.475 | 2008-06-17 | -------- | 8.1127 | 1.9672 | 10.9136 | 9635.12 | 12.7959
39616.4757 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.9431 | 9656.58 | 12.7634
39616.4764 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.9284 | 9635.12 | 12.8426
39616.4771 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.9284 | 9640.48 | 12.877
39616.4778 | 2008-06-17 | -------- | 8.1126 | 1.9672 | 10.9136 | 9635.12 | 13.0044
39616.4785 | 2008-06-17 | -------- | 8.1124 | 1.9672 | 10.9136 | 9645.85 | 13.0594
39616.4792 | 2008-06-17 | -------- | 8.1125 | 1.9672 | 10.8988 | 9635.12 | 13.0446
39616.4799 | 2008-06-17 | -------- | 8.114 | 1.9672 | 10.9284 | 9640.48 | 13.0467
39616.4806 | 2008-06-17 | -------- | 8.1108 | 1.9672 | 10.9284 | 9635.12 | 12.8467
39616.4812 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.9136 | 9635.12 | 12.7349
39616.4819 | 2008-06-17 | -------- | 8.1111 | 1.9672 | 10.8988 | 9629.75 | 12.873
39616.4826 | 2008-06-17 | -------- | 8.11 | 1.9672 | 10.9136 | 9613.65 | 13.0002
39616.4833 | 2008-06-17 | -------- | 8.1128 | 1.9672 | 10.8988 | 9629.75 | 12.8385
39616.484 | 2008-06-17 | -------- | 8.1106 | 1.9672 | 10.9136 | 9640.48 | 12.7797
39616.4847 | 2008-06-17 | -------- | 8.1128 | 1.9672 | 10.9136 | 9629.75 | 12.974
39616.4854 | 2008-06-17 | -------- | 8.1132 | 1.9672 | 10.9136 | 9629.75 | 13.0763
39616.4861 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.9136 | 9635.12 | 13.0763
39616.4868 | 2008-06-17 | -------- | 8.1152 | 1.9672 | 10.9136 | 9651.21 | 13.0826
39616.4875 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.8988 | 9640.48 | 13.1206
39616.4882 | 2008-06-17 | -------- | 8.114 | 1.9672 | 10.9284 | 9645.85 | 13.1016
39616.4889 | 2008-06-17 | -------- | 8.1126 | 1.9672 | 10.9136 | 9640.48 | 13.1311
39616.4896 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.9284 | 9640.48 | 12.9195
39616.4903 | 2008-06-17 | -------- | 8.1122 | 1.9672 | 10.8988 | 9645.85 | 12.7776
39616.491 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.9431 | 9640.48 | 12.8325
39616.4917 | 2008-06-17 | -------- | 8.1108 | 1.9672 | 10.9431 | 9629.75 | 12.9881
39616.4924 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.9284 | 9624.39 | 13.0319
39616.4931 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.9136 | 9619.02 | 12.8203
39616.4938 | 2008-06-17 | -------- | 8.1133 | 1.9672 | 10.8988 | 9635.12 | 12.7349
39616.4944 | 2008-06-17 | -------- | 8.1098 | 1.9672 | 10.8841 | 9602.92 | 13.0235
39616.4951 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.9284 | 9619.02 | 12.9377
39616.4958 | 2008-06-17 | -------- | 8.1112 | 1.9672 | 10.9136 | 9629.75 | 12.7634
39616.4965 | 2008-06-17 | -------- | 8.1117 | 1.9672 | 10.9136 | 9619.02 | 13.0319
39616.4972 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.9136 | 9629.75 | 13.0362
39616.4979 | 2008-06-17 | -------- | 8.1118 | 1.9672 | 10.8988 | 9635.12 | 12.9357
39616.4986 | 2008-06-17 | -------- | 8.1142 | 1.9672 | 10.8988 | 9635.12 | 12.8345
39616.4993 | 2008-06-17 | -------- | 8.1133 | 1.9672 | 10.9284 | 9640.48 | 12.7104
39616.5 | 2008-06-17 | -------- | 8.1127 | 1.9672 | 10.9284 | 9635.12 | 12.7613
39616.5007 | 2008-06-17 | -------- | 8.1142 | 1.9672 | 10.8841 | 9640.48 | 12.7349
39616.5014 | 2008-06-17 | -------- | 8.1124 | 1.9672 | 10.8988 | 9640.48 | 12.7736
39616.5021 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.8988 | 9635.12 | 12.8629
39616.5028 | 2008-06-17 | -------- | 8.1127 | 1.9672 | 10.9284 | 9629.75 | 12.978
39616.5035 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.9136 | 9629.75 | 13.0784
39616.5042 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.9284 | 9624.39 | 12.8791
39616.5049 | 2008-06-17 | -------- | 8.1104 | 1.9672 | 10.9136 | 9635.12 | 12.7104
39616.5056 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.9284 | 9635.12 | 12.7878
39616.5062 | 2008-06-17 | -------- | 8.11 | 1.9672 | 10.9136 | 9608.29 | 13.0171
39616.5069 | 2008-06-17 | -------- | 8.1109 | 1.9672 | 10.8841 | 9624.39 | 12.8689
39616.5076 | 2008-06-17 | -------- | 8.1106 | 1.9672 | 10.8988 | 9635.12 | 12.741
39616.5083 | 2008-06-17 | -------- | 8.1111 | 1.9672 | 10.9284 | 9624.39 | 12.9901
39616.509 | 2008-06-17 | -------- | 8.1089 | 1.9672 | 10.9136 | 9619.02 | 13.0383
39616.5097 | 2008-06-17 | -------- | 8.1125 | 1.9672 | 10.9136 | 9629.75 | 12.8912
39616.5104 | 2008-06-17 | -------- | 8.1116 | 1.9672 | 10.9284 | 9645.85 | 12.7878
39616.5111 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.9136 | 9640.48 | 12.7104
39616.5125 | 2008-06-17 | --L-LL-- | 0 | 1.8571 | 10.9874 | 13743.2 | -99
39616.5132 | 2008-06-17 | ?-L-?L-- | 19.4579 | 1.9672 | 0 | 317.936 | 12.6696
39616.5139 | 2008-06-17 | L---LL-- | 9.3613 | 1.9672 | 10.9136 | 9619.02 | 13.0256
39616.5146 | 2008-06-17 | -------- | 8.1116 | 1.9672 | 10.8988 | 9613.65 | 13.0108
39616.5153 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.9284 | 9645.85 | 12.8284
39616.516 | 2008-06-17 | -------- | 8.1129 | 1.9672 | 10.9136 | 9645.85 | 12.7043
39616.5167 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.9136 | 9619.02 | 13.0573
39616.5174 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.9136 | 9629.75 | 13.0319
39616.5181 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.9284 | 9629.75 | 12.9639
39616.5188 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.9284 | 9629.75 | 12.8345
39616.5194 | 2008-06-17 | -------- | 8.1114 | 1.9672 | 10.9284 | 9629.75 | 12.804
39616.5201 | 2008-06-17 | -------- | 8.1111 | 1.9672 | 10.9136 | 9629.75 | 12.9336
39616.5208 | 2008-06-17 | -------- | 8.1112 | 1.9672 | 10.9284 | 9613.65 | 13.0087
39616.5215 | 2008-06-17 | -------- | 8.112 | 1.9672 | 10.8988 | 9629.75 | 12.9316
39616.5222 | 2008-06-17 | -------- | 8.1101 | 1.9672 | 10.9431 | 9629.75 | 12.7797
39616.5229 | 2008-06-17 | -------- | 8.1099 | 1.9672 | 10.9136 | 9629.75 | 12.7491
39616.5236 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.8988 | 9624.39 | 12.9518
39616.5243 | 2008-06-17 | -------- | 8.1115 | 1.9672 | 10.9431 | 9629.75 | 13.0995
39616.525 | 2008-06-17 | -------- | 8.1116 | 1.9672 | 10.9284 | 9629.75 | 13.0129
39616.5257 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.8988 | 9635.12 | 12.9114
39616.5264 | 2008-06-17 | -------- | 8.1131 | 1.9672 | 10.9284 | 9640.48 | 12.8467
39616.5271 | 2008-06-17 | -------- | 8.1131 | 1.9672 | 10.9136 | 9645.85 | 12.7451
39616.5278 | 2008-06-17 | -------- | 8.1154 | 1.9672 | 10.9284 | 9645.85 | 12.7593
39616.5285 | 2008-06-17 | -------- | 8.1146 | 1.9672 | 10.9136 | 9645.85 | 12.8385
39616.5292 | 2008-06-17 | -------- | 8.1131 | 1.9672 | 10.8841 | 9640.48 | 12.871
39616.5299 | 2008-06-17 | -------- | 8.1134 | 1.9672 | 10.9284 | 9651.21 | 12.8101
39616.5306 | 2008-06-17 | -------- | 8.1112 | 1.9672 | 10.8988 | 9624.39 | 13.0594
39616.5312 | 2008-06-17 | -------- | 8.114 | 1.9672 | 10.9284 | 9645.85 | 13.0531
39616.5319 | 2008-06-17 | -------- | 8.1119 | 1.9672 | 10.9284 | 9645.85 | 13.0826
39616.5326 | 2008-06-17 | -------- | 8.1111 | 1.9672 | 10.9136 | 9629.75 | 13.0467
39616.5333 | 2008-06-17 | -------- | 8.1105 | 1.9672 | 10.9136 | 9624.39 | 12.8851
39616.534 | 2008-06-17 | -------- | 8.1123 | 1.9672 | 10.9284 | 9635.12 | 12.7797
39616.5347 | 2008-06-17 | -------- | 8.1125 | 1.9672 | 10.9284 | 9640.48 | 12.6982
39616.5354 | 2008-06-17 | -------- | 8.1084 | 1.9672 | 10.9284 | 9619.02 | 13.07
39616.5361 | 2008-06-17 | -------- | 8.1105 | 1.9672 | 10.9136 | 9613.65 | 12.8892
39616.5368 | 2008-06-17 | -------- | 8.113 | 1.9672 | 10.9136 | 9635.12 | 12.741
39616.5375 | 2008-06-17 | -------- | 8.1108 | 1.9672 | 10.9284 | 9635.12 | 12.7756
39616.5382 | 2008-06-17 | -------- | 8.1105 | 1.9672 | 10.8988 | 9624.39 | 12.9377
39616.5389 | 2008-06-17 | -------- | 8.1112 | 1.9672 | 10.9136 | 9635.12 | 13.0235
39616.5396 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.9136 | 9635.12 | 13.1185
39616.5403 | 2008-06-17 | -------- | 8.1125 | 1.9672 | 10.8988 | 9629.75 | 13.1206
39616.541 | 2008-06-17 | -------- | 8.1112 | 1.9672 | 10.8988 | 9629.75 | 12.9458
39616.5417 | 2008-06-17 | -------- | 8.1111 | 1.9672 | 10.8988 | 9640.48 | 12.8548
39616.5424 | 2008-06-17 | -------- | 8.1114 | 1.9672 | 10.9431 | 9645.85 | 12.7532
39616.5431 | 2008-06-17 | -------- | 8.1093 | 1.9672 | 10.8988 | 9635.12 | 12.7634
39616.5437 | 2008-06-17 | -------- | 8.1113 | 1.9672 | 10.9136 | 9629.75 | 12.9841
39616.5444 | 2008-06-17 | -------- | 8.1092 | 1.9672 | 10.8841 | 9613.65 | 13.0213
39616.5451 | 2008-06-17 | -------- | 8.1095 | 1.9672 | 10.9136 | 9629.75 | 12.8689
39616.5458 | 2008-06-17 | -------- | 8.111 | 1.9672 | 10.9284 | 9635.12 | 12.8426
39616.5465 | 2008-06-17 | -------- | 8.1121 | 1.9672 | 10.9136 | 9640.48 | 12.7104
39616.5472 | 2008-06-17 | -------- | 8.1114 | 1.9672 | 10.8988 | 9635.12 | 12.7328
39616.5479 | 2008-06-17 | -------- | 8.1096 | 1.9672 | 10.8841 | 9629.75 | 12.9114
39616.5486 | 2008-06-17 | -------- | 8.1106 | 1.9672 | 10.9284 | 9619.02 | 13.0298
39616.5493 | 2008-06-17 | -------- | 8.1114 | 1.9672 | 10.9136 | 9635.12 | 13.0657
39616.55 | 2008-06-17 | -------- | 8.11 | 1.9672 | 10.8841 | 9635.12 | 12.9921
39616.5507 | 2008-06-17 | -------- | 8.1122 | 1.9672 | 10.9284 | 9645.85 | 12.7715
39616.5514 | 2008-06-17 | L----L-- | 7.9515 | 1.9672 | 3.8035 | 10872.2 | 12.7308
39616.5521 | 2008-06-17 | -------- | 5.6768 | 1.9672 | 8.7455 | 10760.9 | 12.8527
39616.5528 | 2008-06-17 | -------- | 5.7145 | 1.9672 | 8.7455 | 10760.9 | 12.9175
39616.5535 | 2008-06-17 | -------- | 5.7133 | 1.9672 | 8.7455 | 10766.2 | 12.8811
39616.5542 | 2008-06-17 | -------- | 5.7138 | 1.9672 | 8.7603 | 10766.2 | 13.0594
39616.5549 | 2008-06-17 | -------- | 5.7121 | 1.9672 | 8.7455 | 10755.6 | 13.0065
39616.5556 | 2008-06-17 | -------- | 5.7128 | 1.9672 | 8.7603 | 10766.2 | 12.873
39616.5563 | 2008-06-17 | -------- | 5.7122 | 1.9672 | 8.7603 | 10766.2 | 12.7919
39616.5569 | 2008-06-17 | -------- | 5.7117 | 1.9672 | 8.775 | 10755.6 | 12.7247
39616.5576 | 2008-06-17 | -------- | 5.7127 | 1.9672 | 8.7308 | 10750.2 | 12.8406
39616.5583 | 2008-06-17 | -------- | 5.7098 | 1.9672 | 8.7308 | 10744.9 | 13.0277
39616.559 | 2008-06-17 | -------- | 5.7097 | 1.9672 | 8.7455 | 10729 | 12.9538
39616.5597 | 2008-06-17 | -------- | 5.7126 | 1.9672 | 8.7603 | 10766.2 | 12.7593
39616.5604 | 2008-06-17 | -------- | 5.7111 | 1.9672 | 8.7603 | 10760.9 | 12.7695
39616.5611 | 2008-06-17 | -------- | 5.7108 | 1.9672 | 8.7455 | 10760.9 | 12.739
39616.5618 | 2008-06-17 | -------- | 5.713 | 1.9672 | 8.7897 | 10750.2 | 13.0044
39616.5625 | 2008-06-17 | -------- | 5.7135 | 1.9672 | 8.775 | 10755.6 | 13.1037
39616.5632 | 2008-06-17 | -------- | 5.7134 | 1.9672 | 8.775 | 10766.2 | 13.0552
39616.5639 | 2008-06-17 | -------- | 5.7127 | 1.9672 | 8.775 | 10760.9 | 12.9195
39616.5646 | 2008-06-17 | -------- | 5.713 | 1.9672 | 8.775 | 10766.2 | 12.9861
39616.5653 | 2008-06-17 | -------- | 5.7124 | 1.9672 | 8.7455 | 10760.9 | 12.7654
39616.566 | 2008-06-17 | -------- | 5.7136 | 1.9672 | 8.7455 | 10760.9 | 12.7593
39616.5667 | 2008-06-17 | -------- | 5.7136 | 1.9672 | 8.7455 | 10744.9 | 12.9558
39616.5674 | 2008-06-17 | -------- | 5.712 | 1.9672 | 8.7455 | 10755.6 | 12.9558
39616.5681 | 2008-06-17 | -------- | 5.7119 | 1.9672 | 8.7603 | 10750.2 | 13.0573
39616.5687 | 2008-06-17 | -------- | 5.7114 | 1.9672 | 8.775 | 10760.9 | 13.0911
39616.5694 | 2008-06-17 | -------- | 5.711 | 1.9672 | 8.7455 | 10739.6 | 12.978
39616.5701 | 2008-06-17 | -------- | 5.7116 | 1.9672 | 8.7897 | 10760.9 | 12.8811
39616.5708 | 2008-06-17 | -------- | 5.7111 | 1.9672 | 8.7603 | 10766.2 | 12.7328
39616.5715 | 2008-06-17 | -------- | 5.7099 | 1.9672 | 8.7455 | 10739.6 | 13.0446
39616.5722 | 2008-06-17 | -------- | 5.7102 | 1.9672 | 8.7603 | 10739.6 | 12.9013
39616.5729 | 2008-06-17 | -------- | 5.7102 | 1.9672 | 8.7603 | 10755.6 | 12.7858
39616.5736 | 2008-06-17 | -------- | 5.7101 | 1.9672 | 8.7603 | 10755.6 | 12.7654
39616.5743 | 2008-06-17 | -------- | 5.71 | 1.9672 | 8.7603 | 10729 | 13.0805
39616.575 | 2008-06-17 | -------- | 5.7113 | 1.9672 | 8.7455 | 10750.2 | 12.9579
39616.5757 | 2008-06-17 | -------- | 5.7119 | 1.9672 | 8.7603 | 10760.9 | 12.7491
39616.5764 | 2008-06-17 | -------- | 5.7124 | 1.9672 | 8.7455 | 10750.2 | 12.7573
39616.5771 | 2008-06-17 | -------- | 5.7132 | 1.9672 | 8.7603 | 10766.2 | 12.741
39616.5778 | 2008-06-17 | -------- | 5.7128 | 1.9672 | 8.7603 | 10766.2 | 12.7776
39616.5785 | 2008-06-17 | -------- | 5.7124 | 1.9672 | 8.7455 | 10766.2 | 12.7552
39616.5792 | 2008-06-17 | -------- | 5.7129 | 1.9672 | 8.775 | 10766.2 | 12.9316
39616.5799 | 2008-06-17 | -------- | 5.7109 | 1.9672 | 8.7603 | 10739.6 | 13.0784
39616.5806 | 2008-06-17 | -------- | 5.7123 | 1.9672 | 8.7897 | 10755.6 | 12.9961
39616.5813 | 2008-06-17 | -------- | 5.713 | 1.9672 | 8.7603 | 10760.9 | 12.7593
39616.5819 | 2008-06-17 | L----L-- | 3.2929 | 1.9672 | 6.0369 | 14500.9 | 12.7959
39616.5826 | 2008-06-17 | L----L-- | 2.6978 | 1.9672 | 6.0074 | 14511.4 | 12.7206
39616.5833 | 2008-06-17 | L----L-- | 2.6969 | 1.9672 | 6.0221 | 14485.1 | 13.0678
39616.584 | 2008-06-17 | L----L-- | 2.6968 | 1.9672 | 6.0221 | 14490.3 | 12.9195
39616.5847 | 2008-06-17 | L----L-- | 2.6961 | 1.9672 | 6.0369 | 14506.1 | 12.8162
39616.5854 | 2008-06-17 | L----L-- | 2.6964 | 1.9672 | 6.0369 | 14511.4 | 12.7797
39616.5861 | 2008-06-17 | L----L-- | 2.6959 | 1.9672 | 6.0369 | 14500.9 | 12.8507
39616.5868 | 2008-06-17 | L----L-- | 2.6957 | 1.9672 | 6.0221 | 14469.3 | 13.0065
39616.5875 | 2008-06-17 | L----L-- | 2.6951 | 1.9672 | 6.0369 | 14485.1 | 12.9054
39616.5882 | 2008-06-17 | L----L-- | 2.6961 | 1.9672 | 6.0221 | 14500.9 | 12.7125
39616.5889 | 2008-06-17 | L----L-- | 2.696 | 1.9672 | 6.0221 | 14474.5 | 12.9941
39616.5896 | 2008-06-17 | L----L-- | 2.6973 | 1.9672 | 6.0369 | 14490.3 | 13.0235
39616.5903 | 2008-06-17 | L----L-- | 2.6976 | 1.9672 | 6.0074 | 14500.9 | 12.9054
39616.591 | 2008-06-17 | L----L-- | 2.6969 | 1.9672 | 6.0515 | 14527.2 | 12.804
39616.5917 | 2008-06-17 | L----L-- | 2.6976 | 1.9672 | 6.0369 | 14527.2 | 12.7002
39616.5924 | 2008-06-17 | L----L-- | 2.6975 | 1.9672 | 6.0221 | 14516.7 | 12.7898
39616.5931 | 2008-06-17 | L----L-- | 2.6988 | 1.9672 | 6.0663 | 14511.4 | 12.9074
39616.5938 | 2008-06-17 | L----L-- | 2.6981 | 1.9672 | 6.0369 | 14495.6 | 13.0847
39616.5944 | 2008-06-17 | L----L-- | 2.6969 | 1.9672 | 6.0221 | 14495.6 | 12.9821
39616.5951 | 2008-06-17 | L----L-- | 2.697 | 1.9672 | 6.0221 | 14506.1 | 12.7959
39616.5958 | 2008-06-17 | L----L-- | 2.6962 | 1.9672 | 6.0221 | 14516.7 | 12.7002
39616.5965 | 2008-06-17 | L----L-- | 2.6953 | 1.9672 | 6.0074 | 14469.3 | 13.0065
39616.5972 | 2008-06-17 | L----L-- | 2.6951 | 1.9672 | 6.0221 | 14490.3 | 12.9094
39616.5979 | 2008-06-17 | L----L-- | 2.6966 | 1.9672 | 6.0221 | 14506.1 | 12.7593
39616.5986 | 2008-06-17 | L----L-- | 2.6955 | 1.9672 | 6.0369 | 14516.7 | 12.6798
39616.5993 | 2008-06-17 | L----L-- | 2.6955 | 1.9672 | 6.0221 | 14474.5 | 12.9861
39616.6 | 2008-06-17 | L----L-- | 2.6959 | 1.9672 | 6.0221 | 14511.4 | 12.7776
39616.6007 | 2008-06-17 | L----L-- | 2.6949 | 1.9672 | 5.978 | 14506.1 | 12.7063
39616.6014 | 2008-06-17 | L----L-- | 2.6953 | 1.9672 | 6.0369 | 14469.3 | 12.9982
39616.6021 | 2008-06-17 | L----L-- | 2.6962 | 1.9672 | 6.0074 | 14506.1 | 12.7858
39616.6028 | 2008-06-17 | L----L-- | 2.6963 | 1.9672 | 6.0369 | 14516.7 | 12.739
39616.6035 | 2008-06-17 | L----L-- | 2.6969 | 1.9672 | 6.0221 | 14516.7 | 12.7613
39616.6042 | 2008-06-17 | L----L-- | 2.6966 | 1.9672 | 6.0221 | 14506.1 | 12.9881
39616.6049 | 2008-06-17 | L----L-- | 2.699 | 1.9672 | 6.0663 | 14511.4 | 13.0213
39616.6056 | 2008-06-17 | L----L-- | 2.6984 | 1.9672 | 6.0369 | 14516.7 | 13.0235
39616.6062 | 2008-06-17 | L----L-- | 2.6978 | 1.9672 | 6.0369 | 14506.1 | 13.0911
39616.6069 | 2008-06-17 | L----L-- | 2.6998 | 1.9672 | 6.0221 | 14516.7 | 13.0192
39616.6076 | 2008-06-17 | L----L-- | 2.6968 | 1.9672 | 5.9927 | 14511.4 | 12.9074

## Detector_Raw
;; | Model | 80I- | (prog | iSeries | 80i | 01.00.21.231)
;;
Time | Date | Flags | hg0 | hg2 | hgt | rfint | intt | rctt | prbt | cnvt | umbt | vntp | orfp | dilp | bbkp | edup | vac | smplf | pmtv | pres | dilf
39616.1854 | 2008-06-17 | #NAME? | 0.087597 | 0.48526 | 0.57286 | 91153 | 30.4656 | 45.0293 | 228.126 | 759.194 | 120.232 | -3.0474 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25218 | 582.751 | 33.1626 | 40
39616.1861 | 2008-06-17 | #NAME? | 0.079579 | 0.47527 | 0.55485 | 91147.8 | 30.5372 | 45.1597 | 229.963 | 756.569 | 121.282 | -3.0004 | -16.7218 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26465 | 582.751 | 33.1626 | 40
39616.1868 | 2008-06-17 | #NAME? | 0.071311 | 0.47155 | 0.54286 | 91140.7 | 30.585 | 45.2575 | 235.214 | 756.569 | 118.394 | -3.0098 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.2532 | 583.121 | 33.1626 | 40
39616.1875 | 2008-06-17 | #NAME? | 0.065088 | 0.45719 | 0.52228 | 91132.7 | 30.6328 | 44.9673 | 237.051 | 758.669 | 120.757 | -3.0286 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26211 | 582.751 | 33.1626 | 40
39616.1882 | 2008-06-17 | #NAME? | 0 | 0 | 0 | 91131.1 | 30.7283 | 45.2901 | 236.001 | 761.819 | 119.969 | -3.038 | -16.7486 | 42.2059 | 49.905 | 15.2083 | 18.3547 | 0.25295 | 582.751 | 32.8611 | 40
39616.1889 | 2008-06-17 | #NAME? | 0 | 0 | 0 | 91125.6 | 30.7999 | 45.1923 | 233.376 | 761.819 | 119.444 | -3.0098 | -16.7619 | 42.2059 | 50.8886 | 15.2083 | 18.6758 | 0.26236 | 582.751 | 33.4641 | 40
39616.1896 | 2008-06-17 | #NAME? | 0 | 0 | 0 | 91119.3 | 30.7521 | 45.0619 | 231.801 | 758.407 | 121.807 | -3.0098 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.6861 | 0.2532 | 583.121 | 33.1626 | 40
39616.1903 | 2008-06-17 | #NAME? | -0.013983 | -0.14578 | -0.15977 | 91125.7 | 30.5372 | 45.0945 | 230.488 | 756.306 | 118.657 | -3.038 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26211 | 582.751 | 33.1626 | 40
39616.191 | 2008-06-17 | #NAME? | -0.014216 | 0.16664 | 0.15243 | 91140.6 | 30.394 | 45.2575 | 229.176 | 757.357 | 120.232 | -3.0004 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25371 | 583.121 | 32.8611 | 40
39616.1917 | 2008-06-17 | #NAME? | 0.049643 | 0.42281 | 0.47246 | 91151.9 | 30.1791 | 44.878 | 228.126 | 759.194 | 121.02 | -3.0004 | -16.7619 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26567 | 582.751 | 33.1626 | 40
39616.1924 | 2008-06-17 | #NAME? | 0.11454 | 0.37001 | 0.48455 | 91170.5 | 30.0597 | 44.997 | 231.276 | 761.819 | 118.919 | -3.0004 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25345 | 582.751 | 33.1626 | 40
39616.1931 | 2008-06-17 | #NAME? | 0.095591 | 0.39591 | 0.4915 | 91179.8 | 30.0836 | 45.1271 | 236.264 | 762.344 | 121.282 | -3.0098 | -16.7619 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26262 | 583.121 | 33.1626 | 40
39616.1937 | 2008-06-17 | #NAME? | 0.075249 | 0.43895 | 0.5142 | 91181.2 | 30.1075 | 45.2575 | 237.051 | 758.932 | 119.444 | -3.0192 | -16.7218 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25447 | 583.121 | 33.1626 | 40
39616.1944 | 2008-06-17 | #NAME? | 0.081123 | 0.45604 | 0.53716 | 91177.4 | 30.2268 | 44.9673 | 235.476 | 756.306 | 119.969 | -3.0004 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26312 | 582.751 | 33.1626 | 40
39616.1951 | 2008-06-17 | #NAME? | 0.087433 | 0.48853 | 0.57596 | 91172.2 | 30.2507 | 44.9673 | 233.901 | 757.094 | 121.545 | -3.0098 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25422 | 582.751 | 33.1626 | 40
39616.1958 | 2008-06-17 | #NAME? | 0.087769 | 0.52726 | 0.61503 | 91166 | 30.3462 | 45.0619 | 232.589 | 759.194 | 118.394 | -2.9817 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26236 | 582.751 | 33.1626 | 40
39616.1965 | 2008-06-17 | #NAME? | 0.088008 | 0.49458 | 0.58258 | 91158.7 | 30.3701 | 44.997 | 231.013 | 762.082 | 120.495 | -3.0286 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25498 | 583.121 | 33.1626 | 40
39616.1972 | 2008-06-17 | #NAME? | 0.085179 | 0.46377 | 0.54895 | 91151.5 | 30.4656 | 45.1597 | 229.701 | 761.557 | 120.495 | -3.0004 | -16.7619 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.2644 | 583.121 | 33.1626 | 40
39616.1979 | 2008-06-17 | #NAME? | 0.082301 | 0.4491 | 0.5314 | 91145.9 | 30.5134 | 44.878 | 228.388 | 758.144 | 119.182 | -3.0192 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25244 | 583.121 | 33.1626 | 40
39616.1986 | 2008-06-17 | #NAME? | 0.095102 | 0.41903 | 0.51413 | 91139.5 | 30.6089 | 45.0619 | 229.438 | 756.306 | 121.282 | -3.0004 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26236 | 582.751 | 33.1626 | 40
39616.1993 | 2008-06-17 | #NAME? | 0.10816 | 0.39597 | 0.50413 | 91134 | 30.6328 | 45.1923 | 234.426 | 757.357 | 119.444 | -2.9911 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25422 | 582.751 | 33.1626 | 40
39616.2 | 2008-06-17 | #NAME? | 0.094016 | 0.40024 | 0.49426 | 91128.4 | 30.7521 | 44.997 | 237.051 | 759.982 | 119.969 | -3.0286 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26542 | 582.751 | 33.1626 | 40
39616.2007 | 2008-06-17 | #NAME? | 0.079411 | 0.42431 | 0.50372 | 91120.9 | 30.776 | 45.0945 | 236.526 | 762.607 | 121.282 | -3.0192 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25524 | 582.751 | 32.8611 | 40
39616.2014 | 2008-06-17 | #NAME? | 0.072675 | 0.44084 | 0.51352 | 91117.9 | 30.6566 | 45.2249 | 234.951 | 760.507 | 118.394 | -3.0286 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.2644 | 583.121 | 33.1626 | 40
39616.2021 | 2008-06-17 | #NAME? | 0.066066 | 0.48169 | 0.54775 | 91125.9 | 30.4656 | 45.3227 | 233.376 | 757.357 | 121.02 | -3.0286 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25549 | 583.121 | 32.8611 | 40
39616.2028 | 2008-06-17 | #NAME? | 0.079258 | 0.50314 | 0.58239 | 91141.8 | 30.2746 | 45.1271 | 232.063 | 756.306 | 119.969 | -3.0192 | -16.7218 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26542 | 582.751 | 33.1626 | 40
39616.2035 | 2008-06-17 | #NAME? | 0.092775 | 0.4771 | 0.56988 | 91157.9 | 30.2268 | 44.9673 | 230.488 | 757.619 | 119.182 | -3.0286 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25575 | 582.751 | 32.8611 | 40
39616.2042 | 2008-06-17 | #NAME? | 0.082879 | 0.4737 | 0.55658 | 91165.3 | 30.2746 | 44.9078 | 229.176 | 760.244 | 121.545 | -3.0004 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26567 | 583.121 | 33.1626 | 40
39616.2049 | 2008-06-17 | #NAME? | 0.072595 | 0.47794 | 0.55053 | 91165.9 | 30.2985 | 45.0619 | 228.126 | 762.344 | 119.182 | -3.0098 | -16.7619 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25498 | 583.121 | 33.1626 | 40
39616.2056 | 2008-06-17 | #NAME? | 0.057114 | 0.4875 | 0.54461 | 91162.1 | 30.394 | 44.9375 | 230.751 | 760.507 | 119.969 | -3.0098 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26338 | 582.751 | 33.1626 | 40
39616.2063 | 2008-06-17 | #NAME? | 0.041544 | 0.49075 | 0.53229 | 91157.2 | 30.4417 | 45.1597 | 236.001 | 756.831 | 121.282 | -3.0286 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25524 | 583.121 | 33.1626 | 40
39616.2069 | 2008-06-17 | #NAME? | 0.056242 | 0.46363 | 0.51987 | 91149.8 | 30.4895 | 45.3227 | 237.051 | 756.569 | 118.394 | -3.0098 | -16.7619 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26516 | 583.121 | 33.1626 | 40
39616.2076 | 2008-06-17 | #NAME? | 0.071444 | 0.48933 | 0.56077 | 91148.3 | 30.5611 | 44.9078 | 236.001 | 758.669 | 121.02 | -3.0286 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25447 | 582.751 | 33.1626 | 40
39616.2083 | 2008-06-17 | #NAME? | 0.083173 | 0.51939 | 0.60257 | 91142 | 30.6328 | 45.2575 | 234.426 | 761.557 | 119.969 | -3.0192 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26363 | 582.751 | 33.1626 | 40
39616.209 | 2008-06-17 | #NAME? | 0.094845 | 0.47152 | 0.56636 | 91137.2 | 30.6805 | 45.0293 | 232.851 | 762.344 | 119.707 | -3.0098 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25524 | 583.121 | 32.8611 | 40
39616.2097 | 2008-06-17 | #NAME? | 0.10134 | 0.42752 | 0.52885 | 91133.5 | 30.7521 | 45.3879 | 231.276 | 759.457 | 121.807 | -3.0004 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26414 | 582.751 | 33.4641 | 40
39616.2104 | 2008-06-17 | #NAME? | 0.10774 | 0.43106 | 0.5388 | 91127.9 | 30.7044 | 45.2249 | 229.963 | 756.306 | 118.657 | -2.9911 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.2532 | 583.121 | 32.8611 | 40
39616.2111 | 2008-06-17 | #NAME? | 0.095701 | 0.45384 | 0.54954 | 91132.6 | 30.5372 | 44.878 | 228.651 | 756.831 | 120.495 | -3.0004 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26567 | 582.751 | 33.1626 | 40
39616.2118 | 2008-06-17 | #NAME? | 0.083351 | 0.46449 | 0.54784 | 91146.9 | 30.3223 | 45.5834 | 228.913 | 758.669 | 120.757 | -3.0192 | -16.7619 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25549 | 583.121 | 33.1626 | 40
39616.2125 | 2008-06-17 | #NAME? | 0.077338 | 0.4686 | 0.54594 | 91160.8 | 30.1075 | 44.9673 | 233.376 | 761.557 | 118.919 | -3.038 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26567 | 582.751 | 33.1626 | 40
39616.2132 | 2008-06-17 | #NAME? | 0.071433 | 0.47034 | 0.54177 | 91175.7 | 30.1791 | 44.9673 | 237.051 | 762.344 | 121.282 | -3.0098 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25396 | 582.751 | 33.1626 | 40
39616.2139 | 2008-06-17 | #NAME? | 0.075437 | 0.46213 | 0.53756 | 91176.8 | 30.203 | 45.1271 | 236.526 | 759.194 | 119.707 | -3.0286 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26262 | 582.751 | 33.1626 | 40
39616.2146 | 2008-06-17 | #NAME? | 0.079609 | 0.46055 | 0.54016 | 91174.4 | 30.2985 | 44.878 | 234.951 | 756.306 | 119.444 | -3.0474 | -16.7753 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25575 | 582.751 | 33.1626 | 40
39616.2153 | 2008-06-17 | #NAME? | 0.07671 | 0.46616 | 0.54287 | 91169 | 30.3462 | 45.2901 | 233.376 | 756.831 | 121.807 | -3.0286 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26185 | 582.751 | 33.4641 | 40
39616.216 | 2008-06-17 | #NAME? | 0.073694 | 0.48068 | 0.55438 | 91163.3 | 30.3701 | 45.0293 | 232.063 | 759.457 | 118.657 | -3.0098 | -16.7619 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25396 | 582.751 | 32.8611 | 40
39616.2167 | 2008-06-17 | #NAME? | 0.071682 | 0.49435 | 0.56603 | 91160.8 | 30.4656 | 44.997 | 230.488 | 762.344 | 120.232 | -3.0286 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26542 | 582.751 | 33.1626 | 40
39616.2174 | 2008-06-17 | #NAME? | 0.069686 | 0.4971 | 0.56678 | 91156 | 30.5134 | 45.2901 | 229.176 | 761.294 | 120.495 | -3.038 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25549 | 582.751 | 32.8611 | 40
39616.2181 | 2008-06-17 | #NAME? | 0.076401 | 0.49095 | 0.56735 | 91151 | 30.5611 | 45.1271 | 228.126 | 757.882 | 118.919 | -3.0568 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26465 | 582.751 | 33.1626 | 40
39616.2188 | 2008-06-17 | #NAME? | 0.083261 | 0.45237 | 0.53563 | 91145.2 | 30.6328 | 45.5183 | 230.488 | 756.306 | 121.282 | -3.0192 | -16.7152 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25549 | 583.121 | 33.1626 | 40
39616.2194 | 2008-06-17 | #NAME? | 0.074939 | 0.42843 | 0.50337 | 91139.3 | 30.7044 | 44.9078 | 235.739 | 757.357 | 119.444 | -3.0098 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26287 | 582.751 | 33.1626 | 40
39616.2201 | 2008-06-17 | #NAME? | 0.066362 | 0.47833 | 0.54469 | 91135.4 | 30.8238 | 45.1923 | 237.051 | 760.244 | 119.707 | -3.0192 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25549 | 583.121 | 33.1626 | 40
39616.2208 | 2008-06-17 | #NAME? | 0.050877 | 0.53637 | 0.58725 | 91128.4 | 30.7283 | 45.0619 | 236.001 | 762.344 | 121.282 | -2.9911 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26516 | 583.121 | 33.1626 | 40
39616.2215 | 2008-06-17 | #NAME? | 0.035273 | 0.55589 | 0.59116 | 91137.3 | 30.5372 | 45.1271 | 234.426 | 760.507 | 118.657 | -3.0098 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25269 | 583.121 | 33.1626 | 40
39616.2222 | 2008-06-17 | #NAME? | 0.064422 | 0.53001 | 0.59443 | 91150 | 30.2985 | 45.4531 | 232.851 | 756.831 | 120.757 | -3.0192 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26363 | 582.751 | 33.1626 | 40
39616.2229 | 2008-06-17 | #NAME? | 0.094308 | 0.47132 | 0.56563 | 91167.1 | 30.1552 | 44.9078 | 231.538 | 756.306 | 119.969 | -3.0192 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25295 | 583.121 | 33.1626 | 40
39616.2236 | 2008-06-17 | #NAME? | 0.090148 | 0.44617 | 0.53632 | 91181.1 | 30.1791 | 45.4857 | 229.963 | 757.619 | 119.444 | -2.9817 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26211 | 582.751 | 33.1626 | 40
39616.2243 | 2008-06-17 | #NAME? | 0.085423 | 0.47343 | 0.55886 | 91184 | 30.203 | 45.2575 | 228.913 | 760.244 | 121.545 | -3.0098 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25473 | 583.121 | 33.1626 | 40
39616.225 | 2008-06-17 | #NAME? | 0.081357 | 0.50091 | 0.58226 | 91183.3 | 30.2746 | 45.1597 | 228.388 | 762.344 | 118.657 | -3.0192 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26312 | 583.121 | 33.4641 | 40
39616.2257 | 2008-06-17 | #NAME? | 0.077302 | 0.50989 | 0.58719 | 91179.5 | 30.3701 | 45.3553 | 232.326 | 761.032 | 120.495 | -3.0098 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25396 | 583.121 | 33.1626 | 40
39616.2264 | 2008-06-17 | #NAME? | 0.099908 | 0.4919 | 0.59181 | 91174.1 | 30.4179 | 44.997 | 236.789 | 757.357 | 121.02 | -3.038 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26414 | 582.751 | 33.1626 | 40
39616.2271 | 2008-06-17 | #NAME? | 0.12296 | 0.4707 | 0.59367 | 91165.5 | 30.4895 | 45.0619 | 236.789 | 756.044 | 118.919 | -3.0098 | -16.7619 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25524 | 583.121 | 33.1626 | 40
39616.2278 | 2008-06-17 | #NAME? | 0.11084 | 0.48463 | 0.59547 | 91158.4 | 30.5134 | 44.878 | 235.476 | 757.357 | 121.282 | -3.038 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26516 | 582.751 | 33.1626 | 40
39616.2285 | 2008-06-17 | #NAME? | 0.098129 | 0.5146 | 0.61273 | 91155 | 30.6328 | 45.3553 | 233.901 | 759.719 | 119.444 | -3.0098 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.25498 | 583.121 | 33.1626 | 40
39616.2292 | 2008-06-17 | #NAME? | 0.083562 | 0.53635 | 0.61991 | 91147.4 | 30.6805 | 45.2575 | 232.326 | 762.344 | 119.444 | -3.038 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.7068 | 0.26287 | 582.751 | 33.1626 | 40
39616.2299 | 2008-06-17 | #NAME? | 0 | 0 | 0 | 91144.6 | 30.7283 | 44.9078 | 230.488 | 761.032 | 121.807 | -3.0474 | -16.7553 | 42.2059 | 49.905 | 15.2083 | 18.3547 | 0.25345 | 582.751 | 33.1626 | 40
39616.2306 | 2008-06-17 | #NAME? | 0 | 0 | 0 | 91139 | 30.8238 | 45.0945 | 228.388 | 757.094 | 118.394 | -3.0004 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.6654 | 0.26262 | 582.751 | 33.4641 | 40
39616.2312 | 2008-06-17 | #NAME? | 0 | 0 | 0 | 91135 | 30.7044 | 44.997 | 229.438 | 756.569 | 120.757 | -3.0286 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.6861 | 0.25295 | 582.751 | 33.1626 | 40
39616.2319 | 2008-06-17 | #NAME? | 0.021159 | -0.15538 | -0.13423 | 91141.6 | 30.5372 | 45.4205 | 234.426 | 758.144 | 119.969 | -3.0004 | -16.7619 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26287 | 583.121 | 33.1626 | 40
39616.2326 | 2008-06-17 | #NAME? | 0.021512 | 0.13423 | 0.15574 | 91152 | 30.394 | 45.0619 | 237.051 | 760.769 | 119.182 | -3.0286 | -16.7619 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25447 | 583.121 | 32.8611 | 40
39616.2333 | 2008-06-17 | #NAME? | 0.069618 | 0.38311 | 0.45273 | 91171.4 | 30.203 | 45.3227 | 236.526 | 762.607 | 121.545 | -3.0474 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26211 | 583.121 | 33.1626 | 40
39616.234 | 2008-06-17 | #NAME? | 0.11851 | 0.40463 | 0.52314 | 91190.8 | 30.0597 | 44.997 | 234.951 | 760.507 | 119.182 | -3.0192 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.256 | 583.121 | 32.8611 | 40
39616.2347 | 2008-06-17 | #NAME? | 0.11532 | 0.47443 | 0.58976 | 91197.7 | 30.0836 | 44.997 | 233.114 | 757.094 | 120.232 | -3.0474 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26211 | 582.751 | 33.1626 | 40
39616.2354 | 2008-06-17 | #NAME? | 0.11127 | 0.46449 | 0.57576 | 91199.4 | 30.1552 | 45.3553 | 231.801 | 756.306 | 121.282 | -3.0098 | -16.7619 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25345 | 582.751 | 32.8611 | 40
39616.2361 | 2008-06-17 | #NAME? | 0.11028 | 0.45015 | 0.56043 | 91195.7 | 30.203 | 45.1271 | 230.488 | 758.144 | 118.657 | -3.0004 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26236 | 583.121 | 33.1626 | 40
39616.2368 | 2008-06-17 | #NAME? | 0.10934 | 0.46173 | 0.57107 | 91186.9 | 30.2985 | 45.1271 | 228.913 | 760.769 | 121.02 | -3.0192 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25269 | 582.751 | 33.1626 | 40
39616.2375 | 2008-06-17 | #NAME? | 0.10249 | 0.47966 | 0.58215 | 91178.1 | 30.3701 | 45.0619 | 228.126 | 762.344 | 119.707 | -3.0474 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26542 | 582.751 | 33.4641 | 40
39616.2382 | 2008-06-17 | #NAME? | 0.095551 | 0.4893 | 0.58485 | 91172.6 | 30.4417 | 44.878 | 231.276 | 760.244 | 119.444 | -3.038 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25295 | 583.121 | 33.1626 | 40
39616.2389 | 2008-06-17 | #NAME? | 0.10421 | 0.48321 | 0.58741 | 91165.4 | 30.5134 | 45.0945 | 236.264 | 756.831 | 121.545 | -2.9911 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26542 | 582.751 | 33.4641 | 40
39616.2396 | 2008-06-17 | #NAME? | 0.11313 | 0.47931 | 0.59244 | 91160.7 | 30.5372 | 44.9673 | 237.051 | 756.569 | 118.657 | -3.0286 | -16.7218 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25269 | 583.121 | 33.1626 | 40
39616.2403 | 2008-06-17 | #NAME? | 0.10592 | 0.49158 | 0.5975 | 91155.5 | 30.6089 | 45.0945 | 235.739 | 758.669 | 120.495 | -3.0286 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26236 | 583.121 | 33.4641 | 40
39616.241 | 2008-06-17 | #NAME? | 0.098443 | 0.50584 | 0.60428 | 91149.5 | 30.6805 | 45.3553 | 234.426 | 761.294 | 120.757 | -3.0286 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25549 | 582.751 | 33.1626 | 40
39616.2417 | 2008-06-17 | #NAME? | 0.081027 | 0.53007 | 0.61109 | 91142.3 | 30.7283 | 45.1271 | 232.589 | 762.344 | 118.919 | -3.0098 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26491 | 582.751 | 33.1626 | 40
39616.2424 | 2008-06-17 | #NAME? | 0.063442 | 0.55921 | 0.62265 | 91137.8 | 30.7999 | 45.0945 | 231.013 | 759.719 | 121.282 | -3.0286 | -16.7218 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25244 | 582.751 | 33.1626 | 40
39616.2431 | 2008-06-17 | #NAME? | 0.073659 | 0.56064 | 0.63429 | 91137.9 | 30.5611 | 45.0945 | 229.701 | 756.569 | 119.707 | -3.0286 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26516 | 582.751 | 33.1626 | 40
39616.2438 | 2008-06-17 | #NAME? | 0.08434 | 0.53077 | 0.61511 | 91149.1 | 30.4417 | 45.3553 | 228.651 | 756.306 | 119.444 | -3.0098 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.2532 | 582.751 | 33.1626 | 40
39616.2444 | 2008-06-17 | #NAME? | 0.080601 | 0.51481 | 0.59541 | 91164.1 | 30.203 | 45.1597 | 228.913 | 758.144 | 121.545 | -3.0098 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26567 | 583.121 | 33.1626 | 40
39616.2451 | 2008-06-17 | #NAME? | 0.076623 | 0.51671 | 0.59333 | 91183.6 | 30.0358 | 45.3553 | 233.901 | 760.769 | 118.919 | -3.0098 | -16.7686 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25575 | 582.751 | 33.1626 | 40
39616.2458 | 2008-06-17 | #NAME? | 0.080644 | 0.51089 | 0.59154 | 91197.6 | 30.0358 | 45.1597 | 237.051 | 762.344 | 120.495 | -3.0474 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26516 | 583.121 | 33.1626 | 40
39616.2465 | 2008-06-17 | #NAME? | 0.084799 | 0.52187 | 0.60666 | 91202.6 | 30.0836 | 44.9673 | 236.264 | 760.244 | 120.757 | -3.0286 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25295 | 582.751 | 33.1626 | 40
39616.2472 | 2008-06-17 | #NAME? | 0.10299 | 0.51908 | 0.62207 | 91198.6 | 30.1313 | 45.2901 | 234.951 | 756.831 | 118.919 | -2.9911 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26287 | 582.751 | 33.1626 | 40
39616.2479 | 2008-06-17 | #NAME? | 0.12142 | 0.50567 | 0.62709 | 91192.4 | 30.2507 | 44.9078 | 233.114 | 756.569 | 121.282 | -3.0192 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25498 | 583.121 | 32.8611 | 40
39616.2486 | 2008-06-17 | #NAME? | 0.11343 | 0.51851 | 0.63194 | 91185.1 | 30.3223 | 45.2901 | 231.801 | 758.669 | 119.182 | -3.038 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26389 | 582.751 | 33.1626 | 40
39616.2493 | 2008-06-17 | #NAME? | 0.10501 | 0.5041 | 0.60911 | 91176.2 | 30.3701 | 44.878 | 230.488 | 761.819 | 120.232 | -3.0098 | -16.7686 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25549 | 583.121 | 33.1626 | 40
39616.25 | 2008-06-17 | #NAME? | 0.075565 | 0.51024 | 0.58581 | 91170.7 | 30.4417 | 44.997 | 229.176 | 762.082 | 121.545 | -3.0286 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26516 | 582.751 | 33.1626 | 40
39616.2507 | 2008-06-17 | #NAME? | 0.045768 | 0.52066 | 0.56643 | 91163.6 | 30.5134 | 44.9673 | 228.126 | 759.194 | 118.657 | -3.0192 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25244 | 583.121 | 33.1626 | 40
39616.2514 | 2008-06-17 | #NAME? | 0.072306 | 0.47481 | 0.54712 | 91158.9 | 30.5372 | 44.997 | 231.276 | 756.569 | 121.02 | -3.0474 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26592 | 582.751 | 33.1626 | 40
39616.2521 | 2008-06-17 | #NAME? | 0.099787 | 0.46022 | 0.56001 | 91155.4 | 30.6328 | 44.997 | 236.264 | 756.569 | 119.969 | -3.0286 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25345 | 583.121 | 33.1626 | 40
39616.2528 | 2008-06-17 | #NAME? | 0.096725 | 0.47672 | 0.57344 | 91147.1 | 30.7044 | 44.997 | 236.789 | 758.407 | 119.444 | -3.0286 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26363 | 582.751 | 33.1626 | 40
39616.2535 | 2008-06-17 | #NAME? | 0.093155 | 0.49394 | 0.5871 | 91142.3 | 30.776 | 45.2901 | 235.739 | 761.032 | 121.545 | -3.0192 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25575 | 582.751 | 32.8611 | 40
39616.2542 | 2008-06-17 | #NAME? | 0.083116 | 0.51764 | 0.60076 | 91140.5 | 30.585 | 44.9375 | 234.164 | 762.607 | 118.657 | -3.0098 | -16.7619 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26287 | 582.751 | 33.4641 | 40
39616.2549 | 2008-06-17 | #NAME? | 0.072966 | 0.51118 | 0.58414 | 91151.4 | 30.3462 | 45.3227 | 232.589 | 760.507 | 120.495 | -3.0004 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25295 | 582.751 | 32.8611 | 40
39616.2556 | 2008-06-17 | #NAME? | 0.044353 | 0.52268 | 0.56703 | 91166.2 | 30.1791 | 44.878 | 231.276 | 757.094 | 120.757 | -3.0474 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26592 | 582.751 | 33.1626 | 40
39616.2562 | 2008-06-17 | #NAME? | 0.015443 | 0.56441 | 0.57985 | 91182.1 | 30.0597 | 44.9673 | 229.701 | 756.569 | 118.657 | -3.0286 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25396 | 582.751 | 32.8611 | 40
39616.2569 | 2008-06-17 | #NAME? | 0.033423 | 0.55975 | 0.59317 | 91189.9 | 30.1313 | 45.1597 | 228.651 | 757.882 | 121.282 | -3.0192 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26287 | 582.751 | 33.1626 | 40
39616.2576 | 2008-06-17 | #NAME? | 0.052179 | 0.53319 | 0.58537 | 91189.3 | 30.2268 | 45.1923 | 229.176 | 760.244 | 119.707 | -3.0004 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25269 | 583.121 | 33.1626 | 40
39616.2583 | 2008-06-17 | #NAME? | 0.064991 | 0.51223 | 0.57722 | 91182.7 | 30.2985 | 44.9673 | 233.901 | 762.344 | 119.707 | -3.0192 | -16.7285 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26236 | 582.751 | 33.1626 | 40
39616.259 | 2008-06-17 | #NAME? | 0.077705 | 0.50788 | 0.58558 | 91179.1 | 30.3462 | 45.2575 | 237.314 | 760.244 | 121.545 | -3.0004 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25524 | 582.751 | 33.1626 | 40
39616.2597 | 2008-06-17 | #NAME? | 0.097316 | 0.4969 | 0.59421 | 91174.4 | 30.4179 | 44.9673 | 236.789 | 756.569 | 118.657 | -3.0098 | -16.7553 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26211 | 583.121 | 33.4641 | 40
39616.2604 | 2008-06-17 | #NAME? | 0.11705 | 0.51429 | 0.63134 | 91168.7 | 30.4895 | 45.0945 | 234.951 | 756.569 | 120.495 | -2.9911 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25575 | 583.121 | 33.1626 | 40
39616.2611 | 2008-06-17 | #NAME? | 0.10194 | 0.56701 | 0.66894 | 91161.9 | 30.585 | 44.997 | 233.639 | 758.669 | 120.757 | -3.0004 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26185 | 582.751 | 33.1626 | 40
39616.2618 | 2008-06-17 | #NAME? | 0.086246 | 0.53506 | 0.62131 | 91153 | 30.6328 | 45.1923 | 232.063 | 761.032 | 118.919 | -2.9911 | -16.7419 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25549 | 582.751 | 33.1626 | 40
39616.2625 | 2008-06-17 | #NAME? | 0.062075 | 0.51017 | 0.57224 | 91148.5 | 30.6805 | 45.2901 | 230.488 | 762.607 | 121.282 | -3.0286 | -16.7352 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.26211 | 582.751 | 33.1626 | 40
39616.2632 | 2008-06-17 | #NAME? | 0.037762 | 0.54002 | 0.57778 | 91142.6 | 30.7521 | 45.0293 | 229.176 | 759.982 | 119.444 | -3.0286 | -16.7486 | 42.2059 | 50.9225 | 15.2083 | 18.6965 | 0.25244 | 582.751 | 33.1626 | 40
39616.2639 | 2008-06-17 | #NAME? | 0.040516 | 0.54371 | 0.58423 | 91135.6 | 30.8238 | 45.1923 | 228.126 | 756.569 | 119.707 | -3.0004 | -16.7553 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26516 | 582.751 | 33.4641 | 40
39616.2646 | 2008-06-17 | #NAME? | 0.043721 | 0.52468 | 0.5684 | 91133 | 30.776 | 44.9078 | 230.488 | 756.831 | 121.545 | -2.9911 | -16.7419 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25269 | 582.751 | 33.1626 | 40
39616.2653 | 2008-06-17 | #NAME? | 0.053519 | 0.49868 | 0.5522 | 91138.1 | 30.5611 | 45.0945 | 236.001 | 759.194 | 118.394 | -3.0004 | -16.7419 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26363 | 582.751 | 33.4641 | 40
39616.266 | 2008-06-17 | #NAME? | 0.063427 | 0.49625 | 0.55967 | 91152.8 | 30.394 | 44.9673 | 237.051 | 762.344 | 120.757 | -3.0004 | -16.7619 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25295 | 582.751 | 33.1626 | 40
39616.2667 | 2008-06-17 | #NAME? | 0.077638 | 0.48991 | 0.56755 | 91168.8 | 30.2268 | 45.1597 | 236.001 | 761.294 | 120.495 | -3.0192 | -16.7285 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.2644 | 582.751 | 33.1626 | 40
39616.2674 | 2008-06-17 | #NAME? | 0.091915 | 0.49772 | 0.58963 | 91183.1 | 30.1791 | 45.1923 | 234.426 | 757.357 | 119.182 | -3.038 | -16.7419 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25549 | 582.751 | 33.1626 | 40
39616.2681 | 2008-06-17 | #NAME? | 0.073138 | 0.53881 | 0.61195 | 91188.5 | 30.2746 | 45.0945 | 232.851 | 756.044 | 121.282 | -3.0004 | -16.7218 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26312 | 582.751 | 33.1626 | 40
39616.2688 | 2008-06-17 | #NAME? | 0.053817 | 0.55906 | 0.61288 | 91185.2 | 30.2985 | 45.2901 | 231.538 | 757.619 | 119.444 | -3.038 | -16.7619 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25269 | 583.121 | 33.1626 | 40
39616.2694 | 2008-06-17 | #NAME? | 0.058987 | 0.55446 | 0.61345 | 91180.4 | 30.3701 | 44.9375 | 229.963 | 760.244 | 119.969 | -3.0286 | -16.7619 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26312 | 582.751 | 33.1626 | 40
39616.2701 | 2008-06-17 | #NAME? | 0.064566 | 0.51824 | 0.58281 | 91175.4 | 30.4656 | 45.1923 | 228.651 | 762.607 | 121.545 | -3.0004 | -16.7486 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25244 | 582.751 | 33.1626 | 40
39616.2708 | 2008-06-17 | #NAME? | 0.060383 | 0.48223 | 0.54261 | 91170.5 | 30.5134 | 45.0945 | 228.388 | 760.769 | 118.394 | -3.0192 | -16.7218 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26542 | 583.121 | 33.1626 | 40
39616.2715 | 2008-06-17 | #NAME? | 0 | 0 | 0 | 91163.6 | 30.6328 | 45.0619 | 232.326 | 757.357 | 120.757 | -3.0286 | -16.7352 | 42.2059 | 49.939 | 15.1743 | 18.3443 | 0.25498 | 582.751 | 33.1626 | 40
39616.2722 | 2008-06-17 | #NAME? | 0 | 0 | 0 | 91157.5 | 30.6566 | 45.2249 | 235.739 | 756.306 | 120.232 | -3.0192 | -16.7352 | 42.2059 | 50.9225 | 15.1743 | 18.6654 | 0.26516 | 582.751 | 33.1626 | 40
39616.2729 | 2008-06-17 | #NAME? | 0 | 0 | 0 | 91154.2 | 30.776 | 45.2575 | 235.476 | 756.831 | 119.444 | -3.0286 | -16.7419 | 42.2059 | 50.9225 | 15.1743 | 18.6861 | 0.25371 | 583.121 | 33.1626 | 40
39616.2736 | 2008-06-17 | #NAME? | 0.0029979 | -0.12826 | -0.12526 | 91146 | 30.7999 | 44.9078 | 234.164 | 759.194 | 121.545 | -2.9817 | -16.7419 | 42.2059 | 50.9225 | 15.1743 | 18.6861 | 0.26185 | 582.751 | 33.4641 | 40
39616.2743 | 2008-06-17 | #NAME? | 0.0030478 | 0.186 | 0.18905 | 91141.2 | 30.776 | 45.0293 | 232.589 | 762.082 | 118.919 | -3.0286 | -16.7285 | 42.2059 | 50.9225 | 15.1743 | 18.6861 | 0.25269 | 583.121 | 33.1626 | 40
39616.275 | 2008-06-17 | #NAME? | 0.053804 | 0.45697 | 0.51077 | 91145.7 | 30.5134 | 45.2901 | 231.013 | 762.082 | 120.495 | -3.038 | -16.7553 | 42.2059 | 50.9225 | 15.1743 | 18.6861 | 0.26516 | 582.751 | 33.4641 | 40
39616.2757 | 2008-06-17 | #NAME? | 0.1054 | 0.43602 | 0.54142 | 91160.7 | 30.3462 | 45.1923 | 229.701 | 758.669 | 121.02 | -3.038 | -16.7285 | 42.2059 | 50.9225 | 15.1743 | 18.6861 | 0.25345 | 583.121 | 33.1626 | 40
39616.2764 | 2008-06-17 | #NAME? | 0.11528 | 0.45193 | 0.56721 | 91178.4 | 30.1791 | 45.0945 | 228.388 | 756.306 | 118.657 | -2.9911 | -16.7619 | 42.2059 | 50.9225 | 15.1743 | 18.6861 | 0.26262 | 582.751 | 33.1626 | 40
39616.2771 | 2008-06-17 | #NAME? | 0.12445 | 0.43201 | 0.55646 | 91189.9 | 30.2507 | 44.878 | 229.176 | 757.357 | 121.282 | -3.0098 | -16.7352 | 42.2059 | 50.9225 | 15.1743 | 18.6861 | 0.25422 | 583.121 | 33.1626 | 40
39616.2778 | 2008-06-17 | #NAME? | 0.11838 | 0.42673 | 0.54511 | 91190.7 | 30.2985 | 45.1923 | 234.164 | 759.982 | 119.707 | -2.9817 | -16.7486 | 42.2059 | 50.9225 | 15.1743 | 18.6861 | 0.26491 | 583.121 | 33.1626 | 40
39616.2785 | 2008-06-17 | #NAME? | 0.11205 | 0.43766 | 0.54971 | 91186.2 | 30.3701 | 44.9078 | 237.051 | 762.607 | 119.444 | -3.0004 | -16.7619 | 42.2059 | 50.9225 | 15.1743 | 18.6861 | 0.25422 | 583.121 | 33.1626 | 40
39616.2792 | 2008-06-17 | #NAME? | 0.10381 | 0.45077 | 0.55458 | 91181.1 | 30.4895 | 45.2249 | 236.526 | 761.032 | 121.545 | -3.0098 | -16.7486 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26363 | 582.751 | 33.1626 | 40
39616.2799 | 2008-06-17 | #NAME? | 0.09554 | 0.51083 | 0.60637 | 91175.1 | 30.5372 | 45.0945 | 235.214 | 757.357 | 118.657 | -3.038 | -16.7686 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25396 | 583.121 | 33.1626 | 40
39616.2806 | 2008-06-17 | #NAME? | 0.10189 | 0.55706 | 0.65895 | 91167.7 | 30.585 | 45.0293 | 233.376 | 756.044 | 120.495 | -3.0192 | -16.7486 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26516 | 582.751 | 33.1626 | 40
39616.2812 | 2008-06-17 | #NAME? | 0.10848 | 0.55208 | 0.66056 | 91162.4 | 30.6805 | 45.4857 | 232.063 | 757.882 | 120.757 | -2.9911 | -16.7619 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25244 | 582.751 | 33.1626 | 40
39616.2819 | 2008-06-17 | #NAME? | 0.11279 | 0.54854 | 0.66133 | 91154.5 | 30.7283 | 45.1923 | 230.488 | 760.507 | 118.919 | -3.0474 | -16.7285 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26516 | 583.121 | 33.1626 | 40
39616.2826 | 2008-06-17 | #NAME? | 0.11705 | 0.53356 | 0.65061 | 91152 | 30.8477 | 44.997 | 229.176 | 762.607 | 121.282 | -3.0192 | -16.7419 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25371 | 582.751 | 33.1626 | 40
39616.2833 | 2008-06-17 | #NAME? | 0.12006 | 0.51964 | 0.6397 | 91144.9 | 30.8715 | 44.997 | 228.126 | 760.769 | 119.444 | -2.9911 | -16.7486 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26262 | 582.751 | 33.4641 | 40
39616.284 | 2008-06-17 | #NAME? | 0.12304 | 0.51236 | 0.6354 | 91142.2 | 30.776 | 44.878 | 231.013 | 757.094 | 119.707 | -3.0192 | -16.7486 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25295 | 582.751 | 33.1626 | 40
39616.2847 | 2008-06-17 | #NAME? | 0.11732 | 0.5139 | 0.63122 | 91148.6 | 30.6328 | 45.4857 | 236.264 | 756.306 | 121.807 | -3.0004 | -16.7553 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26491 | 582.751 | 33.1626 | 40
39616.2854 | 2008-06-17 | #NAME? | 0.11145 | 0.52485 | 0.63631 | 91161.5 | 30.4179 | 44.878 | 237.051 | 757.619 | 118.394 | -3.0192 | -16.7285 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25549 | 583.121 | 33.1626 | 40
39616.2861 | 2008-06-17 | #NAME? | 0.10373 | 0.53782 | 0.64155 | 91177.8 | 30.3223 | 45.3227 | 235.739 | 760.507 | 120.495 | -3.0192 | -16.7352 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26211 | 582.751 | 33.1626 | 40
39616.2868 | 2008-06-17 | #NAME? | 0.095981 | 0.55953 | 0.65551 | 91193.1 | 30.1075 | 45.0945 | 234.426 | 762.607 | 120.757 | -3.0286 | -16.7553 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25396 | 583.121 | 32.8611 | 40
39616.2875 | 2008-06-17 | #NAME? | 0.094861 | 0.57476 | 0.66963 | 91206.8 | 30.1552 | 44.997 | 232.851 | 760.244 | 119.182 | -3.0098 | -16.7553 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26465 | 583.121 | 33.1626 | 40
39616.2882 | 2008-06-17 | #NAME? | 0.093849 | 0.57344 | 0.66729 | 91207.5 | 30.2268 | 45.3879 | 231.276 | 756.569 | 121.545 | -3.0286 | -16.7619 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25549 | 583.121 | 33.1626 | 40
39616.2889 | 2008-06-17 | #NAME? | 0.10475 | 0.55993 | 0.66468 | 91201.5 | 30.3223 | 44.9078 | 229.963 | 756.306 | 119.182 | -2.9911 | -16.7352 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26567 | 583.121 | 33.1626 | 40
39616.2896 | 2008-06-17 | #NAME? | 0.11585 | 0.54042 | 0.65627 | 91195.2 | 30.4656 | 45.0945 | 228.651 | 758.932 | 120.232 | -3.0286 | -16.7686 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25575 | 583.121 | 33.1626 | 40
39616.2903 | 2008-06-17 | #NAME? | 0.10189 | 0.54588 | 0.64776 | 91185 | 30.4895 | 44.9078 | 228.651 | 761.819 | 121.02 | -3.038 | -16.7352 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26516 | 582.751 | 33.1626 | 40
39616.291 | 2008-06-17 | #NAME? | 0.087498 | 0.54684 | 0.63434 | 91177.3 | 30.585 | 45.1597 | 233.114 | 761.819 | 118.657 | -3.0192 | -16.7218 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.256 | 582.751 | 33.1626 | 40
39616.2917 | 2008-06-17 | #NAME? | 0.097257 | 0.52357 | 0.62083 | 91168.1 | 30.6328 | 44.997 | 236.789 | 758.144 | 121.02 | -2.9911 | -16.7619 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26516 | 582.751 | 33.4641 | 40
39616.2924 | 2008-06-17 | #NAME? | 0.10742 | 0.52577 | 0.63319 | 91162.8 | 30.7283 | 45.3227 | 236.789 | 756.569 | 119.969 | -3.0286 | -16.7419 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.256 | 582.751 | 33.1626 | 40
39616.2931 | 2008-06-17 | #NAME? | 0.099198 | 0.54678 | 0.64597 | 91155.5 | 30.8238 | 45.0293 | 234.951 | 757.619 | 119.444 | -3.0192 | -16.7352 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26338 | 583.121 | 33.1626 | 40
39616.2937 | 2008-06-17 | #NAME? | 0.090668 | 0.56971 | 0.66038 | 91149.1 | 30.7999 | 45.0619 | 233.376 | 760.244 | 121.545 | -2.9911 | -16.7419 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25244 | 582.751 | 33.1626 | 40
39616.2944 | 2008-06-17 | #NAME? | 0.11812 | 0.5567 | 0.67482 | 91149.9 | 30.6328 | 45.1923 | 232.063 | 762.607 | 118.394 | -3.0192 | -16.7553 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26542 | 582.751 | 33.1626 | 40
39616.2951 | 2008-06-17 | #NAME? | 0.14618 | 0.52218 | 0.66836 | 91162.8 | 30.4179 | 44.9673 | 230.488 | 760.769 | 120.495 | -3.0286 | -16.7619 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25575 | 582.751 | 33.1626 | 40
39616.2958 | 2008-06-17 | #NAME? | 0.13104 | 0.53051 | 0.66155 | 91178 | 30.203 | 44.997 | 229.176 | 757.357 | 120.495 | -3.0192 | -16.7352 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26363 | 582.751 | 33.1626 | 40
39616.2965 | 2008-06-17 | #NAME? | 0.11519 | 0.53411 | 0.6493 | 91197.6 | 30.0597 | 45.0619 | 228.126 | 756.569 | 118.919 | -3.0004 | -16.7419 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25295 | 583.121 | 33.1626 | 40
39616.2972 | 2008-06-17 | #NAME? | 0.1198 | 0.51717 | 0.63697 | 91208 | 30.1075 | 45.0293 | 230.751 | 758.144 | 121.282 | -2.9911 | -16.7419 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26262 | 583.121 | 33.1626 | 40
39616.2979 | 2008-06-17 | #NAME? | 0.12475 | 0.51188 | 0.63663 | 91208.6 | 30.2268 | 45.0293 | 235.739 | 761.294 | 119.707 | -3.0004 | -16.7619 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25345 | 583.121 | 33.1626 | 40
39616.2986 | 2008-06-17 | #NAME? | 0.12502 | 0.51148 | 0.6365 | 91197.8 | 30.3223 | 45.3879 | 237.051 | 762.607 | 119.707 | -3.0192 | -16.7619 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26542 | 583.121 | 33.1626 | 40
39616.2993 | 2008-06-17 | #NAME? | 0.12521 | 0.52774 | 0.65295 | 91189 | 30.4179 | 45.0945 | 236.001 | 759.457 | 121.545 | -3.0192 | -16.7352 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25575 | 582.751 | 33.1626 | 40
39616.3 | 2008-06-17 | #NAME? | 0.13153 | 0.53814 | 0.66968 | 91181.7 | 30.4656 | 44.997 | 234.426 | 756.569 | 118.394 | -3.0098 | -16.7553 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26465 | 582.751 | 33.1626 | 40
39616.3007 | 2008-06-17 | #NAME? | 0.13797 | 0.53331 | 0.67127 | 91173.6 | 30.5611 | 45.1271 | 232.851 | 756.831 | 120.757 | -3.0474 | -16.7486 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25575 | 583.121 | 33.1626 | 40
39616.3014 | 2008-06-17 | #NAME? | 0.11509 | 0.55752 | 0.67261 | 91164.9 | 30.6805 | 45.0293 | 231.276 | 759.719 | 120.495 | -3.0286 | -16.7486 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26389 | 583.121 | 33.1626 | 40
39616.3021 | 2008-06-17 | #NAME? | 0.091721 | 0.57644 | 0.66816 | 91156.3 | 30.6805 | 45.1597 | 229.963 | 762.344 | 119.182 | -3.0004 | -16.7352 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25244 | 583.121 | 33.1626 | 40
39616.3028 | 2008-06-17 | #NAME? | 0.07996 | 0.58364 | 0.66361 | 91153 | 30.4895 | 45.1923 | 228.651 | 761.557 | 121.282 | -3.0098 | -16.7352 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26592 | 582.751 | 33.4641 | 40
39616.3035 | 2008-06-17 | #NAME? | 0.068392 | 0.56107 | 0.62946 | 91164.2 | 30.394 | 45.2575 | 229.176 | 757.882 | 119.707 | -3.038 | -16.7419 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25244 | 582.751 | 33.1626 | 40
39616.3042 | 2008-06-17 | #NAME? | 0.091899 | 0.50293 | 0.59483 | 91172.8 | 30.2268 | 44.9375 | 233.639 | 756.306 | 119.707 | -3.0286 | -16.7486 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26465 | 582.751 | 33.1626 | 40
39616.3049 | 2008-06-17 | #NAME? | 0.11598 | 0.48685 | 0.60283 | 91182.5 | 30.0597 | 45.3227 | 237.051 | 757.882 | 121.807 | -2.9911 | -16.7419 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.25345 | 582.751 | 33.1626 | 40
39616.3056 | 2008-06-17 | #NAME? | 0.12937 | 0.48216 | 0.61153 | 91195.5 | 30.0119 | 44.9673 | 236.789 | 760.769 | 118.657 | -3.0098 | -16.7486 | 42.2059 | 50.9225 | 15.1743 | 18.6965 | 0.26542 | 582.751 | 33.1626 | 40
39616.3063 | 2008-06-17 | #NAME? | 0.14258 | 0.48624 | 0.62882 | 91199.5 | 30.1552 | 45.3879 | 234.689 | 762.082 | 120.757 | -3.0286 | -16.7419 | 4.9993 | 50.9225 | 15.1743 | 7.0953 | 0.25345 | 582.751 | 32.8611 | 40
39616.3069 | 2008-06-17 | #NAME? | 0.12359 | 0.52266 | 0.64625 | 91191.9 | 30.203 | 45.1923 | 233.639 | 759.194 | 120.232 | -3.0004 | -16.7352 | 1.5398 | 50.9225 | 15.1743 | 1.4812 | 0.25269 | 582.751 | 32.5596 | 40
39616.3076 | 2008-06-17 | #NAME? | 0.10405 | 0.2031 | 0.30716 | 91184 | 30.2985 | 45.0945 | 232.063 | 756.306 | 119.182 | -3.0098 | -16.7352 | 1.5398 | 50.9225 | 15.1743 | 1.4812 | 0.25244 | 582.751 | 32.5596 | 40
39616.3083 | 2008-06-17 | #NAME? | 1.5249 | -1.5628 | -0.037936 | 91173.3 | 30.4179 | 45.2249 | 230.751 | 756.831 | 121.545 | -3.0192 | -16.7285 | 1.5398 | 50.9225 | 15.1743 | 1.4709 | 0.25498 | 582.751 | 32.5596 | 40
39616.309 | 2008-06-17 | #NAME? | 2.9699 | 1.6356 | 4.6054 | 91163.4 | 30.4895 | 44.878 | 229.438 | 759.457 | 119.182 | -2.9817 | -16.7553 | 1.5398 | 50.9225 | 15.1743 | 1.4709 | 0.25295 | 582.751 | 32.5596 | 40
39616.3097 | 2008-06-17 | #NAME? | 2.9527 | 6.3799 | 9.3326 | 91152.6 | 30.6328 | 45.4531 | 228.388 | 762.344 | 119.969 | -3.0004 | -16.7352 | 1.5398 | 50.9225 | 15.1743 | 1.4709 | 0.25524 | 582.751 | 32.5596 | 40
39616.3104 | 2008-06-17 | #NAME? | 2.9113 | 6.4741 | 9.3853 | 91141.5 | 30.6805 | 45.1597 | 229.701 | 761.032 | 121.02 | -3.0098 | -16.7486 | 1.5398 | 50.9225 | 15.1743 | 1.4709 | 0.25244 | 583.121 | 32.5596 | 40
39616.3111 | 2008-06-17 | #NAME? | 2.942 | 6.4181 | 9.3601 | 91137.1 | 30.6089 | 45.4205 | 234.689 | 757.094 | 118.657 | -3.038 | -16.7419 | 1.5398 | 50.9225 | 15.1743 | 1.4709 | 0.25676 | 582.751 | 32.5596 | 40
39616.3118 | 2008-06-17 | #NAME? | 2.9739 | 6.3919 | 9.3658 | 91141.5 | 30.4417 | 45.0945 | 237.314 | 756.306 | 121.282 | -3.038 | -16.7352 | 1.5398 | 50.9225 | 15.1743 | 1.4709 | 0.24989 | 583.121 | 32.5596 | 40
39616.3125 | 2008-06-17 | #NAME? | 2.9117 | 6.3066 | 9.2183 | 91153 | 30.2746 | 45.0293 | 236.789 | 758.407 | 119.444 | -3.0192 | -16.7285 | 1.5398 | 50.9225 | 15.1743 | 1.4709 | 0.25498 | 582.751 | 32.5596 | 40
39616.3132 | 2008-06-17 | #NAME? | 0 | 0 | 0 | 91169.8 | 30.0358 | 45.2901 | 234.951 | 761.294 | 119.707 | -3.0004 | -16.7486 | 1.5398 | 49.905 | 15.1743 | 1.3983 | 0.25396 | 583.121 | 32.5596 | 40
39616.3139 | 2008-06-17 | #NAME? | 0 | 0 | 0 | 91186.9 | 29.8753 | 44.9375 | 232.589 | 762.344 | 121.807 | -3.0098 | -16.7352 | 1.5398 | 50.8886 | 15.1743 | 1.4191 | 0.25956 | 582.751 | 33.1626 | 40
39616.3146 | 2008-06-17 | #NAME? | 0 | 0 | 0 | 91198.7 | 29.9206 | 45.1923 | 231.538 | 759.982 | 118.657 | -3.0098 | -16.7352 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25269 | 583.121 | 32.5596 | 40
39616.3153 | 2008-06-17 | #NAME? | 2.4753 | -1.2059 | 1.2695 | 91196.1 | 30.0358 | 45.4205 | 230.226 | 756.569 | 120.495 | -3.0098 | -16.7285 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25676 | 583.121 | 32.5596 | 40
39616.316 | 2008-06-17 | #NAME? | 2.5166 | 2.7019 | 5.2184 | 91190.1 | 30.1313 | 44.8186 | 229.176 | 756.569 | 120.495 | -3.0004 | -16.7486 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25193 | 582.751 | 32.2582 | 40
39616.3167 | 2008-06-17 | #NAME? | 2.7532 | 6.4593 | 9.2125 | 91182.6 | 30.1552 | 45.1271 | 228.126 | 759.457 | 118.919 | -3.0286 | -16.7553 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25829 | 582.751 | 32.5596 | 40
39616.3174 | 2008-06-17 | #NAME? | 2.9939 | 6.3045 | 9.2984 | 91172.1 | 30.3462 | 44.9673 | 231.801 | 762.607 | 121.282 | -3.0286 | -16.7285 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25244 | 582.751 | 32.5596 | 40
39616.3181 | 2008-06-17 | #NAME? | 2.9591 | 6.36 | 9.3192 | 91159.1 | 30.4656 | 45.3879 | 236.789 | 761.294 | 119.707 | -3.0004 | -16.7218 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25778 | 582.751 | 32.5596 | 40
39616.3187 | 2008-06-17 | #NAME? | 2.9197 | 6.367 | 9.2867 | 91147.5 | 30.3701 | 45.4531 | 237.314 | 757.357 | 119.707 | -3.038 | -16.7285 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25015 | 583.121 | 32.5596 | 40
39616.3194 | 2008-06-17 | #NAME? | 2.9177 | 6.3357 | 9.2534 | 91147.7 | 30.2746 | 44.878 | 235.739 | 756.044 | 121.807 | -3.0098 | -16.7285 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25625 | 582.751 | 32.5596 | 40
39616.3201 | 2008-06-17 | #NAME? | 2.9163 | 6.3038 | 9.2201 | 91156.9 | 30.1791 | 45.1597 | 234.426 | 757.619 | 118.657 | -3.0192 | -16.7285 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25218 | 582.751 | 32.2582 | 40
39616.3208 | 2008-06-17 | #NAME? | 2.9483 | 6.2385 | 9.1867 | 91168.5 | 29.9433 | 44.997 | 232.851 | 760.769 | 120.495 | -2.9911 | -16.7419 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.2588 | 582.751 | 32.5596 | 40
39616.3215 | 2008-06-17 | #NAME? | 2.9809 | 6.2035 | 9.1844 | 91184.2 | 29.7846 | 45.0945 | 231.801 | 762.869 | 120.757 | -3.0098 | -16.7285 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.24913 | 583.121 | 32.2582 | 40
39616.3222 | 2008-06-17 | #NAME? | 2.9654 | 6.2171 | 9.1825 | 91196.1 | 29.6939 | 44.9673 | 230.488 | 760.244 | 118.919 | -3.0098 | -16.7619 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25727 | 582.751 | 32.5596 | 40
39616.3229 | 2008-06-17 | #NAME? | 2.9491 | 6.2762 | 9.2253 | 91207.2 | 29.5125 | 44.878 | 229.438 | 756.569 | 121.282 | -2.9911 | -16.7486 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25066 | 583.121 | 32.5596 | 40
39616.3236 | 2008-06-17 | #NAME? | 2.9541 | 6.3148 | 9.2689 | 91218.6 | 29.5125 | 45.4531 | 228.651 | 756.831 | 119.707 | -3.0098 | -16.7218 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25651 | 583.121 | 32.5596 | 40
39616.3243 | 2008-06-17 | #NAME? | 2.9594 | 6.3075 | 9.267 | 91217.4 | 29.5805 | 44.9673 | 229.963 | 758.407 | 119.444 | -3.0098 | -16.7486 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.24913 | 582.751 | 32.5596 | 40
39616.325 | 2008-06-17 | #NAME? | 2.9313 | 6.333 | 9.2643 | 91209.7 | 29.6712 | 45.2575 | 235.476 | 761.557 | 121.807 | -2.9911 | -16.7285 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.2588 | 582.751 | 32.5596 | 40
39616.3257 | 2008-06-17 | #NAME? | 2.9025 | 6.4504 | 9.3529 | 91198.4 | 29.7166 | 45.4205 | 237.314 | 762.344 | 118.657 | -3.0192 | -16.7352 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25244 | 582.751 | 32.2582 | 40
39616.3264 | 2008-06-17 | #NAME? | 2.87 | 6.573 | 9.443 | 91187.5 | 29.7619 | 44.878 | 236.526 | 759.719 | 120.495 | -3.0004 | -16.7619 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.256 | 582.751 | 32.5596 | 40
39616.3271 | 2008-06-17 | #NAME? | 2.8373 | 6.583 | 9.4203 | 91178 | 29.9433 | 45.3227 | 235.214 | 756.569 | 120.495 | -3.038 | -16.7619 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.2504 | 583.491 | 32.2582 | 40
39616.3278 | 2008-06-17 | #NAME? | 2.9005 | 6.4953 | 9.3958 | 91163.2 | 30.0119 | 45.0945 | 233.901 | 756.569 | 118.919 | -3.0098 | -16.7619 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25575 | 583.121 | 32.5596 | 40
39616.3285 | 2008-06-17 | #NAME? | 2.9652 | 6.448 | 9.4132 | 91153.5 | 30.1313 | 45.3227 | 232.589 | 758.407 | 121.545 | -3.0474 | -16.7419 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25193 | 582.751 | 32.2582 | 40
39616.3292 | 2008-06-17 | #NAME? | 2.9351 | 6.4962 | 9.4313 | 91143 | 30.0358 | 45.0293 | 231.276 | 761.819 | 119.182 | -3.0004 | -16.7285 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25625 | 583.121 | 32.5596 | 40
39616.3299 | 2008-06-17 | #NAME? | 2.9034 | 6.5573 | 9.4608 | 91142 | 29.9433 | 44.7591 | 230.226 | 762.344 | 119.969 | -3.038 | -16.7619 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25015 | 582.381 | 32.2582 | 40
39616.3306 | 2008-06-17 | #NAME? | 2.8366 | 6.6538 | 9.4905 | 91145.4 | 29.8073 | 44.8483 | 228.913 | 759.194 | 121.545 | -3.038 | -16.7218 | 1.5398 | 50.9225 | 15.1743 | 1.4294 | 0.25727 | 583.121 | 32.5596 | 40
39616.3313 | 2008-06-17 | #NAME? | 2.7693 | 6.6089 | 9.3781 | 91149.8 | 29.7166 | 45.4205 | 228.651 | 756.569 | 118.657 | -3.0192 | -16.7419 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.25142 | 582.381 | 32.2582 | 40
39616.3319 | 2008-06-17 | #NAME? | 2.8157 | 6.4478 | 9.2635 | 91155.3 | 29.5805 | 44.9375 | 231.538 | 757.094 | 121.02 | -3.0192 | -16.7285 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.25829 | 582.381 | 32.5596 | 40
39616.3326 | 2008-06-17 | #NAME? | 2.864 | 6.387 | 9.251 | 91167.9 | 29.3991 | 44.9673 | 236.789 | 759.194 | 119.969 | -3.0192 | -16.7553 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.25244 | 582.381 | 32.5596 | 40
39616.3333 | 2008-06-17 | #NAME? | 2.8205 | 6.4197 | 9.2402 | 91179.4 | 29.4444 | 45.2249 | 237.314 | 762.082 | 119.182 | -3.038 | -16.7419 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.25625 | 583.121 | 32.5596 | 40
39616.334 | 2008-06-17 | #NAME? | 2.7754 | 6.4645 | 9.24 | 91177.9 | 29.6032 | 44.8483 | 236.264 | 761.819 | 121.282 | -3.0192 | -16.7553 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.24913 | 583.121 | 32.5596 | 40
39616.3347 | 2008-06-17 | #NAME? | 2.8209 | 6.4191 | 9.24 | 91167.3 | 29.6939 | 45.4531 | 234.951 | 757.882 | 119.182 | -3.038 | -16.7285 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.25855 | 582.751 | 32.2582 | 40
39616.3354 | 2008-06-17 | #NAME? | 2.8678 | 6.3616 | 9.2294 | 91158.2 | 29.7846 | 44.9375 | 233.639 | 756.306 | 119.969 | -2.9911 | -16.7619 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.24888 | 582.751 | 32.2582 | 40
39616.3361 | 2008-06-17 | #NAME? | 2.8796 | 6.3391 | 9.2187 | 91144.7 | 29.9206 | 45.2249 | 232.326 | 757.619 | 121.282 | -3.0098 | -16.7352 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.2588 | 582.751 | 32.5596 | 40
39616.3368 | 2008-06-17 | #NAME? | 2.8908 | 6.3727 | 9.2634 | 91128.8 | 30.0358 | 45.1597 | 231.013 | 760.244 | 118.394 | -3.0004 | -16.7753 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.24862 | 583.121 | 32.2582 | 40
39616.3375 | 2008-06-17 | #NAME? | 0.70307 | 1.5855 | 5.3928 | 91115.3 | 30.1552 | 44.997 | 229.963 | 762.607 | 120.757 | -3.0192 | -16.7553 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.25829 | 582.751 | 32.5596 | 40
39616.3382 | 2008-06-17 | #NAME? | 0 | 0 | 9.4344 | 91104.4 | 30.1075 | 45.1923 | 228.913 | 760.244 | 120.495 | -3.0098 | -16.7553 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.26465 | 582.751 | 32.8611 | 40
39616.3389 | 2008-06-17 | #NAME? | 0 | 0 | 9.4216 | 91102.1 | 30.0358 | 44.878 | 228.388 | 756.831 | 119.182 | -3.0098 | -16.7619 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.26262 | 582.751 | 32.5596 | 40
39616.3396 | 2008-06-17 | #NAME? | 0 | 0 | 9.4084 | 91110.4 | 29.898 | 45.0945 | 232.063 | 756.306 | 121.545 | -3.0286 | -16.7285 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.26389 | 583.121 | 32.5596 | 40
39616.3403 | 2008-06-17 | #NAME? | 0 | 0 | 9.4647 | 91115.9 | 29.8526 | 45.4205 | 237.051 | 758.669 | 119.182 | -3.0098 | -16.7285 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.26542 | 582.751 | 32.8611 | 40
39616.341 | 2008-06-17 | #NAME? | 0 | 0 | 9.507 | 91124.7 | 29.6712 | 44.7888 | 237.314 | 761.819 | 120.232 | -3.0004 | -16.7486 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.26465 | 582.751 | 32.8611 | 40
39616.3417 | 2008-06-17 | #NAME? | 0 | 0 | 9.6168 | 91132 | 29.7392 | 45.3227 | 236.001 | 762.082 | 121.282 | -3.0192 | -16.7352 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.26287 | 582.751 | 32.8611 | 40
39616.3424 | 2008-06-17 | #NAME? | 0 | 0 | 9.4994 | 91129.7 | 29.8073 | 44.9673 | 234.689 | 758.407 | 118.657 | -3.0474 | -16.7419 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.26287 | 583.121 | 32.8611 | 40
39616.3431 | 2008-06-17 | #NAME? | 0 | 0 | 9.4758 | 91126.8 | 29.9206 | 45.0293 | 233.376 | 756.569 | 120.757 | -3.0474 | -16.7553 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.26312 | 582.751 | 32.8611 | 40
39616.3438 | 2008-06-17 | #NAME? | 0 | 0 | 9.5044 | 91117.9 | 29.966 | 44.9673 | 232.063 | 757.882 | 120.232 | -3.0192 | -16.7352 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.26236 | 583.121 | 32.8611 | 40
39616.3451 | 2008-06-17 | #NAME? | 0 | 0 | 6.4972 | 91106.8 | 30.0836 | 45.2901 | 230.488 | 760.769 | 118.657 | -3.0098 | -16.7285 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.34302 | 582.751 | 37.0818 | 40
39616.3458 | 2008-06-17 | #NAME? | 0 | 0 | 0.078811 | 91097.2 | 30.1552 | 44.997 | 229.438 | 756.831 | 120.495 | -3.0004 | -16.7553 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.33895 | 582.751 | 37.0818 | 40
39616.3465 | 2008-06-17 | #NAME? | 0 | 0 | -0.0023959 | 91090.8 | 30.2746 | 44.878 | 228.651 | 756.569 | 120.757 | -3.0004 | -16.7419 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.34022 | 582.751 | 37.0818 | 40
39616.3472 | 2008-06-17 | #NAME? | 0 | 0 | 0.00015453 | 91077.3 | 30.3701 | 45.4531 | 229.438 | 757.882 | 118.919 | -3.0098 | -16.7486 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.34048 | 583.121 | 37.3833 | 40
39616.3479 | 2008-06-17 | #NAME? | 0 | 0 | -0.00038781 | 91072.6 | 30.2507 | 44.878 | 234.689 | 761.294 | 121.282 | -3.0192 | -16.7686 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.33946 | 582.751 | 37.3833 | 40
39616.3486 | 2008-06-17 | #NAME? | 0 | 0 | -7.3488e-05 | 91079.8 | 30.203 | 44.997 | 237.314 | 762.344 | 119.444 | -3.0004 | -16.7486 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.3359 | 583.121 | 37.3833 | 40
39616.3493 | 2008-06-17 | #NAME? | 0 | 0 | 0.00031099 | 91085.6 | 30.1552 | 45.2575 | 236.789 | 759.457 | 119.969 | -3.0004 | -16.7486 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.33768 | 582.751 | 37.6848 | 40
39616.35 | 2008-06-17 | #NAME? | 0 | 0 | 0.17405 | 91093.8 | 29.9887 | 44.9673 | 235.476 | 756.306 | 121.545 | -3.038 | -16.7352 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.32623 | 582.751 | 36.7803 | 40
39616.3507 | 2008-06-17 | #NAME? | 0 | 0 | 5.4185 | 91104.7 | 29.9433 | 45.4205 | 234.164 | 757.357 | 118.394 | -3.0004 | -16.7285 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.32623 | 582.751 | 36.7803 | 40
39616.3514 | 2008-06-17 | #NAME? | 0 | 0 | 8.0193 | 91113.7 | 29.966 | 44.8186 | 232.851 | 759.982 | 121.02 | -3.0286 | -16.7553 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.32547 | 583.121 | 36.7803 | 40
39616.3521 | 2008-06-17 | #NAME? | 0 | 0 | 8.0795 | 91115.1 | 30.2507 | 45.4531 | 231.538 | 762.607 | 119.969 | -3.0098 | -16.7419 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.32394 | 582.751 | 36.7803 | 40
39616.3528 | 2008-06-17 | #NAME? | 0 | 0 | 8.0656 | 91102.5 | 30.394 | 45.0619 | 230.488 | 760.507 | 119.444 | -3.0098 | -16.7419 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.32318 | 583.121 | 36.7803 | 40
39616.3535 | 2008-06-17 | #NAME? | 0 | 0 | 8.0663 | 91091.9 | 30.585 | 45.0293 | 229.438 | 756.569 | 121.807 | -3.0192 | -16.7553 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.32318 | 583.121 | 36.7803 | 40
39616.3542 | 2008-06-17 | #NAME? | 0 | 0 | 8.0509 | 91079.3 | 30.7283 | 45.1597 | 228.388 | 756.831 | 118.657 | -3.0568 | -16.7553 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.32598 | 583.121 | 36.7803 | 40
39616.3549 | 2008-06-17 | #NAME? | 0 | 0 | 8.0279 | 91069.7 | 30.8715 | 45.2249 | 230.226 | 759.982 | 120.495 | -3.0192 | -16.7352 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.32241 | 583.121 | 36.7803 | 40
39616.3556 | 2008-06-17 | #NAME? | 0 | 0 | 7.9802 | 91064 | 30.967 | 45.1271 | 235.476 | 762.607 | 120.757 | -3.0098 | -16.7419 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.32216 | 583.121 | 36.7803 | 40
39616.3562 | 2008-06-17 | #NAME? | 0 | 0 | 7.8327 | 91054.5 | 31.1103 | 45.0945 | 237.576 | 761.032 | 118.919 | -2.9911 | -16.7486 | 1.5398 | 50.9565 | 15.1743 | 1.4294 | 0.32318 | 583.491 | 36.7803 | 40
39616.3569 | 2008-06-17 | #NAME? | 0 | 0 | 7.8555 | 91049.1 | 31.0864 | 44.9078 | 236.789 | 757.094 | 121.02 | -3.0004 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.32521 | 582.381 | 36.7803 | 40
39616.3576 | 2008-06-17 | #NAME? | 0 | 0 | 7.8476 | 91047.3 | 31.0626 | 45.1923 | 235.476 | 756.306 | 119.707 | -3.0004 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.32648 | 583.491 | 36.7803 | 40
39616.3583 | 2008-06-17 | #NAME? | 0 | 0 | 7.8331 | 91052.7 | 30.9432 | 44.9375 | 234.164 | 758.144 | 119.444 | -3.0286 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.32267 | 583.491 | 36.7803 | 40
39616.359 | 2008-06-17 | #NAME? | 0 | 0 | 7.8304 | 91062.4 | 30.8715 | 45.3879 | 232.589 | 760.507 | 121.545 | -3.0192 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.3247 | 583.121 | 36.7803 | 40
39616.3597 | 2008-06-17 | #NAME? | 0 | 0 | 7.8221 | 91073.5 | 30.7521 | 45.2249 | 231.538 | 762.607 | 118.657 | -3.0192 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.32598 | 583.121 | 36.7803 | 40
39616.3604 | 2008-06-17 | #NAME? | 0 | 0 | 7.9495 | 91086.6 | 30.6566 | 45.2249 | 230.226 | 760.507 | 120.495 | -3.038 | -16.7218 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.32699 | 582.751 | 36.7803 | 40
39616.3611 | 2008-06-17 | #NAME? | 0 | 0 | 7.9544 | 91091.8 | 30.7999 | 45.0945 | 229.438 | 757.094 | 120.495 | -3.038 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.32674 | 583.121 | 36.7803 | 40
39616.3618 | 2008-06-17 | #NAME? | 0 | 0 | 7.9426 | 91085.8 | 30.8715 | 44.9673 | 228.126 | 756.569 | 118.919 | -3.038 | -16.7553 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.32241 | 583.121 | 36.7803 | 40
39616.3625 | 2008-06-17 | #NAME? | 0 | 0 | 7.9412 | 91077.7 | 31.0148 | 45.4205 | 230.751 | 758.669 | 121.545 | -3.0286 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.32216 | 583.121 | 36.7803 | 40
39616.3632 | 2008-06-17 | #NAME? | 0 | 0 | 7.9293 | 91066.8 | 31.1103 | 44.878 | 236.001 | 761.294 | 119.182 | -3.0192 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.32343 | 582.751 | 36.7803 | 40
39616.3639 | 2008-06-17 | #NAME? | 0 | 0 | 7.8924 | 91059.4 | 31.2297 | 45.2249 | 237.314 | 762.607 | 119.707 | -2.9911 | -16.7285 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.32419 | 583.121 | 36.7803 | 40
39616.3646 | 2008-06-17 | #NAME? | 0 | 0 | 7.7919 | 91049.3 | 31.3252 | 45.2901 | 236.264 | 759.194 | 121.545 | -2.9911 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.32318 | 582.751 | 36.7803 | 40
39616.3653 | 2008-06-17 | #NAME? | 0 | 0 | 7.7943 | 91042 | 31.4207 | 44.878 | 235.214 | 756.569 | 118.394 | -2.9911 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.32648 | 582.751 | 36.7803 | 40
39616.366 | 2008-06-17 | #NAME? | 0 | 0 | 7.798 | 91031.6 | 31.5401 | 45.2575 | 233.639 | 756.831 | 120.757 | -3.0098 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4294 | 0.32216 | 582.751 | 36.7803 | 40
39616.3667 | 2008-06-17 | #NAME? | 0 | 0 | 7.7841 | 91022 | 31.5879 | 45.0945 | 232.589 | 759.457 | 120.232 | -3.0098 | -16.7619 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32318 | 583.121 | 36.7803 | 40
39616.3674 | 2008-06-17 | #NAME? | 0 | 0 | 7.7985 | 91017.6 | 31.564 | 45.2249 | 231.013 | 762.344 | 119.182 | -2.9911 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32674 | 582.751 | 36.7803 | 40
39616.3681 | 2008-06-17 | #NAME? | 0 | 0 | 7.8043 | 91021 | 31.3968 | 45.1271 | 229.963 | 761.557 | 121.545 | -3.0474 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32572 | 582.751 | 36.7803 | 40
39616.3688 | 2008-06-17 | #NAME? | 0 | 0 | 7.8844 | 91031.4 | 31.3013 | 44.9375 | 229.176 | 758.144 | 119.182 | -3.0098 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32496 | 582.751 | 36.7803 | 40
39616.3694 | 2008-06-17 | #NAME? | 0 | 0 | 7.8951 | 91043.1 | 31.1103 | 45.0619 | 228.126 | 756.044 | 119.969 | -3.0098 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32292 | 582.751 | 36.7803 | 40
39616.3701 | 2008-06-17 | #NAME? | 0 | 0 | 7.8903 | 91054.7 | 30.9432 | 44.9375 | 231.538 | 757.619 | 121.545 | -3.0098 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32318 | 583.121 | 37.0818 | 40
39616.3708 | 2008-06-17 | #NAME? | 0 | 0 | 7.8812 | 91065.3 | 30.8238 | 45.2249 | 236.526 | 760.244 | 118.657 | -3.0098 | -16.7285 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.3247 | 582.751 | 36.7803 | 40
39616.3715 | 2008-06-17 | #NAME? | 0 | 0 | 7.8487 | 91078.7 | 30.8715 | 44.878 | 237.051 | 762.607 | 121.02 | -3.0192 | -16.7285 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32725 | 582.751 | 36.7803 | 40
39616.3722 | 2008-06-17 | #NAME? | 0 | 0 | 7.7666 | 91078.8 | 30.9909 | 45.3553 | 236.001 | 760.507 | 119.969 | -3.0568 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32292 | 582.751 | 36.7803 | 40
39616.3729 | 2008-06-17 | #NAME? | 0 | 0 | 7.738 | 91069.8 | 31.0626 | 44.9673 | 234.689 | 756.569 | 119.444 | -3.0098 | -16.7285 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32318 | 582.751 | 36.7803 | 40
39616.3736 | 2008-06-17 | #NAME? | 0 | 0 | 6.3192 | 91062.6 | 31.1581 | 45.0293 | 233.114 | 756.569 | 121.545 | -3.0004 | -16.7553 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.33819 | 582.751 | 37.3833 | 40
39616.3743 | 2008-06-17 | #NAME? | 0 | 0 | 0.67644 | 91051.2 | 31.2297 | 44.997 | 232.063 | 758.669 | 118.919 | -3.0192 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.33946 | 582.751 | 37.6848 | 40
39616.375 | 2008-06-17 | #NAME? | 0 | 0 | 0.57141 | 91040.2 | 31.3252 | 44.9375 | 230.751 | 761.557 | 119.969 | -3.0098 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32674 | 582.751 | 36.7803 | 40
39616.3757 | 2008-06-17 | #NAME? | 0 | 0 | 6.2563 | 91027.6 | 31.4207 | 45.5834 | 229.701 | 762.082 | 121.282 | -3.038 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32216 | 582.751 | 36.7803 | 40
39616.3764 | 2008-06-17 | #NAME? | 0 | 0 | 7.9879 | 91019.1 | 31.4924 | 45.2575 | 228.651 | 758.932 | 118.394 | -3.0004 | -16.7285 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32623 | 582.751 | 36.7803 | 40
39616.3771 | 2008-06-17 | #NAME? | 0 | 0 | 8.0812 | 91012.6 | 31.5162 | 45.2575 | 229.176 | 756.569 | 121.282 | -3.038 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32394 | 583.121 | 36.7803 | 40
39616.3778 | 2008-06-17 | #NAME? | 0 | 0 | 8.0781 | 91009.1 | 31.4685 | 45.4205 | 234.164 | 757.619 | 119.707 | -3.0192 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32394 | 583.121 | 36.7803 | 40
39616.3785 | 2008-06-17 | #NAME? | 0 | 0 | 8.0685 | 91012 | 31.3013 | 44.8483 | 237.051 | 760.507 | 119.707 | -3.0098 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32318 | 583.121 | 36.7803 | 40
39616.3792 | 2008-06-17 | #NAME? | 0 | 0 | 8.0588 | 91024.9 | 31.1342 | 45.0619 | 236.789 | 762.607 | 121.545 | -3.0098 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32547 | 582.751 | 36.7803 | 40
39616.3799 | 2008-06-17 | #NAME? | 0 | 0 | 7.997 | 91037.7 | 30.9909 | 45.1923 | 235.739 | 760.769 | 118.394 | -3.038 | -16.7619 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32725 | 583.121 | 36.7803 | 40
39616.3806 | 2008-06-17 | #NAME? | 0 | 0 | 7.9978 | 91051.7 | 30.8238 | 45.0945 | 234.164 | 757.094 | 120.495 | -3.0286 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.3275 | 582.751 | 36.7803 | 40
39616.3812 | 2008-06-17 | #NAME? | 0 | 0 | 7.986 | 91062.2 | 30.8715 | 45.2575 | 232.851 | 756.831 | 120.495 | -3.0192 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32343 | 583.121 | 36.7803 | 40
39616.3819 | 2008-06-17 | #NAME? | 0 | 0 | 7.9811 | 91061.8 | 30.967 | 44.9673 | 231.538 | 758.932 | 118.919 | -3.0286 | -16.7553 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32598 | 582.751 | 37.0818 | 40
39616.3826 | 2008-06-17 | #NAME? | 0 | 0 | 7.9906 | 91054.3 | 31.0387 | 45.0619 | 230.488 | 762.082 | 121.545 | -3.038 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32394 | 582.751 | 36.7803 | 40
39616.3833 | 2008-06-17 | #NAME? | 0 | 0 | 7.9882 | 91045.5 | 31.1103 | 44.9078 | 229.176 | 762.082 | 118.919 | -2.9911 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32496 | 582.751 | 36.7803 | 40
39616.384 | 2008-06-17 | #NAME? | 0 | 0 | 8.0026 | 91035.9 | 31.2775 | 45.3227 | 228.388 | 758.407 | 120.232 | -3.0004 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32648 | 583.121 | 36.7803 | 40
39616.3847 | 2008-06-17 | #NAME? | 0 | 0 | 7.9953 | 91026.6 | 31.3252 | 45.0945 | 230.488 | 756.306 | 121.282 | -3.0286 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32241 | 583.121 | 36.7803 | 40
39616.3854 | 2008-06-17 | #NAME? | 0 | 0 | 8.0051 | 91018.6 | 31.4446 | 44.9375 | 236.001 | 757.619 | 118.657 | -3.0286 | -16.7285 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32674 | 582.751 | 36.7803 | 40
39616.3861 | 2008-06-17 | #NAME? | 0 | 0 | 8.0038 | 91006.6 | 31.3968 | 45.3879 | 237.314 | 761.032 | 121.02 | -3.038 | -16.7553 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32292 | 583.121 | 36.7803 | 40
39616.3868 | 2008-06-17 | #NAME? | 0 | 0 | 8.0055 | 91007.3 | 31.3013 | 44.997 | 236.526 | 762.344 | 119.444 | -3.0004 | -16.7553 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32292 | 582.751 | 36.7803 | 40
39616.3875 | 2008-06-17 | #NAME? | 0 | 0 | 7.9992 | 91011.6 | 31.1819 | 45.0945 | 235.476 | 759.982 | 119.707 | -3.038 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32648 | 582.751 | 36.7803 | 40
39616.3882 | 2008-06-17 | #NAME? | 0 | 0 | 7.9812 | 91022.7 | 31.0387 | 44.9078 | 233.901 | 756.569 | 121.545 | -3.0098 | -16.7619 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32496 | 582.751 | 36.7803 | 40
39616.3889 | 2008-06-17 | #NAME? | 0 | 0 | 7.9807 | 91034.5 | 30.9193 | 45.3227 | 232.589 | 756.569 | 118.394 | -3.0192 | -16.7619 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32343 | 583.121 | 36.7803 | 40
39616.3896 | 2008-06-17 | #NAME? | 0 | 0 | 7.9617 | 91041.9 | 30.9432 | 45.3227 | 231.276 | 758.669 | 120.495 | -3.0192 | -16.7285 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.3275 | 582.751 | 36.7803 | 40
39616.3903 | 2008-06-17 | #NAME? | 0 | 0 | 7.9685 | 91043.9 | 30.8954 | 44.9078 | 230.226 | 761.819 | 120.757 | -3.0192 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32445 | 583.121 | 36.7803 | 40
39616.391 | 2008-06-17 | #NAME? | 0 | 0 | 7.9606 | 91040.7 | 31.1103 | 45.1597 | 229.176 | 762.082 | 118.657 | -3.038 | -16.7285 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32292 | 583.121 | 36.7803 | 40
39616.3917 | 2008-06-17 | #NAME? | 0 | 0 | 7.9374 | 91028.7 | 31.1581 | 44.9375 | 228.126 | 758.932 | 121.02 | -3.0192 | -16.7152 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32419 | 582.751 | 36.7803 | 40
39616.3924 | 2008-06-17 | #NAME? | 0 | 0 | 7.9527 | 91024.8 | 31.3252 | 45.0945 | 231.013 | 756.569 | 119.969 | -3.0474 | -16.7686 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32521 | 583.121 | 37.0818 | 40
39616.3931 | 2008-06-17 | #NAME? | 0 | 0 | 7.9408 | 91013.2 | 31.3968 | 45.1923 | 236.001 | 757.094 | 119.707 | -3.0286 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32547 | 582.751 | 36.7803 | 40
39616.3938 | 2008-06-17 | #NAME? | 0 | 0 | 7.9559 | 90999.2 | 31.4924 | 45.1923 | 237.051 | 759.719 | 121.545 | -3.0192 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.3247 | 583.121 | 36.7803 | 40
39616.3944 | 2008-06-17 | #NAME? | 0 | 0 | 7.9434 | 90990.7 | 31.4446 | 45.4531 | 236.001 | 762.344 | 118.657 | -3.0004 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32598 | 582.751 | 36.7803 | 40
39616.3951 | 2008-06-17 | #NAME? | 0 | 0 | 7.9419 | 90994.6 | 31.2775 | 44.8483 | 234.689 | 761.032 | 120.495 | -2.9911 | -16.7619 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32318 | 582.751 | 36.7803 | 40
39616.3958 | 2008-06-17 | #NAME? | 0 | 0 | 7.9877 | 91003.2 | 31.2058 | 45.3553 | 233.376 | 757.094 | 120.757 | -3.0004 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32674 | 582.751 | 36.7803 | 40
39616.3965 | 2008-06-17 | #NAME? | 0 | 0 | 7.9748 | 91015.5 | 31.0148 | 44.9078 | 232.063 | 756.569 | 118.919 | -3.0286 | -16.7285 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32547 | 583.121 | 36.7803 | 40
39616.3972 | 2008-06-17 | #NAME? | 0 | 0 | 7.9757 | 91029.7 | 30.8715 | 45.1923 | 230.751 | 758.407 | 121.282 | -3.0098 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32419 | 583.121 | 36.7803 | 40
39616.3979 | 2008-06-17 | #NAME? | 0 | 0 | 7.9508 | 91034 | 31.0387 | 45.0293 | 229.701 | 761.557 | 119.444 | -3.038 | -16.7553 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32572 | 582.751 | 36.7803 | 40
39616.3986 | 2008-06-17 | #NAME? | 0 | 0 | 7.9698 | 91027.4 | 31.0864 | 44.8483 | 228.651 | 762.344 | 119.969 | -3.0098 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32598 | 582.751 | 36.7803 | 40
39616.3993 | 2008-06-17 | #NAME? | 0 | 0 | 7.9745 | 91025 | 31.0626 | 45.0619 | 228.913 | 758.669 | 121.545 | -3.038 | -16.7619 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32419 | 582.751 | 36.7803 | 40
39616.4 | 2008-06-17 | #NAME? | 0 | 0 | 7.9411 | 91022.7 | 31.0626 | 44.9078 | 233.901 | 756.569 | 118.657 | -3.0098 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32623 | 582.751 | 36.7803 | 40
39616.4007 | 2008-06-17 | #NAME? | 0 | 0 | 7.8955 | 91018.2 | 31.1581 | 45.1271 | 237.051 | 758.144 | 120.495 | -3.0004 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32674 | 582.751 | 36.7803 | 40
39616.4014 | 2008-06-17 | #NAME? | 0 | 0 | 7.8796 | 91011 | 31.1581 | 45.0293 | 237.051 | 761.294 | 120.757 | -3.0192 | -16.7553 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32699 | 582.751 | 36.7803 | 40
39616.4021 | 2008-06-17 | #NAME? | 0 | 0 | 7.8989 | 91007.3 | 31.1581 | 45.0293 | 235.739 | 762.344 | 118.919 | -2.9911 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32445 | 583.121 | 36.7803 | 40
39616.4028 | 2008-06-17 | #NAME? | 0 | 0 | 7.9054 | 91011.6 | 30.8238 | 45.2249 | 234.426 | 759.719 | 121.282 | -3.0286 | -16.7553 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32394 | 582.751 | 36.7803 | 40
39616.4035 | 2008-06-17 | #NAME? | 0 | 0 | 7.8882 | 91024.3 | 30.7044 | 44.8483 | 233.114 | 756.831 | 119.182 | -3.0004 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32369 | 582.751 | 36.7803 | 40
39616.4042 | 2008-06-17 | #NAME? | 0 | 0 | 7.8946 | 91035.3 | 30.2985 | 45.2575 | 231.801 | 756.569 | 119.969 | -3.0192 | -16.7553 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32725 | 582.751 | 36.7803 | 40
39616.4049 | 2008-06-17 | #NAME? | 0 | 0 | 7.874 | 91059.1 | 30.2507 | 44.878 | 230.488 | 758.669 | 121.545 | -3.0004 | -16.7553 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32343 | 582.751 | 36.7803 | 40
39616.4056 | 2008-06-17 | #NAME? | 0 | 0 | 7.8694 | 91071.7 | 30.1075 | 45.2575 | 229.438 | 761.819 | 118.657 | -3.0098 | -16.7619 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32547 | 582.751 | 36.7803 | 40
39616.4062 | 2008-06-17 | #NAME? | 0 | 0 | 7.831 | 91078.2 | 29.9433 | 45.0293 | 228.388 | 762.344 | 120.757 | -2.9911 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.3275 | 582.751 | 36.7803 | 40
39616.4069 | 2008-06-17 | #NAME? | 0 | 0 | 7.8284 | 91090.6 | 29.8526 | 45.0945 | 230.226 | 759.194 | 120.232 | -3.0098 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32801 | 582.751 | 36.7803 | 40
39616.4076 | 2008-06-17 | #NAME? | 0 | 0 | 7.8224 | 91095.1 | 29.8753 | 44.997 | 235.214 | 756.306 | 119.182 | -3.0098 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32674 | 582.751 | 36.7803 | 40
39616.4083 | 2008-06-17 | #NAME? | 0 | 0 | 7.8967 | 91094.7 | 29.8299 | 44.878 | 237.314 | 756.831 | 121.545 | -3.0286 | -16.7553 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32776 | 583.121 | 36.7803 | 40
39616.409 | 2008-06-17 | #NAME? | 0 | 0 | 7.9467 | 91087.2 | 29.898 | 45.0945 | 236.526 | 759.719 | 119.182 | -3.0004 | -16.7553 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32343 | 583.121 | 36.7803 | 40
39616.4097 | 2008-06-17 | #NAME? | 0 | 0 | 7.9421 | 91079.8 | 29.898 | 44.9078 | 235.214 | 762.607 | 119.707 | -3.0004 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.3275 | 582.751 | 36.7803 | 40
39616.4104 | 2008-06-17 | #NAME? | 0 | 0 | 7.9661 | 91071.8 | 30.0119 | 45.1923 | 233.901 | 761.032 | 121.545 | -2.9817 | -16.7619 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32369 | 583.121 | 36.7803 | 40
39616.4111 | 2008-06-17 | #NAME? | 0 | 0 | 7.9646 | 91058.6 | 30.0597 | 44.997 | 232.851 | 757.094 | 118.657 | -3.038 | -16.7619 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32699 | 582.751 | 36.7803 | 40
39616.4118 | 2008-06-17 | #NAME? | 0 | 0 | 7.9825 | 91048.2 | 30.1313 | 45.0293 | 231.276 | 756.569 | 120.232 | -3.0474 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32445 | 582.751 | 36.7803 | 40
39616.4125 | 2008-06-17 | #NAME? | 0 | 0 | 7.9931 | 91040.5 | 30.0119 | 45.2575 | 230.488 | 758.144 | 120.757 | -3.038 | -16.7152 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32292 | 582.751 | 36.7803 | 40
39616.4132 | 2008-06-17 | #NAME? | 0 | 0 | 7.9846 | 91039.3 | 29.9433 | 45.1597 | 229.176 | 761.557 | 118.657 | -3.0192 | -16.7419 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32369 | 582.751 | 36.7803 | 40
39616.4139 | 2008-06-17 | #NAME? | 0 | 0 | 7.9736 | 91043 | 29.8526 | 45.2249 | 228.126 | 762.344 | 121.282 | -3.0192 | -16.7486 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32674 | 583.121 | 36.7803 | 40
39616.4146 | 2008-06-17 | #NAME? | 0 | 0 | 7.98 | 91052.9 | 29.7392 | 45.0293 | 231.276 | 758.669 | 119.444 | -3.0004 | -16.7352 | 1.5059 | 50.9565 | 15.1743 | 1.4398 | 0.32394 | 582.751 | 36.7803 | 40
39616.4153 | 2008-06-17 | #NAME? | 0 | 0 | 7.9694 | 91057.6 | 29.5805 | 45.2901 | 236.526 | 756.569 | 119.444 | -3.0192 | -16.7419 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32598 | 582.751 | 36.7803 | 40
39616.416 | 2008-06-17 | #NAME? | 0 | 0 | 7.8889 | 91064.1 | 29.4444 | 44.9375 | 237.314 | 757.619 | 121.807 | -2.9817 | -16.7553 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.3275 | 582.381 | 36.7803 | 40
39616.4167 | 2008-06-17 | #NAME? | 0 | 0 | 7.9364 | 91074.5 | 29.3311 | 45.3879 | 236.264 | 760.244 | 118.657 | -3.0192 | -16.7218 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.3275 | 583.121 | 36.7803 | 40
39616.4174 | 2008-06-17 | #NAME? | 0 | 0 | 7.9112 | 91078.3 | 29.3764 | 44.9673 | 234.951 | 762.607 | 120.232 | -3.0192 | -16.7352 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32598 | 582.381 | 36.4789 | 40
39616.4181 | 2008-06-17 | #NAME? | 0 | 0 | 7.7749 | 91073.9 | 29.4671 | 45.1597 | 233.376 | 760.507 | 121.02 | -3.0004 | -16.7686 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32318 | 583.121 | 36.7803 | 40
39616.4187 | 2008-06-17 | #NAME? | 0 | 0 | 7.6722 | 91062.5 | 29.6032 | 45.3879 | 232.063 | 756.831 | 118.657 | -3.0286 | -16.7486 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32674 | 583.121 | 36.4789 | 40
39616.4194 | 2008-06-17 | #NAME? | 0 | 0 | 7.7034 | 91048.7 | 29.6939 | 44.9078 | 231.013 | 757.094 | 121.02 | -3.0192 | -16.7619 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32598 | 583.121 | 36.7803 | 40
39616.4201 | 2008-06-17 | #NAME? | 0 | 0 | 7.7151 | 91032.7 | 29.8073 | 45.4531 | 229.963 | 758.932 | 119.969 | -3.0098 | -16.7352 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32598 | 582.751 | 36.4789 | 40
39616.4208 | 2008-06-17 | #NAME? | 0 | 0 | 7.717 | 91017.4 | 29.898 | 44.9375 | 228.913 | 762.344 | 119.182 | -3.0004 | -16.7619 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32267 | 582.751 | 36.7803 | 40
39616.4215 | 2008-06-17 | #NAME? | 0 | 0 | 7.7483 | 91004.6 | 30.0119 | 45.0945 | 228.388 | 762.082 | 121.282 | -3.0192 | -16.7352 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32572 | 582.751 | 36.4789 | 40
39616.4222 | 2008-06-17 | #NAME? | 0 | 0 | 7.7487 | 90992.6 | 29.9887 | 45.3553 | 232.589 | 758.669 | 119.182 | -3.0004 | -16.7486 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32547 | 582.751 | 36.7803 | 40
39616.4229 | 2008-06-17 | #NAME? | 0 | 0 | 7.7615 | 90989.1 | 29.9433 | 44.9375 | 236.789 | 756.306 | 119.969 | -3.0098 | -16.7619 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32267 | 582.751 | 36.7803 | 40
39616.4236 | 2008-06-17 | #NAME? | 0 | 0 | 7.7585 | 90989.2 | 29.9433 | 45.4531 | 236.789 | 756.831 | 121.545 | -3.038 | -16.7553 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32674 | 583.121 | 36.7803 | 40
39616.4243 | 2008-06-17 | #NAME? | 0 | 0 | 7.7448 | 90989.1 | 29.8526 | 44.8483 | 235.739 | 759.719 | 118.394 | -3.0192 | -16.7619 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32445 | 582.751 | 36.7803 | 40
39616.425 | 2008-06-17 | #NAME? | 0 | 0 | 7.7406 | 90994.8 | 29.7846 | 45.2901 | 234.426 | 762.607 | 120.495 | -3.0286 | -16.7352 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32267 | 582.751 | 36.4789 | 40
39616.4257 | 2008-06-17 | #NAME? | 0 | 0 | 7.7228 | 90999.5 | 29.7846 | 45.0293 | 233.114 | 761.032 | 120.495 | -3.0004 | -16.7619 | 1.5059 | 50.9904 | 15.1743 | 1.4398 | 0.32521 | 582.751 | 36.4789 | 40
39616.4264 | 2008-06-17 | #NAME? | 0 | 0 | 7.7106 | 90999.5 | 29.8299 | 45.3227 | 231.538 | 757.357 | 118.919 | -3.038 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.3247 | 582.751 | 36.4789 | 40
39616.4271 | 2008-06-17 | #NAME? | 0 | 0 | 7.6997 | 90991.5 | 29.898 | 45.0293 | 230.488 | 756.306 | 121.545 | -3.0286 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32241 | 582.751 | 36.4789 | 40
39616.4278 | 2008-06-17 | #NAME? | 0 | 0 | 7.7036 | 90986.4 | 29.966 | 45.1271 | 229.438 | 757.882 | 119.444 | -3.0004 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32623 | 583.121 | 36.7803 | 40
39616.4285 | 2008-06-17 | #NAME? | 0 | 0 | 7.6908 | 90979 | 30.0836 | 45.3227 | 228.388 | 761.032 | 119.707 | -3.0286 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32419 | 582.751 | 36.7803 | 40
39616.4292 | 2008-06-17 | #NAME? | 0 | 0 | 7.6873 | 90971.9 | 30.0836 | 44.9673 | 229.963 | 762.082 | 121.545 | -2.9911 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32648 | 583.121 | 36.7803 | 40
39616.4299 | 2008-06-17 | #NAME? | 0 | 0 | 7.6958 | 90966.7 | 30.1791 | 45.4531 | 235.476 | 759.457 | 118.394 | -3.0474 | -16.7285 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32572 | 582.751 | 36.4789 | 40
39616.4306 | 2008-06-17 | #NAME? | 0 | 0 | 7.6877 | 90959.9 | 30.2268 | 44.9375 | 237.576 | 756.569 | 120.757 | -3.0004 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32216 | 583.121 | 36.7803 | 40
39616.4313 | 2008-06-17 | #NAME? | 0 | 0 | 7.7044 | 90951.7 | 30.2746 | 44.9673 | 236.789 | 757.094 | 120.232 | -3.0192 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.3247 | 582.751 | 36.7803 | 40
39616.4319 | 2008-06-17 | #NAME? | 0 | 0 | 7.6994 | 90945.1 | 30.2746 | 45.2575 | 235.214 | 758.669 | 119.182 | -2.9911 | -16.7218 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.3219 | 582.751 | 36.7803 | 40
39616.4326 | 2008-06-17 | #NAME? | 0 | 0 | 7.7067 | 90944.7 | 30.1552 | 44.8483 | 233.901 | 761.557 | 121.545 | -3.038 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32241 | 582.751 | 36.7803 | 40
39616.4333 | 2008-06-17 | #NAME? | 0 | 0 | 7.7048 | 90947.7 | 30.203 | 45.1597 | 232.589 | 762.082 | 119.182 | -3.038 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32216 | 582.751 | 36.7803 | 40
39616.434 | 2008-06-17 | #NAME? | 0 | 0 | 7.6907 | 90950.6 | 30.0597 | 45.0945 | 231.276 | 758.669 | 119.969 | -3.038 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32343 | 583.121 | 36.7803 | 40
39616.4347 | 2008-06-17 | #NAME? | 0 | 0 | 7.7025 | 90960.3 | 29.9433 | 44.997 | 230.226 | 756.044 | 121.545 | -3.0192 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32394 | 582.751 | 36.4789 | 40
39616.4354 | 2008-06-17 | #NAME? | 0 | 0 | 7.6741 | 90966.7 | 29.8753 | 45.0293 | 229.176 | 757.094 | 118.394 | -3.0474 | -16.7619 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32623 | 583.121 | 36.7803 | 40
39616.4361 | 2008-06-17 | #NAME? | 0 | 0 | 7.656 | 90975.9 | 29.8073 | 44.9673 | 228.126 | 759.719 | 120.757 | -3.0286 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32674 | 583.121 | 36.7803 | 40
39616.4368 | 2008-06-17 | #NAME? | 0 | 0 | 7.6527 | 90984.1 | 29.7166 | 45.3879 | 231.276 | 762.344 | 119.969 | -3.0098 | -16.7619 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32648 | 582.751 | 36.7803 | 40
39616.4375 | 2008-06-17 | #NAME? | 0 | 0 | 7.6231 | 90985.7 | 29.7846 | 44.9673 | 236.001 | 760.769 | 119.182 | -3.038 | -16.7152 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32699 | 583.121 | 36.7803 | 40
39616.4382 | 2008-06-17 | #NAME? | 0 | 0 | 7.6133 | 90984.6 | 29.8073 | 45.0293 | 236.789 | 757.094 | 121.545 | -3.038 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32394 | 582.751 | 36.7803 | 40
39616.4389 | 2008-06-17 | #NAME? | 0 | 0 | 7.597 | 90978.4 | 29.898 | 45.1923 | 235.739 | 756.306 | 118.657 | -3.0286 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32241 | 582.751 | 36.7803 | 40
39616.4396 | 2008-06-17 | #NAME? | 0 | 0 | 7.5989 | 90973.6 | 29.966 | 45.3553 | 234.689 | 758.144 | 120.495 | -3.0004 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32547 | 583.121 | 36.7803 | 40
39616.4403 | 2008-06-17 | #NAME? | 0 | 0 | 7.5904 | 90965.2 | 30.0358 | 45.1597 | 233.114 | 761.032 | 120.757 | -3.0474 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32623 | 583.121 | 36.7803 | 40
39616.441 | 2008-06-17 | #NAME? | 0 | 0 | 7.6005 | 90958.3 | 30.1075 | 45.1597 | 231.801 | 762.607 | 118.919 | -2.9817 | -16.7285 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32648 | 582.751 | 36.7803 | 40
39616.4417 | 2008-06-17 | #NAME? | 0 | 0 | 7.5986 | 90947.8 | 30.1791 | 45.3879 | 230.751 | 759.982 | 121.545 | -2.9911 | -16.7152 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32267 | 583.121 | 36.7803 | 40
39616.4424 | 2008-06-17 | #NAME? | 0 | 0 | 7.6077 | 90940.2 | 30.2268 | 44.9673 | 229.701 | 756.831 | 119.444 | -3.0474 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32648 | 582.751 | 36.7803 | 40
39616.4431 | 2008-06-17 | #NAME? | 0 | 0 | 7.62 | 90934 | 30.3462 | 45.1923 | 228.651 | 756.569 | 119.707 | -2.9817 | -16.7285 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.3219 | 582.751 | 36.7803 | 40
39616.4437 | 2008-06-17 | #NAME? | 0 | 0 | 7.6206 | 90923.1 | 30.3223 | 45.1271 | 229.176 | 758.669 | 121.807 | -3.038 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.3219 | 583.121 | 36.7803 | 40
39616.4444 | 2008-06-17 | #NAME? | 0 | 0 | 7.6352 | 90918.8 | 30.2985 | 45.1923 | 234.164 | 762.082 | 118.394 | -3.0192 | -16.7686 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32241 | 583.121 | 36.7803 | 40
39616.4451 | 2008-06-17 | #NAME? | 0 | 0 | 7.627 | 90919.2 | 30.2268 | 45.2901 | 237.314 | 762.344 | 121.02 | -3.0004 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32241 | 583.121 | 36.7803 | 40
39616.4458 | 2008-06-17 | #NAME? | 0 | 0 | 7.625 | 90925.7 | 30.1075 | 44.878 | 237.051 | 759.719 | 119.969 | -3.038 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32623 | 583.121 | 36.7803 | 40
39616.4465 | 2008-06-17 | #NAME? | 0 | 0 | 7.6094 | 90929 | 30.1552 | 45.4531 | 235.739 | 756.831 | 119.444 | -3.0192 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32241 | 582.751 | 36.7803 | 40
39616.4472 | 2008-06-17 | #NAME? | 0 | 0 | 7.5967 | 90931.3 | 30.0119 | 44.878 | 234.164 | 756.831 | 121.282 | -3.038 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32419 | 583.121 | 36.7803 | 40
39616.4479 | 2008-06-17 | #NAME? | 0 | 0 | 7.5986 | 90938.8 | 29.9887 | 45.3879 | 232.851 | 758.932 | 119.182 | -3.0192 | -16.7285 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32369 | 583.121 | 36.7803 | 40
39616.4486 | 2008-06-17 | #NAME? | 0 | 0 | 7.573 | 90945.6 | 29.8526 | 45.3227 | 231.538 | 761.819 | 119.969 | -3.0192 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32496 | 583.121 | 36.7803 | 40
39616.4493 | 2008-06-17 | #NAME? | 0 | 0 | 7.5651 | 90955.2 | 29.898 | 44.997 | 230.488 | 761.819 | 121.282 | -2.9817 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32572 | 582.751 | 36.7803 | 40
39616.45 | 2008-06-17 | #NAME? | 0 | 0 | 7.5614 | 90955.4 | 29.9433 | 45.4205 | 229.701 | 758.407 | 118.657 | -3.0192 | -16.7218 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32241 | 582.751 | 36.4789 | 40
39616.4507 | 2008-06-17 | #NAME? | 0 | 0 | 7.5474 | 90948.2 | 30.0358 | 44.8483 | 228.388 | 756.306 | 121.02 | -3.0004 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32496 | 583.121 | 36.7803 | 40
39616.4514 | 2008-06-17 | #NAME? | 0 | 0 | 7.5447 | 90940.4 | 30.1075 | 45.3879 | 229.963 | 757.619 | 119.969 | -3.0098 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32241 | 583.121 | 36.7803 | 40
39616.4521 | 2008-06-17 | #NAME? | 0 | 0 | 7.5325 | 90932 | 30.1791 | 44.997 | 234.951 | 760.507 | 119.444 | -3.0286 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32343 | 583.121 | 36.7803 | 40
39616.4528 | 2008-06-17 | #NAME? | 0 | 0 | 7.5543 | 90924.8 | 30.2507 | 44.9673 | 237.576 | 762.607 | 121.807 | -3.038 | -16.7285 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32343 | 582.751 | 36.7803 | 40
39616.4535 | 2008-06-17 | #NAME? | 0 | 0 | 7.5533 | 90914.5 | 30.3462 | 45.1597 | 236.789 | 760.769 | 118.919 | -3.0004 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32216 | 582.751 | 36.7803 | 40
39616.4542 | 2008-06-17 | #NAME? | 0 | 0 | 7.5622 | 90906.6 | 30.3462 | 44.8483 | 235.476 | 756.831 | 120.232 | -3.0098 | -16.7619 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32419 | 583.121 | 36.7803 | 40
39616.4549 | 2008-06-17 | #NAME? | 0 | 0 | 7.5608 | 90904.1 | 30.2507 | 45.4857 | 234.164 | 756.569 | 121.02 | -3.0286 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32496 | 583.121 | 36.7803 | 40
39616.4556 | 2008-06-17 | #NAME? | 0 | 0 | 7.5595 | 90905.7 | 30.2268 | 45.0293 | 232.851 | 758.407 | 118.657 | -2.9911 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32419 | 582.751 | 36.7803 | 40
39616.4563 | 2008-06-17 | #NAME? | 0 | 0 | 7.5646 | 90911.4 | 30.203 | 44.997 | 231.538 | 761.032 | 120.757 | -3.0098 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32267 | 583.121 | 36.7803 | 40
39616.4569 | 2008-06-17 | #NAME? | 0 | 0 | 7.5378 | 90915.7 | 30.1313 | 45.2249 | 230.488 | 762.607 | 119.969 | -3.0004 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32445 | 583.121 | 36.7803 | 40
39616.4576 | 2008-06-17 | #NAME? | 0 | 0 | 7.5338 | 90918.9 | 30.0358 | 44.8483 | 229.176 | 759.982 | 119.444 | -3.0098 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32699 | 582.751 | 36.4789 | 40
39616.4583 | 2008-06-17 | #NAME? | 0 | 0 | 7.5262 | 90923.4 | 29.9887 | 45.4857 | 228.388 | 756.569 | 121.807 | -3.0192 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32343 | 583.121 | 36.7803 | 40
39616.459 | 2008-06-17 | #NAME? | 0 | 0 | 7.5059 | 90927.4 | 29.8299 | 45.0293 | 229.701 | 756.831 | 118.657 | -3.0192 | -16.7285 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32699 | 582.751 | 36.7803 | 40
39616.4597 | 2008-06-17 | #NAME? | 0 | 0 | 7.5082 | 90932.6 | 29.8299 | 45.3227 | 234.689 | 759.457 | 120.232 | -2.9817 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32674 | 582.751 | 36.7803 | 40
39616.4604 | 2008-06-17 | #NAME? | 0 | 0 | 7.4841 | 90935.7 | 29.8073 | 44.9375 | 237.314 | 762.344 | 121.02 | -3.0098 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32267 | 582.751 | 36.7803 | 40
39616.4611 | 2008-06-17 | #NAME? | 0 | 0 | 7.4908 | 90935.7 | 29.9433 | 45.2575 | 236.526 | 761.557 | 118.657 | -2.9911 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32292 | 583.121 | 36.7803 | 40
39616.4618 | 2008-06-17 | #NAME? | 0 | 0 | 7.4821 | 90926.7 | 30.0119 | 45.4205 | 235.214 | 757.882 | 121.02 | -3.038 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32241 | 582.751 | 36.7803 | 40
39616.4625 | 2008-06-17 | #NAME? | 0 | 0 | 7.4793 | 90919.3 | 30.0597 | 44.9375 | 234.164 | 756.306 | 119.969 | -3.0098 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32292 | 583.121 | 36.7803 | 40
39616.4632 | 2008-06-17 | #NAME? | 0 | 0 | 7.482 | 90910.8 | 30.203 | 45.4531 | 232.589 | 757.619 | 119.182 | -3.038 | -16.7285 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32241 | 583.121 | 36.7803 | 40
39616.4639 | 2008-06-17 | #NAME? | 0 | 0 | 7.479 | 90899.2 | 30.2746 | 45.2249 | 231.538 | 760.769 | 121.545 | -3.0098 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32318 | 582.751 | 36.7803 | 40
39616.4646 | 2008-06-17 | #NAME? | 0 | 0 | 7.4986 | 90888.5 | 30.3462 | 45.0945 | 230.488 | 762.607 | 118.919 | -3.0286 | -16.7218 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32547 | 583.121 | 36.7803 | 40
39616.4653 | 2008-06-17 | #NAME? | 0 | 0 | 7.4998 | 90880.2 | 30.3223 | 45.0945 | 229.176 | 760.507 | 120.232 | -3.0192 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32521 | 582.751 | 36.7803 | 40
39616.466 | 2008-06-17 | #NAME? | 0 | 0 | 7.5089 | 90879.8 | 30.2268 | 45.1271 | 228.388 | 756.831 | 121.02 | -3.0192 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32292 | 582.751 | 36.7803 | 40
39616.4667 | 2008-06-17 | #NAME? | 0 | 0 | 7.5032 | 90884.6 | 30.1075 | 45.2249 | 231.013 | 756.569 | 118.657 | -3.0192 | -16.7619 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32598 | 582.751 | 36.7803 | 40
39616.4674 | 2008-06-17 | #NAME? | 0 | 0 | 7.4873 | 90887.5 | 30.0597 | 44.9078 | 236.001 | 758.144 | 121.282 | -3.0098 | -16.7285 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32292 | 583.121 | 36.4789 | 40
39616.4681 | 2008-06-17 | #NAME? | 0 | 0 | 7.4946 | 90892.1 | 30.0119 | 45.3227 | 237.051 | 761.557 | 119.444 | -3.0098 | -16.7218 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32318 | 583.121 | 36.7803 | 40
39616.4688 | 2008-06-17 | #NAME? | 0 | 0 | 7.4716 | 90894.3 | 29.898 | 44.9673 | 236.001 | 762.082 | 119.444 | -3.0192 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32369 | 582.751 | 36.7803 | 40
39616.4694 | 2008-06-17 | #NAME? | 0 | 0 | 7.4614 | 90901.9 | 29.8753 | 45.1923 | 234.426 | 758.407 | 121.545 | -3.0474 | -16.7686 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32648 | 582.751 | 36.7803 | 40
39616.4701 | 2008-06-17 | #NAME? | 0 | 0 | 7.4491 | 90907.7 | 29.8073 | 45.2249 | 233.376 | 756.306 | 118.657 | -2.9911 | -16.7619 | 1.5059 | 51.0243 | 15.1743 | 1.4398 | 0.32318 | 583.121 | 36.7803 | 40
39616.4708 | 2008-06-17 | #NAME? | 0 | 0 | 7.4391 | 90912.5 | 29.6485 | 44.997 | 232.063 | 757.619 | 120.232 | -3.0004 | -16.7218 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32674 | 583.121 | 36.7803 | 40
39616.4715 | 2008-06-17 | #NAME? | 0 | 0 | 7.4295 | 90916.9 | 29.7846 | 45.2575 | 230.751 | 760.244 | 121.02 | -3.0474 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32267 | 583.121 | 36.7803 | 40
39616.4722 | 2008-06-17 | #NAME? | 0 | 0 | 7.4239 | 90907.2 | 29.8526 | 44.9673 | 229.701 | 762.607 | 118.657 | -3.0098 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.3247 | 582.751 | 36.7803 | 40
39616.4729 | 2008-06-17 | #NAME? | 0 | 0 | 7.423 | 90899.2 | 29.898 | 45.1923 | 228.651 | 760.244 | 121.282 | -3.0098 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32394 | 583.121 | 36.7803 | 40
39616.4736 | 2008-06-17 | #NAME? | 0 | 0 | 7.4197 | 90890 | 30.0358 | 44.9078 | 229.176 | 757.094 | 119.707 | -3.038 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32419 | 582.751 | 36.7803 | 40
39616.4743 | 2008-06-17 | #NAME? | 0 | 0 | 7.4299 | 90878.7 | 30.1075 | 45.4205 | 233.639 | 756.306 | 119.444 | -3.0192 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32648 | 583.121 | 36.7803 | 40
39616.475 | 2008-06-17 | #NAME? | 0 | 0 | 7.4305 | 90867 | 30.1791 | 45.1923 | 237.314 | 758.407 | 121.807 | -3.038 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32292 | 583.121 | 36.7803 | 40
39616.4757 | 2008-06-17 | #NAME? | 0 | 0 | 7.4525 | 90856.5 | 30.2746 | 44.9078 | 237.314 | 761.294 | 118.657 | -3.0286 | -16.7619 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32267 | 583.121 | 36.7803 | 40
39616.4764 | 2008-06-17 | #NAME? | 0 | 0 | 7.4532 | 90847.6 | 30.2507 | 45.3553 | 235.739 | 762.344 | 120.232 | -3.0286 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32267 | 583.121 | 36.7803 | 40
39616.4771 | 2008-06-17 | #NAME? | 0 | 0 | 7.4464 | 90846.8 | 30.1791 | 44.8186 | 234.426 | 759.982 | 120.757 | -3.038 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32267 | 583.121 | 36.7803 | 40
39616.4778 | 2008-06-17 | #NAME? | 0 | 0 | 7.4505 | 90850.7 | 30.0597 | 45.2249 | 233.114 | 756.831 | 118.919 | -3.038 | -16.7285 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32267 | 583.121 | 36.7803 | 40
39616.4785 | 2008-06-17 | #NAME? | 0 | 0 | 7.4424 | 90852.3 | 30.0597 | 44.9673 | 231.801 | 756.569 | 121.282 | -3.0098 | -16.7619 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32445 | 582.751 | 36.4789 | 40
39616.4792 | 2008-06-17 | #NAME? | 0 | 0 | 7.4362 | 90854 | 29.966 | 45.1923 | 230.751 | 758.407 | 119.707 | -3.0192 | -16.7285 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32267 | 582.751 | 36.7803 | 40
39616.4799 | 2008-06-17 | #NAME? | 0 | 0 | 7.4243 | 90857.5 | 29.898 | 44.9078 | 229.438 | 761.294 | 119.444 | -3.0004 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32267 | 583.121 | 36.7803 | 40
39616.4806 | 2008-06-17 | #NAME? | 0 | 0 | 7.4201 | 90858.3 | 29.9206 | 45.1597 | 228.388 | 762.344 | 121.545 | -2.9911 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32623 | 582.751 | 36.7803 | 40
39616.4812 | 2008-06-17 | #NAME? | 0 | 0 | 7.4025 | 90859 | 29.8526 | 45.3227 | 229.176 | 759.982 | 118.919 | -3.0192 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32445 | 582.751 | 36.7803 | 40
39616.4819 | 2008-06-17 | #NAME? | 0 | 0 | 7.3984 | 90863 | 29.7619 | 44.9673 | 234.164 | 756.831 | 120.232 | -3.0192 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32648 | 582.751 | 36.4789 | 40
39616.4826 | 2008-06-17 | #NAME? | 0 | 0 | 7.3833 | 90863.2 | 29.8753 | 45.6812 | 237.314 | 756.831 | 120.757 | -3.038 | -16.7218 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32241 | 583.121 | 36.7803 | 40
39616.4833 | 2008-06-17 | #NAME? | 0 | 0 | 7.3652 | 90855.2 | 29.9433 | 45.4531 | 236.789 | 758.669 | 118.657 | -3.0004 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32648 | 583.121 | 36.7803 | 40
39616.484 | 2008-06-17 | #NAME? | 0 | 0 | 7.3832 | 90847 | 29.9887 | 44.9673 | 235.476 | 762.082 | 121.282 | -3.038 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32343 | 582.751 | 36.7803 | 40
39616.4847 | 2008-06-17 | #NAME? | 0 | 0 | 7.3829 | 90840.1 | 30.1075 | 45.2901 | 234.164 | 761.819 | 119.444 | -3.0192 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32572 | 582.751 | 36.4789 | 40
39616.4854 | 2008-06-17 | #NAME? | 0 | 0 | 7.3803 | 90828.3 | 30.1552 | 44.9673 | 232.851 | 758.407 | 119.707 | -3.0286 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32369 | 583.121 | 36.7803 | 40
39616.4861 | 2008-06-17 | #NAME? | 0 | 0 | 7.4014 | 90816.8 | 30.2268 | 44.7591 | 231.538 | 756.306 | 121.545 | -3.0474 | -16.7285 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.3247 | 582.751 | 36.7803 | 40
39616.4868 | 2008-06-17 | #NAME? | 0 | 0 | 7.4065 | 90806.4 | 30.3701 | 45.0619 | 230.751 | 757.357 | 118.657 | -3.0004 | -16.7218 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32623 | 582.751 | 36.7803 | 40
39616.4875 | 2008-06-17 | #NAME? | 0 | 0 | 7.4043 | 90799.6 | 30.2746 | 45.0945 | 229.438 | 759.982 | 120.232 | -3.0192 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32343 | 582.751 | 36.7803 | 40
39616.4882 | 2008-06-17 | #NAME? | 0 | 0 | 7.4224 | 90799.2 | 30.203 | 44.878 | 228.651 | 762.344 | 120.757 | -2.9911 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32623 | 582.751 | 36.7803 | 40
39616.4889 | 2008-06-17 | #NAME? | 0 | 0 | 7.4113 | 90801.3 | 30.1791 | 45.616 | 229.176 | 759.719 | 118.657 | -3.038 | -16.7218 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32572 | 582.751 | 36.7803 | 40
39616.4896 | 2008-06-17 | #NAME? | 0 | 0 | 7.3929 | 90801.6 | 30.1791 | 45.3879 | 233.901 | 756.569 | 121.02 | -3.0098 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32674 | 583.121 | 36.4789 | 40
39616.4903 | 2008-06-17 | #NAME? | 0 | 0 | 7.4089 | 90803.7 | 30.0358 | 44.9673 | 237.051 | 756.569 | 119.707 | -3.0286 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32648 | 582.751 | 36.7803 | 40
39616.491 | 2008-06-17 | #NAME? | 0 | 0 | 7.395 | 90809 | 29.898 | 44.9673 | 236.264 | 758.407 | 119.444 | -3.0004 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.3247 | 582.751 | 36.4789 | 40
39616.4917 | 2008-06-17 | #NAME? | 0 | 0 | 7.3838 | 90816 | 29.8526 | 44.7293 | 234.951 | 760.769 | 121.545 | -3.0286 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32598 | 583.121 | 36.4789 | 40
39616.4924 | 2008-06-17 | #NAME? | 0 | 0 | 7.3838 | 90822 | 29.7619 | 44.9673 | 233.639 | 762.344 | 118.657 | -3.0098 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32318 | 583.121 | 36.7803 | 40
39616.4931 | 2008-06-17 | #NAME? | 0 | 0 | 7.3496 | 90824.8 | 29.8526 | 45.3553 | 232.326 | 760.507 | 120.495 | -3.0192 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32699 | 583.121 | 36.4789 | 40
39616.4938 | 2008-06-17 | #NAME? | 0 | 0 | 7.3353 | 90823.7 | 29.6259 | 45.1597 | 231.013 | 756.831 | 121.02 | -3.0098 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32445 | 582.751 | 36.4789 | 40
39616.4944 | 2008-06-17 | #NAME? | 0 | 0 | 7.3484 | 90833.1 | 29.6032 | 44.997 | 229.963 | 756.569 | 118.657 | -3.0192 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32369 | 583.121 | 36.7803 | 40
39616.4951 | 2008-06-17 | #NAME? | 0 | 0 | 7.3335 | 90835.6 | 29.7166 | 45.2575 | 228.913 | 758.932 | 121.02 | -3.0098 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32394 | 582.751 | 36.4789 | 40
39616.4958 | 2008-06-17 | #NAME? | 0 | 0 | 7.3114 | 90824.6 | 29.7392 | 45.2249 | 228.388 | 762.082 | 119.969 | -3.0192 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32267 | 583.121 | 36.7803 | 40
39616.4965 | 2008-06-17 | #NAME? | 0 | 0 | 7.3253 | 90819.6 | 29.8299 | 45.2575 | 232.326 | 762.082 | 119.444 | -3.0098 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32369 | 583.121 | 36.7803 | 40
39616.4972 | 2008-06-17 | #NAME? | 0 | 0 | 7.3206 | 90807 | 29.9433 | 45.1271 | 236.789 | 758.407 | 121.545 | -3.038 | -16.7619 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32648 | 583.121 | 36.4789 | 40
39616.4979 | 2008-06-17 | #NAME? | 0 | 0 | 7.3225 | 90793.3 | 30.0119 | 45.0293 | 237.051 | 756.306 | 119.182 | -3.0098 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32241 | 582.751 | 36.7803 | 40
39616.4986 | 2008-06-17 | #NAME? | 0 | 0 | 7.3428 | 90784.3 | 30.1075 | 44.9375 | 235.739 | 757.357 | 120.232 | -3.038 | -16.7619 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32674 | 582.751 | 36.7803 | 40
39616.4993 | 2008-06-17 | #NAME? | 0 | 0 | 7.3418 | 90770.2 | 30.2507 | 45.3227 | 234.426 | 760.244 | 121.282 | -3.0004 | -16.7218 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32521 | 583.121 | 36.4789 | 40
39616.5 | 2008-06-17 | #NAME? | 0 | 0 | 7.343 | 90754.9 | 30.2507 | 45.0619 | 233.114 | 762.869 | 118.394 | -2.9911 | -16.7218 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32216 | 582.751 | 36.4789 | 40
39616.5007 | 2008-06-17 | #NAME? | 0 | 0 | 7.366 | 90750.8 | 30.1075 | 44.8186 | 231.801 | 761.819 | 121.282 | -3.0004 | -16.7285 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32292 | 583.121 | 36.7803 | 40
39616.5014 | 2008-06-17 | #NAME? | 0 | 0 | 7.3681 | 90753.5 | 30.0358 | 45.0945 | 231.013 | 758.144 | 119.707 | -3.0098 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32241 | 583.121 | 36.7803 | 40
39616.5021 | 2008-06-17 | #NAME? | 0 | 0 | 7.3539 | 90754.1 | 30.0358 | 44.8483 | 229.438 | 756.306 | 119.707 | -3.0192 | -16.7285 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32292 | 583.121 | 36.7803 | 40
39616.5028 | 2008-06-17 | #NAME? | 0 | 0 | 7.3602 | 90755.9 | 29.966 | 45.0293 | 228.651 | 757.094 | 121.807 | -3.038 | -16.7419 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32699 | 583.121 | 36.7803 | 40
39616.5035 | 2008-06-17 | #NAME? | 0 | 0 | 7.3439 | 90760 | 29.9206 | 45.2901 | 228.913 | 760.244 | 118.394 | -3.0098 | -16.7553 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32318 | 582.751 | 36.7803 | 40
39616.5042 | 2008-06-17 | #NAME? | 0 | 0 | 7.3298 | 90764.2 | 29.7846 | 44.9673 | 233.114 | 762.607 | 121.02 | -3.0098 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4501 | 0.32648 | 582.751 | 36.4789 | 40
39616.5049 | 2008-06-17 | #NAME? | 0 | 0 | 7.3315 | 90772.8 | 29.7392 | 45.3879 | 237.051 | 759.982 | 120.232 | -2.9911 | -16.7352 | 1.5059 | 51.0243 | 15.1743 | 1.4605 | 0.32598 | 582.751 | 36.7803 | 40
39616.5056 | 2008-06-17 | #NAME? | 0 | 0 | 7.3086 | 90778.1 | 29.6485 | 44.8483 | 236.789 | 756.831 | 119.182 | -3.0004 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4605 | 0.3247 | 582.751 | 36.7803 | 40
39616.5062 | 2008-06-17 | #NAME? | 0 | 0 | 7.3102 | 90782.3 | 29.6032 | 45.1271 | 235.476 | 756.306 | 121.282 | -3.0192 | -16.7619 | 1.5059 | 51.0243 | 15.1743 | 1.4605 | 0.32343 | 583.121 | 36.7803 | 40
39616.5069 | 2008-06-17 | #NAME? | 0 | 0 | 7.2874 | 90784.8 | 29.4671 | 45.0293 | 234.164 | 758.144 | 119.444 | -3.0192 | -16.7486 | 1.5059 | 51.0243 | 15.1743 | 1.4605 | 0.32648 | 583.121 | 36.7803 | 40
39616.5076 | 2008-06-17 | #NAME? | 0 | 0 | 7.2771 | 90792 | 29.5805 | 44.878 | 232.851 | 761.294 | 119.444 | -3.0004 | -16.7553 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32623 | 583.121 | 36.7803 | 40
39616.5083 | 2008-06-17 | #NAME? | 0 | 0 | 7.288 | 90782.8 | 29.7166 | 45.2575 | 231.538 | 762.082 | 121.545 | -3.0098 | -16.7218 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32318 | 582.751 | 36.4789 | 40
39616.509 | 2008-06-17 | #NAME? | 0 | 0 | 7.2787 | 90766.8 | 29.8073 | 45.2249 | 230.488 | 758.932 | 118.657 | -3.0004 | -16.7419 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32496 | 583.121 | 36.4789 | 40
39616.5097 | 2008-06-17 | #NAME? | 0 | 0 | 7.2868 | 90754.4 | 29.8753 | 45.1271 | 229.438 | 756.306 | 120.757 | -3.0192 | -16.7419 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32547 | 582.751 | 36.4789 | 40
39616.5104 | 2008-06-17 | #NAME? | 0 | 0 | 7.2964 | 90741.1 | 29.966 | 44.997 | 228.651 | 756.831 | 120.232 | -3.0474 | -16.7419 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32598 | 583.121 | 36.4789 | 40
39616.5111 | 2008-06-17 | #NAME? | 0 | 0 | 7.3012 | 90727.8 | 30.1313 | 45.1597 | 230.226 | 758.669 | 119.182 | -3.0098 | -16.7419 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32241 | 582.751 | 36.7803 | 40
39616.5118 | 2008-06-17 | #NAME? | 0 | 0 | 7.7089 | 90710.9 | 30.2507 | 44.9673 | 235.476 | 762.082 | 121.545 | -3.0098 | -16.7619 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32114 | 583.121 | 36.4789 | 40
39616.5125 | 2008-06-17 | #NAME? | 0 | 0 | 17.1589 | 90699.6 | 30.2985 | 44.9078 | 237.051 | 761.557 | 119.444 | -3.0192 | -16.7619 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32114 | 583.121 | 36.4789 | 40
39616.5132 | 2008-06-17 | #NAME? | 0 | 0 | 10.2009 | 90690.8 | 30.1552 | 45.2901 | 236.264 | 757.882 | 119.707 | -3.0004 | -16.7619 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32114 | 582.751 | 36.7803 | 40
39616.5139 | 2008-06-17 | #NAME? | 0 | 0 | 9.0258 | 90694.5 | 30.203 | 44.9375 | 234.951 | 756.044 | 121.545 | -3.0286 | -16.7486 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32572 | 582.751 | 36.4789 | 40
39616.5146 | 2008-06-17 | #NAME? | 0 | 0 | 8.2412 | 90692.7 | 30.1075 | 45.1923 | 233.639 | 757.619 | 118.394 | -2.9911 | -16.7218 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32598 | 583.121 | 36.7803 | 40
39616.5153 | 2008-06-17 | #NAME? | 0 | 0 | 7.1112 | 90696.4 | 29.898 | 44.9078 | 232.326 | 759.719 | 120.757 | -3.0004 | -16.7553 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.3247 | 583.121 | 36.7803 | 40
39616.516 | 2008-06-17 | #NAME? | 0 | 0 | 3.4903 | 90704.2 | 29.8753 | 45.4857 | 231.013 | 762.607 | 120.232 | -3.0286 | -16.7419 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.3415 | 582.751 | 37.3833 | 40
39616.5167 | 2008-06-17 | #NAME? | 0 | 0 | 0.0099562 | 90712 | 29.6939 | 45.4205 | 229.963 | 761.032 | 119.182 | -2.9911 | -16.7753 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.33742 | 582.751 | 37.3833 | 40
39616.5174 | 2008-06-17 | #NAME? | 0 | 0 | -0.0022099 | 90721.3 | 29.4898 | 45.1923 | 228.651 | 757.357 | 121.545 | -3.0004 | -16.7218 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.3415 | 582.751 | 37.3833 | 40
39616.5181 | 2008-06-17 | #NAME? | 0 | 0 | -0.0015078 | 90734.1 | 29.3084 | 45.2901 | 228.651 | 756.044 | 119.182 | -2.9911 | -16.7352 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.3415 | 583.121 | 37.3833 | 40
39616.5188 | 2008-06-17 | #NAME? | 0 | 0 | -0.0030689 | 90742.5 | 29.1043 | 44.9375 | 232.851 | 757.882 | 119.969 | -2.9911 | -16.7619 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.34022 | 583.121 | 37.3833 | 40
39616.5194 | 2008-06-17 | #NAME? | 0 | 0 | -0.0019485 | 90757.5 | 28.9909 | 45.2901 | 236.789 | 760.244 | 121.282 | -3.0474 | -16.7352 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.33793 | 582.751 | 37.3833 | 40
39616.5201 | 2008-06-17 | #NAME? | 0 | 0 | 0.54076 | 90765 | 28.7415 | 45.0293 | 237.051 | 762.607 | 118.394 | -3.0286 | -16.7285 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32267 | 583.121 | 36.4789 | 40
39616.5208 | 2008-06-17 | #NAME? | 0 | 0 | 5.8181 | 90773.3 | 28.8095 | 45.2249 | 235.739 | 761.294 | 121.02 | -3.0192 | -16.7553 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32598 | 582.751 | 36.7803 | 40
39616.5215 | 2008-06-17 | #NAME? | 0 | 0 | 7.9336 | 90774.8 | 28.5828 | 44.9673 | 234.426 | 757.882 | 120.232 | -3.0098 | -16.7686 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32343 | 582.751 | 36.4789 | 40
39616.5222 | 2008-06-17 | #NAME? | 0 | 0 | 8.0496 | 90777.3 | 28.7642 | 44.9375 | 232.851 | 756.306 | 118.919 | -3.0286 | -16.7553 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32521 | 583.121 | 36.4789 | 40
39616.5229 | 2008-06-17 | #NAME? | 0 | 0 | 8.0431 | 90765.2 | 28.9229 | 44.997 | 231.801 | 757.094 | 121.545 | -3.0004 | -16.7553 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32648 | 583.121 | 36.4789 | 40
39616.5236 | 2008-06-17 | #NAME? | 0 | 0 | 8.0556 | 90745.3 | 29.1043 | 44.9375 | 230.751 | 759.719 | 119.444 | -3.0192 | -16.7419 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32267 | 583.121 | 36.4789 | 40
39616.5243 | 2008-06-17 | #NAME? | 0 | 0 | 8.0627 | 90727.8 | 29.263 | 45.1597 | 229.438 | 762.344 | 119.707 | -3.0192 | -16.7619 | 1.5059 | 51.0582 | 15.1743 | 1.4605 | 0.32267 | 583.121 | 36.4789 | 40
39616.525 | 2008-06-17 | #NAME? | 0 | 0 | 8.1658 | 90709.5 | 29.3537 | 44.9673 | 228.388 | 760.507 | 121.545 | -3.0098 | -16.7352 | 1.5059 | 51.0582 | 15.2083 | 1.4605 | 0.32369 | 583.121 | 36.7803 | 40
39616.5257 | 2008-06-17 | #NAME? | 0 | 0 | 8.6765 | 90695.6 | 29.4898 | 45.0945 | 229.438 | 757.094 | 118.657 | -2.9911 | -16.7553 | 1.5059 | 51.0582 | 15.2083 | 1.4605 | 0.32241 | 582.751 | 36.4789 | 40
39616.5264 | 2008-06-17 | #NAME? | 0 | 0 | 8.7494 | 90679.9 | 29.6485 | 45.0619 | 234.164 | 756.306 | 120.495 | -2.9911 | -16.7419 | 1.5059 | 51.0582 | 15.2083 | 1.4605 | 0.32598 | 583.121 | 36.7803 | 40
39616.5271 | 2008-06-17 | #NAME? | 0 | 0 | 8.7747 | 90665.2 | 29.7619 | 45.1597 | 237.314 | 758.407 | 120.757 | -3.0098 | -16.7285 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32521 | 583.121 | 36.4789 | 40
39616.5278 | 2008-06-17 | #NAME? | 0 | 0 | 8.7806 | 90652.9 | 29.7166 | 45.0945 | 236.526 | 760.769 | 118.919 | -3.0474 | -16.7553 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32419 | 583.121 | 36.4789 | 40
39616.5285 | 2008-06-17 | #NAME? | 0 | 0 | 8.7778 | 90649.7 | 29.6485 | 44.9078 | 235.476 | 762.607 | 121.282 | -3.0004 | -16.7285 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32267 | 583.121 | 36.7803 | 40
39616.5292 | 2008-06-17 | #NAME? | 0 | 0 | 8.6256 | 90647.9 | 29.7166 | 45.4205 | 233.901 | 760.244 | 119.444 | -3.038 | -16.7352 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32598 | 583.121 | 36.7803 | 40
39616.5299 | 2008-06-17 | #NAME? | 0 | 0 | 8.1607 | 90643.4 | 29.6259 | 45.2575 | 232.589 | 756.831 | 119.969 | -3.0004 | -16.7352 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32394 | 582.751 | 36.7803 | 40
39616.5306 | 2008-06-17 | #NAME? | 0 | 0 | 8.1357 | 90649 | 29.5351 | 44.878 | 231.276 | 756.569 | 121.282 | -3.0474 | -16.7486 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32598 | 582.751 | 36.7803 | 40
39616.5312 | 2008-06-17 | #NAME? | 0 | 0 | 8.1401 | 90654.6 | 29.4671 | 45.2575 | 230.488 | 758.144 | 118.394 | -3.0286 | -16.7486 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32292 | 582.751 | 36.4789 | 40
39616.5319 | 2008-06-17 | #NAME? | 0 | 0 | 8.1142 | 90659 | 29.2177 | 45.3227 | 228.913 | 761.294 | 121.02 | -3.038 | -16.7352 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32241 | 583.121 | 36.7803 | 40
39616.5326 | 2008-06-17 | #NAME? | 0 | 0 | 8.1082 | 90675.9 | 29.059 | 45.2249 | 228.651 | 762.082 | 119.969 | -2.9911 | -16.7285 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32394 | 583.121 | 36.4789 | 40
39616.5333 | 2008-06-17 | #NAME? | 0 | 0 | 8.1856 | 90690 | 28.9683 | 45.2575 | 231.538 | 759.457 | 119.444 | -3.0192 | -16.7285 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32521 | 582.751 | 36.4789 | 40
39616.534 | 2008-06-17 | #NAME? | 0 | 0 | 8.626 | 90701.1 | 28.9456 | 44.6698 | 236.526 | 756.306 | 121.545 | -3.0474 | -16.7285 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32419 | 583.121 | 36.4789 | 40
39616.5347 | 2008-06-17 | #NAME? | 0 | 0 | 8.6922 | 90700.6 | 28.9002 | 45.4205 | 237.314 | 756.831 | 118.657 | -3.0192 | -16.7352 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32292 | 583.121 | 36.7803 | 40
39616.5354 | 2008-06-17 | #NAME? | 0 | 0 | 8.6614 | 90698.9 | 28.6735 | 45.3553 | 236.001 | 758.407 | 120.495 | -3.0004 | -16.7419 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32292 | 583.121 | 36.4789 | 40
39616.5361 | 2008-06-17 | #NAME? | 0 | 0 | 8.6721 | 90704.8 | 28.9229 | 44.9375 | 234.951 | 761.032 | 120.495 | -3.0286 | -16.7352 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32699 | 582.751 | 36.7803 | 40
39616.5368 | 2008-06-17 | #NAME? | 0 | 0 | 8.6828 | 90690.8 | 29.059 | 45.3879 | 233.376 | 762.607 | 118.919 | -3.0192 | -16.7619 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32674 | 582.751 | 36.7803 | 40
39616.5375 | 2008-06-17 | #NAME? | 0 | 0 | 8.6811 | 90672.2 | 29.2177 | 44.997 | 232.063 | 760.769 | 121.282 | -3.0286 | -16.7486 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32241 | 582.751 | 36.4789 | 40
39616.5382 | 2008-06-17 | #NAME? | 0 | 0 | 8.7126 | 90653.5 | 29.4671 | 45.2575 | 231.013 | 757.094 | 119.444 | -3.0192 | -16.7218 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32496 | 583.121 | 36.7803 | 40
39616.5389 | 2008-06-17 | #NAME? | 0 | 0 | 8.5516 | 90629.3 | 29.6032 | 44.9375 | 229.963 | 756.306 | 119.444 | -3.0098 | -16.7352 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32521 | 583.121 | 36.4789 | 40
39616.5396 | 2008-06-17 | #NAME? | 0 | 0 | 8.117 | 90612.2 | 29.6259 | 44.9673 | 228.913 | 758.407 | 121.545 | -3.038 | -16.7352 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32292 | 582.751 | 36.7803 | 40
39616.5403 | 2008-06-17 | #NAME? | 0 | 0 | 8.0887 | 90605.5 | 29.6485 | 45.1923 | 228.913 | 761.032 | 118.657 | -3.0004 | -16.7285 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32343 | 583.121 | 36.7803 | 40
39616.541 | 2008-06-17 | #NAME? | 0 | 0 | 8.0699 | 90598.1 | 29.5125 | 44.9375 | 232.851 | 762.607 | 120.495 | -3.0004 | -16.7553 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32547 | 582.751 | 36.4789 | 40
39616.5417 | 2008-06-17 | #NAME? | 0 | 0 | 8.0772 | 90604.9 | 29.4671 | 45.1271 | 236.789 | 760.244 | 121.02 | -3.038 | -16.7352 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32445 | 582.751 | 36.7803 | 40
39616.5424 | 2008-06-17 | #NAME? | 0 | 0 | 8.0546 | 90604.9 | 29.5351 | 44.997 | 236.789 | 757.094 | 118.919 | -3.0192 | -16.7352 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32216 | 583.121 | 36.4789 | 40
39616.5431 | 2008-06-17 | #NAME? | 0 | 0 | 8.2164 | 90602.7 | 29.4671 | 44.9078 | 235.476 | 756.569 | 121.282 | -3.038 | -16.7285 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32674 | 583.121 | 36.4789 | 40
39616.5437 | 2008-06-17 | #NAME? | 0 | 0 | 8.662 | 90605.1 | 29.3537 | 45.1923 | 234.426 | 758.144 | 119.444 | -3.038 | -16.7553 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32267 | 583.121 | 36.7803 | 40
39616.5444 | 2008-06-17 | #NAME? | 0 | 0 | 8.6784 | 90610.6 | 29.195 | 44.878 | 232.851 | 761.032 | 119.444 | -3.0286 | -16.7486 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32572 | 582.751 | 36.4789 | 40
39616.5451 | 2008-06-17 | #NAME? | 0 | 0 | 8.6799 | 90617.8 | 29.263 | 45.3553 | 231.538 | 762.344 | 121.545 | -3.0286 | -16.7419 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32419 | 583.121 | 36.4789 | 40
39616.5458 | 2008-06-17 | #NAME? | 0 | 0 | 8.6689 | 90621.2 | 29.059 | 45.0619 | 230.488 | 759.982 | 118.919 | -2.9911 | -16.7553 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32292 | 583.121 | 36.7803 | 40
39616.5465 | 2008-06-17 | #NAME? | 0 | 0 | 8.6705 | 90628.3 | 28.9683 | 45.0619 | 229.438 | 757.094 | 120.495 | -3.0098 | -16.7419 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32292 | 582.751 | 36.7803 | 40
39616.5472 | 2008-06-17 | #NAME? | 0 | 0 | 8.4808 | 90629.7 | 29.0136 | 45.0945 | 228.651 | 756.831 | 120.495 | -2.9911 | -16.7285 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32267 | 582.751 | 36.4789 | 40
39616.5479 | 2008-06-17 | #NAME? | 0 | 0 | 8.0131 | 90627 | 29.0136 | 44.9673 | 230.226 | 757.882 | 118.919 | -3.0192 | -16.7486 | 1.5059 | 51.0921 | 15.2083 | 1.4605 | 0.32394 | 583.121 | 36.7803 | 40
39616.5486 | 2008-06-17 | #NAME? | 0 | 0 | 7.995 | 90617.4 | 29.3311 | 45.4531 | 235.476 | 760.769 | 121.282 | -3.0192 | -16.7419 | 1.5059 | 51.0921 | 15.2083 | 1.4709 | 0.32648 | 583.121 | 36.7803 | 40
39616.5493 | 2008-06-17 | #NAME? | 0 | 0 | 7.9876 | 90596.7 | 29.4218 | 45.4205 | 237.051 | 762.607 | 119.182 | -3.0474 | -16.7419 | 1.5059 | 51.0921 | 15.2083 | 1.4709 | 0.32216 | 583.121 | 36.4789 | 40
39616.55 | 2008-06-17 | #NAME? | 0 | 0 | 7.9986 | 90579.3 | 29.5125 | 44.9673 | 236.264 | 761.032 | 119.969 | -3.0004 | -16.7619 | 1.5059 | 51.0921 | 15.2083 | 1.4709 | 0.32598 | 583.121 | 36.4789 | 40
39616.5507 | 2008-06-17 | #NAME? | 0 | 0 | 8.0245 | 90568.5 | 29.6485 | 45.0293 | 234.951 | 756.831 | 121.282 | -3.0192 | -16.7553 | 1.5059 | 51.0921 | 15.2083 | 1.4709 | 0.32216 | 583.121 | 36.7803 | 40
39616.5514 | 2008-06-17 | #NAME? | 0 | 0 | 8.0157 | 90552.4 | 29.7846 | 45.4205 | 233.901 | 756.569 | 118.394 | -3.0568 | -16.7218 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32267 | 583.121 | 36.4789 | 40
39616.5521 | 2008-06-17 | #NAME? | 0 | 0 | 6.8872 | 90540.4 | 29.7166 | 44.997 | 232.326 | 758.407 | 121.02 | -2.9911 | -16.7553 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32623 | 582.751 | 36.4789 | 40
39616.5528 | 2008-06-17 | #NAME? | 0 | 0 | 5.6372 | 90536.4 | 29.7392 | 45.4857 | 231.276 | 761.032 | 119.969 | -2.9911 | -16.7419 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.3219 | 583.121 | 36.7803 | 40
39616.5535 | 2008-06-17 | #NAME? | 0 | 0 | 5.5968 | 90531.9 | 29.7619 | 44.878 | 229.963 | 762.344 | 119.182 | -3.0286 | -16.7486 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32598 | 582.751 | 36.4789 | 40
39616.5542 | 2008-06-17 | #NAME? | 0 | 0 | 5.5931 | 90535.1 | 29.5351 | 45.0619 | 228.913 | 759.457 | 121.545 | -3.0474 | -16.7553 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32419 | 583.121 | 36.4789 | 40
39616.5549 | 2008-06-17 | #NAME? | 0 | 0 | 5.5857 | 90543.4 | 29.6259 | 45.1597 | 228.388 | 756.569 | 118.919 | -3.0474 | -16.7285 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32241 | 582.751 | 36.4789 | 40
39616.5556 | 2008-06-17 | #NAME? | 0 | 0 | 5.5865 | 90542.2 | 29.5578 | 44.9673 | 231.801 | 756.569 | 120.232 | -2.9817 | -16.7486 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32521 | 583.121 | 36.7803 | 40
39616.5563 | 2008-06-17 | #NAME? | 0 | 0 | 5.5775 | 90554.5 | 29.3084 | 45.2575 | 236.526 | 758.407 | 121.282 | -2.9911 | -16.7419 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32496 | 583.121 | 36.7803 | 40
39616.5569 | 2008-06-17 | #NAME? | 0 | 0 | 5.5761 | 90566.5 | 29.2177 | 44.9078 | 237.051 | 761.032 | 118.394 | -3.0286 | -16.7486 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32394 | 583.121 | 36.4789 | 40
39616.5576 | 2008-06-17 | #NAME? | 0 | 0 | 5.9018 | 90574.5 | 29.1497 | 44.997 | 235.739 | 762.344 | 121.02 | -3.0192 | -16.7419 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32674 | 583.121 | 36.4789 | 40
39616.5583 | 2008-06-17 | #NAME? | 0 | 0 | 6.1306 | 90578 | 29.059 | 44.9078 | 234.426 | 759.719 | 120.232 | -3.0098 | -16.7352 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32521 | 583.121 | 36.4789 | 40
39616.559 | 2008-06-17 | #NAME? | 0 | 0 | 6.1306 | 90583.1 | 29.0136 | 45.3553 | 232.851 | 756.569 | 119.182 | -3.038 | -16.7352 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32674 | 583.121 | 36.4789 | 40
39616.5597 | 2008-06-17 | #NAME? | 0 | 0 | 6.1257 | 90584.2 | 29.0363 | 44.9673 | 231.801 | 756.831 | 121.545 | -3.0004 | -16.7486 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32267 | 583.121 | 36.4789 | 40
39616.5604 | 2008-06-17 | #NAME? | 0 | 0 | 6.1282 | 90573.1 | 29.3084 | 45.2901 | 230.751 | 758.669 | 119.444 | -3.0286 | -16.7285 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32572 | 582.751 | 36.4789 | 40
39616.5611 | 2008-06-17 | #NAME? | 0 | 0 | 6.1284 | 90551.5 | 29.4444 | 45.4857 | 229.438 | 761.557 | 119.707 | -3.0004 | -16.7553 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32394 | 583.121 | 36.4789 | 40
39616.5618 | 2008-06-17 | #NAME? | 0 | 0 | 6.0304 | 90533 | 29.5351 | 44.9078 | 228.651 | 762.082 | 121.545 | -3.0192 | -16.7352 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32572 | 582.751 | 36.4789 | 40
39616.5625 | 2008-06-17 | #NAME? | 0 | 0 | 5.6306 | 90517.2 | 29.7166 | 45.1597 | 229.438 | 759.194 | 118.394 | -3.0192 | -16.7419 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32623 | 583.121 | 36.4789 | 40
39616.5632 | 2008-06-17 | #NAME? | 0 | 0 | 5.5807 | 90499.1 | 29.8073 | 44.9673 | 234.426 | 756.306 | 120.757 | -3.0286 | -16.7419 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32598 | 582.751 | 36.4789 | 40
39616.5639 | 2008-06-17 | #NAME? | 0 | 0 | 5.5882 | 90490.1 | 29.6939 | 44.997 | 237.051 | 756.569 | 120.232 | -3.0098 | -16.7553 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32547 | 583.121 | 36.7803 | 40
39616.5646 | 2008-06-17 | #NAME? | 0 | 0 | 5.5882 | 90490.8 | 29.6485 | 45.2901 | 236.526 | 758.407 | 119.182 | -3.0192 | -16.7352 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32343 | 583.121 | 36.7803 | 40
39616.5653 | 2008-06-17 | #NAME? | 0 | 0 | 5.7397 | 90494.5 | 29.4898 | 44.878 | 235.214 | 761.294 | 121.545 | -2.9911 | -16.7419 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32547 | 583.121 | 36.4789 | 40
39616.566 | 2008-06-17 | #NAME? | 0 | 0 | 6.1352 | 90501.9 | 29.3991 | 45.3879 | 233.901 | 762.607 | 119.182 | -3.0192 | -16.7553 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32623 | 583.121 | 36.4789 | 40
39616.5667 | 2008-06-17 | #NAME? | 0 | 0 | 6.1466 | 90509.1 | 29.3764 | 44.9375 | 232.326 | 760.507 | 119.969 | -3.0098 | -16.7553 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32369 | 583.121 | 36.4789 | 40
39616.5674 | 2008-06-17 | #NAME? | 0 | 0 | 6.1463 | 90514.5 | 29.3084 | 45.3553 | 231.276 | 756.831 | 121.545 | -2.9911 | -16.7486 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32598 | 582.751 | 36.4789 | 40
39616.5681 | 2008-06-17 | #NAME? | 0 | 0 | 6.1308 | 90511.9 | 29.263 | 45.1597 | 230.226 | 756.569 | 118.394 | -3.0192 | -16.7619 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32267 | 583.121 | 36.4789 | 40
39616.5687 | 2008-06-17 | #NAME? | 0 | 0 | 6.135 | 90518.4 | 29.0816 | 45.0293 | 228.913 | 758.144 | 120.757 | -3.0192 | -16.7352 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32674 | 583.121 | 36.4789 | 40
39616.5694 | 2008-06-17 | #NAME? | 0 | 0 | 5.9948 | 90522.1 | 29.0816 | 45.0293 | 228.388 | 760.769 | 120.232 | -3.0004 | -16.7285 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32598 | 583.121 | 36.4789 | 40
39616.5701 | 2008-06-17 | #NAME? | 0 | 0 | 5.5758 | 90526.6 | 29.0363 | 45.0293 | 231.276 | 762.344 | 119.182 | -3.0004 | -16.7486 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32419 | 582.751 | 36.7803 | 40
39616.5708 | 2008-06-17 | #NAME? | 0 | 0 | 5.5396 | 90528.7 | 29.0136 | 45.4531 | 236.264 | 759.719 | 121.545 | -3.0098 | -16.7419 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.3247 | 583.121 | 36.4789 | 40
39616.5715 | 2008-06-17 | #NAME? | 0 | 0 | 5.5231 | 90527.7 | 28.8776 | 44.9375 | 237.576 | 756.569 | 118.657 | -3.0286 | -16.7352 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32394 | 582.751 | 36.4789 | 40
39616.5722 | 2008-06-17 | #NAME? | 0 | 0 | 5.5255 | 90531.1 | 28.9229 | 45.1597 | 236.264 | 756.831 | 120.495 | -3.0192 | -16.7285 | 1.5059 | 51.126 | 15.2083 | 1.4709 | 0.32674 | 582.751 | 36.7803 | 40
39616.5729 | 2008-06-17 | #NAME? | 0 | 0 | 5.5138 | 90530.3 | 28.8776 | 44.9673 | 234.951 | 758.669 | 120.757 | -3.0286 | -16.7419 | 1.5059 | 51.126 | 15.2083 | 1.4916 | 0.32369 | 583.121 | 36.7803 | 40
39616.5736 | 2008-06-17 | #NAME? | 0 | 0 | 5.6402 | 90523.3 | 29.0363 | 45.0293 | 233.376 | 761.557 | 118.919 | -3.0192 | -16.7553 | 1.5059 | 51.126 | 15.2083 | 1.4916 | 0.32394 | 583.121 | 36.7803 | 40
39616.5743 | 2008-06-17 | #NAME? | 0 | 0 | 6.0829 | 90507.7 | 29.2404 | 45.3879 | 232.063 | 762.344 | 121.282 | -3.0192 | -16.7486 | 1.5059 | 51.126 | 15.2083 | 1.4916 | 0.32292 | 583.121 | 36.4789 | 40
39616.575 | 2008-06-17 | #NAME? | 0 | 0 | 6.135 | 90488.1 | 29.3537 | 44.997 | 231.013 | 759.719 | 119.444 | -3.0286 | -16.7486 | 1.5059 | 51.126 | 15.2083 | 1.4916 | 0.32241 | 583.121 | 36.4789 | 40
39616.5757 | 2008-06-17 | #NAME? | 0 | 0 | 6.1463 | 90473.3 | 29.4898 | 45.0619 | 229.701 | 756.569 | 119.444 | -3.0004 | -16.7352 | 1.5059 | 51.126 | 15.2083 | 1.4916 | 0.32598 | 583.121 | 36.4789 | 40
39616.5764 | 2008-06-17 | #NAME? | 0 | 0 | 6.1501 | 90456.1 | 29.5805 | 45.1597 | 228.913 | 756.569 | 121.807 | -3.0286 | -16.7619 | 1.5059 | 51.126 | 15.2083 | 1.4916 | 0.32572 | 582.381 | 36.4789 | 40
39616.5771 | 2008-06-17 | #NAME? | 0 | 0 | 6.1568 | 90443.4 | 29.6259 | 45.2901 | 228.913 | 759.194 | 118.657 | -3.0098 | -16.7419 | 1.5059 | 51.16 | 15.2083 | 1.4916 | 0.3247 | 582.751 | 36.4789 | 40
39616.5778 | 2008-06-17 | #NAME? | 0 | 0 | 6.1573 | 90438.9 | 29.5125 | 45.0619 | 232.851 | 762.344 | 120.495 | -3.0098 | -16.7285 | 1.5059 | 51.16 | 15.2083 | 1.4916 | 0.32623 | 582.751 | 36.4789 | 40
39616.5785 | 2008-06-17 | #NAME? | 0 | 0 | 5.8555 | 90443.7 | 29.4444 | 44.997 | 237.051 | 761.557 | 120.757 | -3.0192 | -16.7285 | 1.5059 | 51.16 | 15.2083 | 1.4916 | 0.3247 | 583.491 | 36.4789 | 40
39616.5792 | 2008-06-17 | #NAME? | 0 | 0 | 5.5526 | 90448.5 | 29.263 | 45.0945 | 237.314 | 757.882 | 118.657 | -3.0286 | -16.7486 | 1.5059 | 51.16 | 15.2083 | 1.4916 | 0.32241 | 583.491 | 36.4789 | 40
39616.5799 | 2008-06-17 | #NAME? | 0 | 0 | 5.5452 | 90455.4 | 29.3084 | 44.997 | 235.739 | 756.569 | 121.282 | -3.0098 | -16.7486 | 1.5059 | 51.16 | 15.2083 | 1.4916 | 0.32216 | 583.121 | 36.4789 | 40
39616.5806 | 2008-06-17 | #NAME? | 0 | 0 | 5.544 | 90454.5 | 29.2404 | 45.0619 | 234.426 | 757.882 | 119.444 | -2.9911 | -16.7285 | 1.472 | 51.16 | 15.2083 | 1.4916 | 0.32292 | 582.751 | 36.4789 | 40
39616.5813 | 2008-06-17 | #NAME? | 0 | 0 | 5.5384 | 90457.9 | 29.1723 | 45.0293 | 233.114 | 760.507 | 119.707 | -3.0474 | -16.7352 | 1.472 | 51.16 | 15.2083 | 1.4916 | 0.32572 | 583.491 | 36.4789 | 40
39616.5819 | 2008-06-17 | #NAME? | 0 | 0 | 4.6214 | 90460.7 | 28.9229 | 45.0619 | 231.801 | 762.607 | 121.545 | -3.0192 | -16.7285 | 1.472 | 51.16 | 15.2083 | 1.4916 | 0.32648 | 582.751 | 36.4789 | 40
39616.5826 | 2008-06-17 | #NAME? | 0 | 0 | 2.4699 | 90471.1 | 28.9456 | 45.2901 | 230.751 | 761.032 | 118.919 | -3.0286 | -16.7553 | 1.472 | 51.16 | 15.2083 | 1.4916 | 0.32623 | 583.121 | 36.4789 | 40
39616.5833 | 2008-06-17 | #NAME? | 0 | 0 | 2.5287 | 90470.4 | 28.9002 | 45.1923 | 229.701 | 757.094 | 120.232 | -3.038 | -16.7486 | 1.472 | 51.16 | 15.2083 | 1.4916 | 0.32445 | 583.121 | 36.4789 | 40
39616.584 | 2008-06-17 | #NAME? | 0 | 0 | 2.5253 | 90472.4 | 28.8549 | 44.997 | 228.651 | 756.306 | 120.757 | -3.038 | -16.7352 | 1.472 | 51.16 | 15.2083 | 1.4916 | 0.32547 | 583.491 | 36.4789 | 40
39616.5847 | 2008-06-17 | #NAME? | 0 | 0 | 2.52 | 90473.9 | 28.7868 | 44.9673 | 228.651 | 758.144 | 118.657 | -3.0286 | -16.7553 | 1.472 | 51.16 | 15.2083 | 1.4916 | 0.32572 | 583.491 | 36.4789 | 40
39616.5854 | 2008-06-17 | #NAME? | 0 | 0 | 2.5182 | 90473.5 | 28.8549 | 45.2575 | 232.851 | 761.032 | 121.02 | -3.038 | -16.6951 | 1.472 | 51.16 | 15.2083 | 1.4916 | 0.32598 | 583.121 | 36.4789 | 40
39616.5861 | 2008-06-17 | #NAME? | 0 | 0 | 2.5161 | 90468.8 | 28.8322 | 44.9375 | 237.051 | 762.344 | 119.969 | -2.9911 | -16.7553 | 1.472 | 51.16 | 15.2083 | 1.4916 | 0.32318 | 583.491 | 36.4789 | 40
39616.5868 | 2008-06-17 | #NAME? | 0 | 0 | 2.6264 | 90464.1 | 28.8322 | 45.1597 | 237.051 | 760.507 | 119.182 | -3.0004 | -16.7419 | 1.472 | 51.16 | 15.2083 | 1.4916 | 0.32343 | 582.381 | 36.4789 | 40
39616.5875 | 2008-06-17 | #NAME? | 0 | 0 | 2.8996 | 90462.8 | 28.7868 | 45.1923 | 235.739 | 756.569 | 121.545 | -3.0098 | -16.7486 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32648 | 583.121 | 36.7803 | 40
39616.5882 | 2008-06-17 | #NAME? | 0 | 0 | 2.9135 | 90456.5 | 28.9229 | 45.3553 | 234.426 | 756.306 | 119.182 | -3.0192 | -16.7352 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32292 | 582.381 | 36.4789 | 40
39616.5889 | 2008-06-17 | #NAME? | 0 | 0 | 2.9175 | 90443.1 | 29.059 | 44.7293 | 233.114 | 758.144 | 119.969 | -2.9817 | -16.7686 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32318 | 583.121 | 36.7803 | 40
39616.5896 | 2008-06-17 | #NAME? | 0 | 0 | 2.9249 | 90423.8 | 29.3537 | 45.2575 | 231.801 | 761.557 | 121.282 | -3.038 | -16.7486 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32292 | 582.381 | 36.7803 | 40
39616.5903 | 2008-06-17 | #NAME? | 0 | 0 | 2.8244 | 90401.2 | 29.4218 | 44.9375 | 230.488 | 762.607 | 118.394 | -3.038 | -16.7619 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32419 | 582.751 | 36.7803 | 40
39616.591 | 2008-06-17 | #NAME? | 0 | 0 | 2.5486 | 90386.8 | 29.5351 | 45.3553 | 229.438 | 759.719 | 120.757 | -3.038 | -16.7553 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32267 | 582.751 | 36.4789 | 40
39616.5917 | 2008-06-17 | #NAME? | 0 | 0 | 2.5254 | 90371 | 29.6259 | 45.1597 | 228.913 | 756.831 | 120.495 | -3.0098 | -16.7553 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32598 | 583.491 | 36.4789 | 40
39616.5924 | 2008-06-17 | #NAME? | 0 | 0 | 2.5315 | 90362.3 | 29.5125 | 44.878 | 229.176 | 756.569 | 119.182 | -3.038 | -16.7419 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32241 | 582.751 | 36.4789 | 40
39616.5931 | 2008-06-17 | #NAME? | 0 | 0 | 2.5301 | 90362.9 | 29.5578 | 45.1923 | 233.901 | 758.932 | 121.282 | -2.9911 | -16.7285 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32623 | 583.121 | 36.7803 | 40
39616.5938 | 2008-06-17 | #NAME? | 0 | 0 | 2.6302 | 90362.1 | 29.3764 | 44.997 | 237.051 | 761.557 | 119.182 | -3.0286 | -16.7686 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32674 | 582.381 | 36.4789 | 40
39616.5944 | 2008-06-17 | #NAME? | 0 | 0 | 2.9086 | 90371.3 | 29.2857 | 44.878 | 236.789 | 762.344 | 119.969 | -3.0098 | -16.7419 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32343 | 582.751 | 36.4789 | 40
39616.5951 | 2008-06-17 | #NAME? | 0 | 0 | 2.9242 | 90376.5 | 29.3991 | 45.2575 | 235.476 | 759.457 | 121.545 | -3.0286 | -16.7553 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32292 | 582.381 | 36.4789 | 40
39616.5958 | 2008-06-17 | #NAME? | 0 | 0 | 2.9124 | 90368.5 | 29.2404 | 45.3227 | 234.426 | 756.831 | 118.657 | -3.0286 | -16.7285 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32292 | 583.491 | 36.4789 | 40
39616.5965 | 2008-06-17 | #NAME? | 0 | 0 | 2.9154 | 90374.7 | 29.1723 | 44.7293 | 232.589 | 756.569 | 120.757 | -3.0192 | -16.7285 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32292 | 583.491 | 36.4789 | 40
39616.5972 | 2008-06-17 | #NAME? | 0 | 0 | 2.8316 | 90380.2 | 29.2177 | 44.997 | 231.801 | 758.669 | 120.495 | -3.0192 | -16.7619 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32598 | 583.491 | 36.4789 | 40
39616.5979 | 2008-06-17 | #NAME? | 0 | 0 | 2.5323 | 90380.7 | 29.0816 | 45.2575 | 230.226 | 760.769 | 118.919 | -3.0098 | -16.7486 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32267 | 582.381 | 36.4789 | 40
39616.5986 | 2008-06-17 | #NAME? | 0 | 0 | 2.5 | 90388.3 | 29.0136 | 44.8483 | 229.176 | 762.607 | 121.545 | -3.0192 | -16.7553 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32547 | 583.491 | 36.7803 | 40
39616.5993 | 2008-06-17 | #NAME? | 0 | 0 | 2.5008 | 90390.7 | 28.9456 | 44.997 | 228.388 | 759.982 | 119.444 | -3.0192 | -16.7352 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32521 | 582.751 | 36.4789 | 40
39616.6 | 2008-06-17 | #NAME? | 0 | 0 | 2.491 | 90396.1 | 28.8549 | 45.1597 | 230.488 | 756.831 | 119.707 | -3.0192 | -16.7553 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32598 | 582.381 | 36.7803 | 40
39616.6007 | 2008-06-17 | #NAME? | 0 | 0 | 2.5856 | 90393.1 | 29.0136 | 45.0293 | 235.739 | 756.569 | 121.282 | -3.0004 | -16.7352 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32394 | 582.381 | 36.4789 | 40
39616.6014 | 2008-06-17 | #NAME? | 0 | 0 | 2.8842 | 90381 | 29.1723 | 45.2901 | 237.314 | 758.932 | 118.394 | -3.0192 | -16.7419 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32674 | 582.751 | 36.4789 | 40
39616.6021 | 2008-06-17 | #NAME? | 0 | 0 | 2.9146 | 90366.2 | 29.1723 | 44.9078 | 236.001 | 762.344 | 121.02 | -3.0192 | -16.7352 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32318 | 582.751 | 36.4789 | 40
39616.6028 | 2008-06-17 | #NAME? | 0 | 0 | 2.9183 | 90360.3 | 29.2404 | 45.5834 | 234.951 | 761.557 | 119.969 | -2.9911 | -16.7285 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32674 | 583.121 | 36.4789 | 40
39616.6035 | 2008-06-17 | #NAME? | 0 | 0 | 2.9165 | 90351.7 | 29.2857 | 44.9375 | 233.376 | 757.619 | 119.444 | -3.0192 | -16.7352 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32521 | 582.751 | 36.7803 | 40
39616.6042 | 2008-06-17 | #NAME? | 0 | 0 | 2.8371 | 90342.6 | 29.4671 | 45.4205 | 232.063 | 756.306 | 121.807 | -2.9911 | -16.7352 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32598 | 583.491 | 36.7803 | 40
39616.6049 | 2008-06-17 | #NAME? | 0 | 0 | 2.5431 | 90326.6 | 29.5578 | 44.997 | 231.013 | 757.882 | 118.919 | -3.0192 | -16.7352 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32598 | 583.121 | 36.4789 | 40
39616.6056 | 2008-06-17 | #NAME? | 0 | 0 | 2.5222 | 90316.2 | 29.5351 | 45.0293 | 229.963 | 760.507 | 120.232 | -3.0004 | -16.7419 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32369 | 583.491 | 36.4789 | 40
39616.6062 | 2008-06-17 | #NAME? | 0 | 0 | 2.5305 | 90312.4 | 29.3311 | 45.3553 | 228.913 | 762.607 | 120.757 | -2.9911 | -16.7285 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32292 | 583.121 | 36.4789 | 40
39616.6069 | 2008-06-17 | #NAME? | 0 | 0 | 2.5289 | 90320.9 | 29.2857 | 44.9673 | 228.651 | 761.032 | 118.919 | -3.0098 | -16.7619 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.32343 | 582.381 | 36.4789 | 40
39616.6076 | 2008-06-17 | #NAME? | 0 | 0 | 2.0146 | 90325 | 29.2857 | 45.5508 | 232.589 | 757.357 | 120.757 | -2.9911 | -16.7419 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.33539 | 582.751 | 37.3833 | 40
39616.6083 | 2008-06-17 | #NAME? | 0 | 0 | 0.1857 | 90322 | 29.3311 | 45.2575 | 237.051 | 756.306 | 120.232 | -2.9817 | -16.7419 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.33615 | 583.121 | 37.0818 | 40
39616.609 | 2008-06-17 | #NAME? | 0 | 0 | -0.004071 | 90319.5 | 29.1497 | 44.9375 | 237.051 | 758.144 | 119.182 | -3.0192 | -16.7486 | 1.472 | 51.16 | 15.2083 | 1.5019 | 0.33615 | 582.751 | 37.3833 | 40
39616.6097 | 2008-06-17 | #NAME? | 0 | 0 | -0.0045309 | 90324.8 | 29.0136 | 45.3879 | 235.476 | 760.769 | 121.282 | -3.0004 | -16.7486 | 40.8832 | 51.16 | 15.2083 | 3.8947 | 0.33895 | 582.011 | 37.3833 | 40
39616.6104 | 2008-06-17 | #NAME? | 0 | 0 | -0.0043467 | 90335.6 | 28.7868 | 44.9673 | 234.426 | 762.344 | 119.444 | -3.0474 | -16.7486 | 42.0364 | 51.16 | 15.2083 | 18.075 | 0.33615 | 583.121 | 37.6848 | 40
39616.6111 | 2008-06-17 | #NAME? | 0 | 0 | -0.0033104 | 90340.9 | 28.9002 | 45.0945 | 232.851 | 760.507 | 119.707 | -3.0286 | -16.7419 | 41.9685 | 51.16 | 15.2083 | 18.6965 | 0.33539 | 582.011 | 37.3833 | 40
39616.6118 | 2008-06-17 | #NAME? | 0 | 0 | -0.003758 | 90336.4 | 28.7868 | 45.1923 | 231.538 | 757.094 | 121.545 | -3.0474 | -16.7486 | 41.9346 | 51.16 | 15.2083 | 18.7276 | 0.33488 | 582.381 | 37.3833 | 40
39616.6125 | 2008-06-17 | #NAME? | 0 | 0 | 2.1441 | 90344.8 | 28.6508 | 44.9673 | 230.488 | 756.569 | 118.394 | -3.0098 | -16.7419 | 41.9346 | 51.16 | 15.2083 | 18.7379 | 0.70104 | 583.121 | 53.9646 | 40
39616.6132 | 2008-06-17 | #NAME? | 0 | 0 | 9.196 | 90345.8 | 28.5828 | 45.1597 | 228.913 | 757.882 | 120.495 | -3.0098 | -16.7553 | 41.9007 | 51.16 | 15.2083 | 18.7483 | 0.34175 | 582.751 | 37.0818 | 40
39616.6139 | 2008-06-17 | #NAME? | 0 | 0 | 4.1435 | 90343.7 | 28.5601 | 45.2901 | 228.388 | 759.719 | 120.757 | -3.0286 | -16.7419 | 41.9007 | 51.16 | 15.2083 | 18.7483 | 0.30562 | 583.121 | 35.273 | 40
39616.6146 | 2008-06-17 | #NAME? | 0 | 0 | 1.9843 | 90342.1 | 28.288 | 44.997 | 231.276 | 762.607 | 118.919 | -2.9723 | -16.7419 | 41.9007 | 51.16 | 15.2083 | 18.7586 | 0.29722 | 582.751 | 34.9715 | 40
39616.6153 | 2008-06-17 | #NAME? | 0 | 0 | 1.6048 | 90346.9 | 28.3333 | 45.0293 | 236.001 | 761.819 | 121.02 | -3.0286 | -16.7419 | 41.9007 | 51.16 | 15.2083 | 18.769 | 0.29748 | 583.121 | 34.67 | 40
39616.616 | 2008-06-17 | #NAME? | 0 | 0 | 1.3986 | 90345 | 28.288 | 45.5834 | 237.051 | 758.144 | 119.444 | -3.038 | -16.7619 | 41.9007 | 51.16 | 15.2083 | 18.769 | 0.29646 | 582.751 | 34.67 | 40
39616.6167 | 2008-06-17 | #NAME? | 0 | 0 | 1.2612 | 90343.2 | 28.1293 | 45.2249 | 235.476 | 756.306 | 119.444 | -2.9911 | -16.7419 | 41.9007 | 51.16 | 15.2083 | 18.8001 | 0.29468 | 582.751 | 34.67 | 40
39616.6174 | 2008-06-17 | #NAME? | 0 | 0 | 1.1691 | 90344.8 | 28.2653 | 44.8483 | 234.164 | 756.569 | 121.807 | -3.0286 | -16.7553 | 41.9007 | 51.16 | 15.2083 | 18.8001 | 0.2929 | 582.751 | 34.67 | 40
39616.6181 | 2008-06-17 | #NAME? | 0 | 0 | 0.6476 | 90331.5 | 28.2653 | 45.2575 | 232.589 | 758.669 | 118.657 | -3.038 | -16.7419 | 41.9007 | 51.16 | 15.2083 | 18.8001 | 0.23259 | 582.751 | 33.1626 | 40
39616.6188 | 2008-06-17 | #NAME? | 0 | 0 | 0 | 90321.8 | 28.356 | 45.0293 | 231.276 | 761.557 | 120.232 | -3.0286 | -16.7486 | 41.9007 | 51.16 | 15.2083 | 18.8001 | 0.25498 | 583.121 | 33.1626 | 40
39616.6194 | 2008-06-17 | #NAME? | 0.052175 | 0.23595 | 0.28812 | 90308.8 | 28.4467 | 45.0293 | 229.963 | 762.344 | 121.02 | -3.0004 | -16.7285 | 41.9007 | 51.16 | 15.2083 | 18.8001 | 0.23539 | 583.121 | 32.8611 | 40
39616.6201 | 2008-06-17 | #NAME? | 0.18577 | 0.84808 | 1.0338 | 90294.8 | 28.6508 | 45.1923 | 228.651 | 759.719 | 118.394 | -3.0004 | -16.7753 | 41.9007 | 51.16 | 15.2083 | 18.8001 | 0.24862 | 582.751 | 33.1626 | 40
39616.6208 | 2008-06-17 | #NAME? | 0.18056 | 0.84735 | 1.0279 | 90277.8 | 28.6281 | 44.9078 | 228.651 | 756.831 | 121.02 | -3.038 | -16.7486 | 41.9007 | 51.16 | 15.2083 | 18.8001 | 0.23564 | 582.751 | 32.8611 | 40
39616.6215 | 2008-06-17 | #NAME? | 0.15693 | 0.84971 | 1.0066 | 90272 | 28.5828 | 45.0293 | 233.376 | 756.306 | 119.969 | -3.0098 | -16.7419 | 41.9007 | 51.16 | 15.2083 | 18.8001 | 0.24837 | 582.751 | 32.8611 | 40
39616.6222 | 2008-06-17 | #NAME? | 0.085582 | 0.89482 | 0.98041 | 90270.3 | 28.5601 | 44.997 | 236.789 | 757.619 | 119.444 | -3.0192 | -16.7285 | 41.9007 | 51.16 | 15.2083 | 18.8001 | 0.23361 | 583.121 | 32.8611 | 40
39616.6229 | 2008-06-17 | #NAME? | 0.028416 | 0.91298 | 0.94139 | 90273.2 | 28.4467 | 44.8483 | 236.526 | 760.244 | 121.545 | -3.0098 | -16.7285 | 41.9007 | 51.16 | 15.2083 | 18.8001 | 0.2476 | 582.751 | 32.8611 | 40
39616.6236 | 2008-06-17 | #NAME? | 0.0079471 | 0.90067 | 0.90861 | 90271.9 | 28.4467 | 45.1597 | 235.214 | 762.344 | 118.394 | -3.0192 | -16.7619 | 41.9007 | 51.16 | 15.2083 | 18.8001 | 0.23514 | 583.121 | 32.8611 | 40
39616.6243 | 2008-06-17 | #NAME? | -0.0043747 | 0.89634 | 0.89197 | 90274.2 | 28.356 | 44.9078 | 233.376 | 761.294 | 120.495 | -3.0098 | -16.7486 | 41.9007 | 51.16 | 15.2083 | 18.8001 | 0.24684 | 583.121 | 32.8611 | 40
39616.625 | 2008-06-17 | #NAME? | 0.0043823 | 0.86979 | 0.87417 | 90276.1 | 28.3787 | 44.997 | 232.326 | 757.882 | 120.495 | -3.0192 | -16.7419 | 41.9007 | 51.16 | 15.2083 | 18.8001 | 0.23488 | 582.751 | 32.8611 | 40

## Det_Parse
Detector Responses
Date & Time | Conc | Reference | Candidate | Zero
2008-06-17 12:24 | 0.0099562
2008-06-17 12:25 | -0.0022099 |  |  | -0.0022099
2008-06-17 12:26 | -0.0015078 |  |  | -0.0015078
2008-06-17 12:27 | -0.0030689 |  |  | -0.0030689
2008-06-17 12:28 | -0.0019485 |  |  | -0.0019485
2008-06-17 12:29 | 0.54076
2008-06-17 12:30 | 5.8181
2008-06-17 12:31 | 7.9336
2008-06-17 12:32 | 8.0496 | 8.0496
2008-06-17 12:33 | 8.0431 | 8.0431
2008-06-17 12:34 | 8.0556 | 8.0556
2008-06-17 12:35 | 8.0627 | 8.0627
2008-06-17 12:36 | 8.1658
2008-06-17 12:37 | 8.6765
2008-06-17 12:38 | 8.7494 |  | 8.7494
2008-06-17 12:39 | 8.7747 |  | 8.7747
2008-06-17 12:40 | 8.7806 |  | 8.7806
2008-06-17 12:41 | 8.7778 |  | 8.7778
2008-06-17 12:42 | 8.6256
2008-06-17 12:43 | 8.1607
2008-06-17 12:44 | 8.1357 | 8.1357
2008-06-17 12:45 | 8.1401 | 8.1401
2008-06-17 12:46 | 8.1142 | 8.1142
2008-06-17 12:47 | 8.1082 | 8.1082
2008-06-17 12:48 | 8.1856
2008-06-17 12:49 | 8.626
2008-06-17 12:50 | 8.6922 |  | 8.6922
2008-06-17 12:51 | 8.6614 |  | 8.6614
2008-06-17 12:52 | 8.6721 |  | 8.6721
2008-06-17 12:53 | 8.6828 |  | 8.6828
2008-06-17 12:54 | 8.6811
2008-06-17 12:55 | 8.7126
2008-06-17 12:56 | 8.5516
2008-06-17 12:57 | 8.117
2008-06-17 12:58 | 8.0887 | 8.0887
2008-06-17 12:59 | 8.0699 | 8.0699
2008-06-17 13:00 | 8.0772 | 8.0772
2008-06-17 13:01 | 8.0546 | 8.0546
2008-06-17 13:02 | 8.2164
2008-06-17 13:03 | 8.662
2008-06-17 13:04 | 8.6784 |  | 8.6784
2008-06-17 13:05 | 8.6799 |  | 8.6799
2008-06-17 13:06 | 8.6689 |  | 8.6689
2008-06-17 13:07 | 8.6705 |  | 8.6705
2008-06-17 13:08 | 8.4808
2008-06-17 13:09 | 8.0131
2008-06-17 13:10 | 7.995 | 7.995
2008-06-17 13:11 | 7.9876 | 7.9876
2008-06-17 13:12 | 7.9986 | 7.9986
2008-06-17 13:13 | 8.0245 | 8.0245
2008-06-17 13:14 | 8.0157
2008-06-17 13:15 | 6.8872
2008-06-17 13:16 | 5.6372
2008-06-17 13:17 | 5.5968
2008-06-17 13:18 | 5.5931 | 5.5931
2008-06-17 13:19 | 5.5857 | 5.5857
2008-06-17 13:20 | 5.5865 | 5.5865
2008-06-17 13:21 | 5.5775 | 5.5775
2008-06-17 13:22 | 5.5761
2008-06-17 13:23 | 5.9018
2008-06-17 13:24 | 6.1306 |  | 6.1306
2008-06-17 13:25 | 6.1306 |  | 6.1306
2008-06-17 13:26 | 6.1257 |  | 6.1257
2008-06-17 13:27 | 6.1282 |  | 6.1282
2008-06-17 13:28 | 6.1284
2008-06-17 13:29 | 6.0304
2008-06-17 13:30 | 5.6306 | 5.6306
2008-06-17 13:31 | 5.5807 | 5.5807
2008-06-17 13:32 | 5.5882 | 5.5882
2008-06-17 13:33 | 5.5882 | 5.5882
2008-06-17 13:34 | 5.7397
2008-06-17 13:35 | 6.1352
2008-06-17 13:36 | 6.1466 |  | 6.1466
2008-06-17 13:37 | 6.1463 |  | 6.1463
2008-06-17 13:38 | 6.1308 |  | 6.1308
2008-06-17 13:39 | 6.135 |  | 6.135
2008-06-17 13:40 | 5.9948
2008-06-17 13:41 | 5.5758
2008-06-17 13:42 | 5.5396 | 5.5396
2008-06-17 13:43 | 5.5231 | 5.5231
2008-06-17 13:44 | 5.5255 | 5.5255
2008-06-17 13:45 | 5.5138 | 5.5138
2008-06-17 13:46 | 5.6402
2008-06-17 13:47 | 6.0829
2008-06-17 13:48 | 6.135 |  | 6.135
2008-06-17 13:49 | 6.1463 |  | 6.1463
2008-06-17 13:50 | 6.1501 |  | 6.1501
2008-06-17 13:51 | 6.1568 |  | 6.1568
2008-06-17 13:52 | 6.1573
2008-06-17 13:53 | 5.8555
2008-06-17 13:54 | 5.5526 | 5.5526
2008-06-17 13:55 | 5.5452 | 5.5452
2008-06-17 13:56 | 5.544 | 5.544
2008-06-17 13:57 | 5.5384 | 5.5384
2008-06-17 13:58 | 4.6214
2008-06-17 13:59 | 2.4699
2008-06-17 14:00 | 2.5287
2008-06-17 14:01 | 2.5253 | 2.5253
2008-06-17 14:02 | 2.52 | 2.52
2008-06-17 14:03 | 2.5182 | 2.5182
2008-06-17 14:04 | 2.5161 | 2.5161
2008-06-17 14:05 | 2.6264
2008-06-17 14:06 | 2.8996 |  | 2.8996
2008-06-17 14:07 | 2.9135 |  | 2.9135
2008-06-17 14:08 | 2.9175 |  | 2.9175
2008-06-17 14:09 | 2.9249 |  | 2.9249
2008-06-17 14:10 | 2.8244
2008-06-17 14:11 | 2.5486 | 2.5486
2008-06-17 14:12 | 2.5254 | 2.5254
2008-06-17 14:13 | 2.5315 | 2.5315
2008-06-17 14:14 | 2.5301 | 2.5301
2008-06-17 14:15 | 2.6302
2008-06-17 14:16 | 2.9086 |  | 2.9086
2008-06-17 14:17 | 2.9242 |  | 2.9242
2008-06-17 14:18 | 2.9124 |  | 2.9124
2008-06-17 14:19 | 2.9154 |  | 2.9154
2008-06-17 14:20 | 2.8316
2008-06-17 14:21 | 2.5323 | 2.5323
2008-06-17 14:22 | 2.5 | 2.5
2008-06-17 14:23 | 2.5008 | 2.5008
2008-06-17 14:24 | 2.491 | 2.491
2008-06-17 14:25 | 2.5856
2008-06-17 14:26 | 2.8842 |  | 2.8842
2008-06-17 14:27 | 2.9146 |  | 2.9146
2008-06-17 14:28 | 2.9183 |  | 2.9183
2008-06-17 14:29 | 2.9165 |  | 2.9165
2008-06-17 14:30 | 2.8371
2008-06-17 14:31 | 2.5431 | 2.5431
2008-06-17 14:32 | 2.5222 | 2.5222
2008-06-17 14:33 | 2.5305 | 2.5305
2008-06-17 14:34 | 2.5289 | 2.5289
2008-06-17 14:35 | 2.0146
2008-06-17 14:36 | 0.1857
2008-06-17 14:37 | -0.004071 |  |  | -0.004071
2008-06-17 14:38 | -0.0045309 |  |  | -0.0045309
2008-06-17 14:39 | -0.0043467 |  |  | -0.0043467
2008-06-17 14:40 | -0.0033104 |  |  | -0.0033104
2008-06-17 14:41 | -0.003758
2008-06-17 14:42 | 2.1441
2008-06-17 14:43 | 9.196
2008-06-17 14:44 | 4.1435
2008-06-17 14:45 | 1.9843
2008-06-17 14:46 | 1.6048
2008-06-17 14:47 | 1.3986
2008-06-17 14:48 | 1.2612
2008-06-17 14:49 | 1.1691
2008-06-17 14:50 | 0.6476
2008-06-17 14:51 | 0
2008-06-17 14:52 | 0.28812
2008-06-17 14:53 | 1.0338
2008-06-17 14:54 | 1.0279
2008-06-17 14:55 | 1.0066
2008-06-17 14:56 | 0.98041
2008-06-17 14:57 | 0.94139
2008-06-17 14:58 | 0.90861
2008-06-17 14:59 | 0.89197
2008-06-17 15:00 | 0.87417
Formulas:
Date & Time = IF([Date & Time, row above]<Test_Info!$C9,MIN(IF(Detector_Time>[Date & Time, row above],Detector_Time,"")),NA())
Recalculate

## Point-1
Single-Point Bracketing Data
Reference Setpoint |  |  | 8.1 | mg/m3 |  | Bracketing Procedure Start |  | 2008-06-17 12:24
Reference Concentration |  |  | 8.06 | mg/m3
[6 empty] | Bracketing Procedure End |  | 2008-06-17 14:40
Candidate Setpoint |  |  | 8.1 | mg/m3
Date/Time |  | Meas. Type | Detector Responses |  |  |  |  | Ref/Cand Ratios |  |  | Detector Measurement Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-06-17 12:27 |  | Zero | -0.0021838 | [7 empty] | 2008-06-17 12:25 | 2008-06-17 12:28
2008-06-17 12:34 |  | Ref | 8.0527 | 8.055 | 4 | 0.020988 |  |  |  |  | 2008-06-17 12:32 | 2008-06-17 12:35
2008-06-17 12:40 |  | Cand | 8.7706 | 8.773 | 4 | 0.014133 | 8.091 | 1.0843 | 0.002647 |  | 2008-06-17 12:38 | 2008-06-17 12:41
2008-06-17 12:45 |  | Ref | 8.1245 | 8.127 | 4 | 0.020988 |  |  |  |  | 2008-06-17 12:44 | 2008-06-17 12:47
2008-06-17 12:52 |  | Cand | 8.6771 | 8.6797 | 4 | 0.014133 | 8.1048 | 1.0709 | 0.0026384 |  | 2008-06-17 12:50 | 2008-06-17 12:53
2008-06-17 12:59 |  | Ref | 8.0726 | 8.0753 | 4 | 0.020988 |  |  |  |  | 2008-06-17 12:58 | 2008-06-17 13:01
2008-06-17 13:05 |  | Cand | 8.6744 | 8.6772 | 4 | 0.014133 | 8.0398 | 1.0793 | 0.0026569 |  | 2008-06-17 13:04 | 2008-06-17 13:07
2008-06-17 13:12 |  | Ref | 8.0014 | 8.0043 | 4 | 0.020988 |  |  |  |  | 2008-06-17 13:10 | 2008-06-17 13:13
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
2008-06-17 14:39 |  | Zero | -0.0040647 | [7 empty] | 2008-06-17 14:37 | 2008-06-17 14:40
Bracketing Intervals
Detector Linearity | Linearity Factor at 8.1 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 0.028265 | 2008-06-17 12:38 | 2008-06-17 13:07
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 0.041976 | 2008-06-17 12:32 | 2008-06-17 13:13
| Ratio Average Uncertainty Components, 1s |  |  | Detector Linearity |  |  |  | 0
|  |  |  | Measurement Stability |  |  |  | 0.0015285
|  |  |  | Ratio Repeatability |  |  |  | 0.0038729
|  | Set Average Candidate/Reference Ratio |  |  |  |  | 1.0782
|  | Number of Brackets in this Set |  |  |  |  | 3
|  | Uncertainty for This Set's Average, 1s |  |  |  |  | 0.0041636
|  | Nested "Level-2" StdErr, s2 |  |  |  |  | 0.0067514
Formulas:
Date/Time = AVERAGE([Detector Measurement Intervals],[Last Record])
Date/Time = IF([SE(mavg)]>0,AVERAGE([Detector Measurement Intervals],[Last Record])," ")
Detector Measurement Intervals = MIN(IF(D_Ref="","",IF(D_Time>[Detector Measurement Intervals, row above],D_Time,"")))
Detector Measurement Intervals = MIN(IF(D_Cand="","",IF(D_Time>[Detector Measurement Intervals, row above],D_Time,"")))
Last Record = MIN(IF(D_Zero="",IF(D_Time>[Detector Measurement Intervals],D_Time1,""),""))
Last Record = MIN(IF(D_Ref="",IF(D_Time>[Detector Measurement Intervals],D_Time1,""),""))
Note on A8: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on D3: Ed Brown: The "nominal" reference setpoint, used for identification, not for calculation
Note on D4: Ed Brown: Actual certified reference concentration
Note on D6: Ed Brown: The "nominal" candidate setpoint, used for identification, not for calculation
Note on D9: Ed Brown: Average response for measurement interval
Note on E9: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E27: Ed Brown: Value from Linearity Chack reponse table
Note on F10: Ed Brown: Number of individual readings in measurement interval
Note on G10: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on H9: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H36: Ed Brown: The "End Result" of this bracketing procedure
Note on H38: Ed Brown: Uncertainty of the ratio result of this bracketing procedure
Note on I9: Ed Brown: Calculated measurement ratio for bracket
Note on J9: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on K27: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K28: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on M26: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Populate Parse Table

## Point-2
Single-Point Bracketing Data
Reference Setpoint |  |  | 5.7 | mg/m3 |  | Bracketing Procedure Start |  | 2008-06-17 12:24
Reference Concentration |  |  | 5.35 | mg/m3
[6 empty] | Bracketing Procedure End |  | 2008-06-17 14:40
Candidate Setpoint |  |  | 5.7 | mg/m3
Date/Time |  | Meas. Type | Detector Responses |  |  |  |  | Ref/Cand Ratios |  |  | Detector Measurement Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-06-17 12:27 |  | Zero | -0.0021838 | [7 empty] | 2008-06-17 12:25 | 2008-06-17 12:28
2008-06-17 13:19 |  | Ref | 5.5857 | 5.5887 | 4 | 0.011805 |  |  |  |  | 2008-06-17 13:18 | 2008-06-17 13:21
2008-06-17 13:25 |  | Cand | 6.1288 | 6.1318 | 4 | 0.0034228 | 5.5943 | 1.0961 | 0.0017462 |  | 2008-06-17 13:24 | 2008-06-17 13:27
2008-06-17 13:32 |  | Ref | 5.5969 | 5.6 | 4 | 0.011805 |  |  |  |  | 2008-06-17 13:30 | 2008-06-17 13:33
2008-06-17 13:37 |  | Cand | 6.1397 | 6.1429 | 4 | 0.0034228 | 5.5644 | 1.104 | 0.0017668 |  | 2008-06-17 13:36 | 2008-06-17 13:39
2008-06-17 13:44 |  | Ref | 5.5255 | 5.5288 | 4 | 0.011805 |  |  |  |  | 2008-06-17 13:42 | 2008-06-17 13:45
2008-06-17 13:50 |  | Cand | 6.1471 | 6.1504 | 4 | 0.0034228 | 5.5386 | 1.1105 | 0.0017841 |  | 2008-06-17 13:48 | 2008-06-17 13:51
2008-06-17 13:55 |  | Ref | 5.5451 | 5.5485 | 4 | 0.011805 |  |  |  |  | 2008-06-17 13:54 | 2008-06-17 13:57
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
2008-06-17 14:39 |  | Zero | -0.0040647 | [7 empty] | 2008-06-17 14:37 | 2008-06-17 14:40
Bracketing Intervals
Detector Linearity | Linearity Factor at 5.7 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 0.0068456 | 2008-06-17 13:24 | 2008-06-17 13:51
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 0.02361 | 2008-06-17 13:18 | 2008-06-17 13:57
| Ratio Average Uncertainty Components, 1s |  |  | Detector Linearity |  |  |  | 0
|  |  |  | Measurement Stability |  |  |  | 0.0010195
|  |  |  | Ratio Repeatability |  |  |  | 0.0041483
|  | Set Average Candidate/Reference Ratio |  |  |  |  | 1.1035
|  | Number of Brackets in this Set |  |  |  |  | 3
|  | Uncertainty for This Set's Average, 1s |  |  |  |  | 0.0042717
|  | Nested "Level-2" StdErr, s2 |  |  |  |  | 0.007203
Formulas:
Date/Time = AVERAGE([Detector Measurement Intervals],[Last Record])
Date/Time = IF([SE(mavg)]>0,AVERAGE([Detector Measurement Intervals],[Last Record])," ")
Detector Measurement Intervals = MIN(IF(D_Cand="","",IF(D_Time>[Detector Measurement Intervals, row above],D_Time,"")))
Detector Measurement Intervals = MIN(IF(D_Ref="","",IF(D_Time>[Detector Measurement Intervals, row above],D_Time,"")))
Last Record = MIN(IF(D_Zero="",IF(D_Time>[Detector Measurement Intervals],D_Time1,""),""))
Last Record = MIN(IF(D_Ref="",IF(D_Time>[Detector Measurement Intervals],D_Time1,""),""))
Note on A8: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on D3: Ed Brown: The "nominal" reference setpoint, used for identification, not for calculation
Note on D4: Ed Brown: Actual certified reference concentration
Note on D6: Ed Brown: The "nominal" candidate setpoint, used for identification, not for calculation
Note on D9: Ed Brown: Average response for measurement interval
Note on E9: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E27: Ed Brown: Value from Linearity Chack reponse table
Note on F10: Ed Brown: Number of individual readings in measurement interval
Note on G10: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on H9: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H36: Ed Brown: The "End Result" of this bracketing procedure
Note on H38: Ed Brown: Uncertainty of the ratio result of this bracketing procedure
Note on I9: Ed Brown: Calculated measurement ratio for bracket
Note on J9: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on K27: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K28: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on M26: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Populate Parse Table

## Point-3
Single-Point Bracketing Data
Reference Setpoint |  |  | 2.7 | mg/m3 |  | Bracketing Procedure Start |  | 2008-06-17 12:24
Reference Concentration |  |  | 2.44 | mg/m3
[6 empty] | Bracketing Procedure End |  | 2008-06-17 14:40
Candidate Setpoint |  |  | 2.7 | mg/m3
Date/Time |  | Meas. Type | Detector Responses |  |  |  |  | Ref/Cand Ratios |  |  | Detector Measurement Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-06-17 12:27 |  | Zero | -0.0021838 | [7 empty] | 2008-06-17 12:25 | 2008-06-17 12:28
2008-06-17 14:02 |  | Ref | 2.5199 | 2.5234 | 4 | 0.0079057 |  |  |  |  | 2008-06-17 14:01 | 2008-06-17 14:04
2008-06-17 14:08 |  | Cand | 2.9139 | 2.9175 | 4 | 0.0057918 | 2.5305 | 1.1529 | 0.0034243 |  | 2008-06-17 14:06 | 2008-06-17 14:09
2008-06-17 14:13 |  | Ref | 2.5339 | 2.5376 | 4 | 0.0079057 |  |  |  |  | 2008-06-17 14:11 | 2008-06-17 14:14
2008-06-17 14:17 |  | Cand | 2.9152 | 2.9189 | 4 | 0.0057918 | 2.5237 | 1.1566 | 0.0034396 |  | 2008-06-17 14:16 | 2008-06-17 14:19
2008-06-17 14:22 |  | Ref | 2.506 | 2.5099 | 4 | 0.0079057 |  |  |  |  | 2008-06-17 14:21 | 2008-06-17 14:24
2008-06-17 14:28 |  | Cand | 2.9084 | 2.9123 | 4 | 0.0057918 | 2.5225 | 1.1545 | 0.0034378 |  | 2008-06-17 14:26 | 2008-06-17 14:29
2008-06-17 14:33 |  | Ref | 2.5312 | 2.5352 | 4 | 0.0079057 |  |  |  |  | 2008-06-17 14:31 | 2008-06-17 14:34
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
2008-06-17 14:39 |  | Zero | -0.0040647 | [7 empty] | 2008-06-17 14:37 | 2008-06-17 14:40
Bracketing Intervals
Detector Linearity | Linearity Factor at 2.7 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 0.011584 | 2008-06-17 14:06 | 2008-06-17 14:29
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 0.015811 | 2008-06-17 14:01 | 2008-06-17 14:34
| Ratio Average Uncertainty Components, 1s |  |  | Detector Linearity |  |  |  | 0
|  |  |  | Measurement Stability |  |  |  | 0.0019825
|  |  |  | Ratio Repeatability |  |  |  | 0.00089269
|  | Set Average Candidate/Reference Ratio |  |  |  |  | 1.1547
|  | Number of Brackets in this Set |  |  |  |  | 3
|  | Uncertainty for This Set's Average, 1s |  |  |  |  | 0.0021743
|  | Nested "Level-2" StdErr, s2 |  |  |  |  | 0.0018367
Formulas:
Date/Time = AVERAGE([Detector Measurement Intervals],[Last Record])
Date/Time = IF([SE(mavg)]>0,AVERAGE([Detector Measurement Intervals],[Last Record])," ")
Detector Measurement Intervals = MIN(IF(D_Cand="","",IF(D_Time>[Detector Measurement Intervals, row above],D_Time,"")))
Detector Measurement Intervals = MIN(IF(D_Ref="","",IF(D_Time>[Detector Measurement Intervals, row above],D_Time,"")))
Last Record = MIN(IF(D_Zero="",IF(D_Time>[Detector Measurement Intervals],D_Time1,""),""))
Last Record = MIN(IF(D_Ref="",IF(D_Time>[Detector Measurement Intervals],D_Time1,""),""))
Note on A8: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on D3: Ed Brown: The "nominal" reference setpoint, used for identification, not for calculation
Note on D4: Ed Brown: Actual certified reference concentration
Note on D6: Ed Brown: The "nominal" candidate setpoint, used for identification, not for calculation
Note on D9: Ed Brown: Average response for measurement interval
Note on E9: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E27: Ed Brown: Value from Linearity Chack reponse table
Note on F10: Ed Brown: Number of individual readings in measurement interval
Note on G10: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on H9: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H36: Ed Brown: The "End Result" of this bracketing procedure
Note on H38: Ed Brown: Uncertainty of the ratio result of this bracketing procedure
Note on I9: Ed Brown: Calculated measurement ratio for bracket
Note on J9: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on K27: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K28: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on M26: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Populate Parse Table

## Results
Bracketing Procedure Results Summary
Operator: |  | Unspecified |  |  | Candidate ID: |  | DL57TB
Company/Lab: |  | Unspecified |  |  | Candidate S/N: |  | Unspecified
Start Date/Time: |  | 2008-06-17 12:24 |  |  | Reference ID: |  | PO24RI
End Date/Time: |  | 2008-06-17 15:00 |  |  | Reference S/N: |  | 0803027443
Reference Generator Certification Parameters |  |  | Bracketing Parameters |  |  |  | Candidate Generator Parameters
|  |  | Uncertainty Terms |  | Ratio Measurements
|  |  | StdErr Among Brackets S2 | Uncertainty of Mean Ratio u(Rmean)
Setpoint ug/scm | Certified ug/scm | Uncertainty 1s ug/scm |  |  | No. of Brackets K | Average Ratio Rmean | Setpoint ug/scm | Bracketing (Ccand)j ug/scm | Spreadsheet Layer
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | Point-1
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | Point-2
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | Point-3
Note on A11: Ed Brown: The "nominal" setpoint, used for identification, not for calculation
Note on H11: Ed Brown: The "nominal" setpoint, used for identification, not for calculation