## Read Me
Instructions for Bracketing & Certification Workbook
This workbook is designed to process the data from a single "bracketing procedure" comparison
between a NIST-traceable "reference" gas generator and "candidate" generator going through
certification. The comparisons are based on the "point-to-point" method of propagating traceability,
where generators are certified at specific setpoints that correspond to the setpoints of the NIST
standard. The generator output at each setpoint is a separate NIST-traceable quantity, independent
of the other setpoint outputs.
The main inputs to this workbook are the reference generator certification and the data files from
the bracketing procedure. The latter input files are recorded during the bracketing by the candidate
generator, reference generator, and mercury detector software. The outputs of the bracketing procedure
are carried forward into a calculation of uncertainty which includes a default ("Type-B") uncertainty component
for the estimated reproducibility of the bracketing results. It is important to note that this workbook is only
intended for certification of field-installed generators (i.e., where the cost of repeating the bracketing
experiment is prohibitive), and should never be used to certify a vendor-prime or field-reference generator.
The workbook is not structurally limited in the number of setpoints or "brackets per setpoint" it
can accommodate. The prototype was designed for a maximum of 7 setpoints and 6 brackets per setpoint,
but the calculations are not disrupted by use of smaller datasets. Experienced Excel users may customize
the "dimensions" of the workbook to their specific application.
All layers that perform calculations have protection enabled, to prevent users from accidentally
overwriting the formulae. User input cells are unprotected, and use a blue display font. Some calculation
cells, mainly those using functions that lookup values in the raw data layers, use a dark blue font to indicate
that the user may overwrite the formulae or values in these cells. No password is required to remove the
protection from any layer in this workbook.
1. Workbook Organization
The workbook consists of several worksheet layers , which are displayed as tabs at the bottom
of the screen. Detailed descriptions of each layer are included below the step-by-step instructions.
Read Me (this sheet) - describes the workbook, explaining how to use the worksheets.
Test_Info - allows the user to input the basic information about the bracketing procedure, the
reference gas generator, and the detector linearity check.
Cand_Raw & Ref_Raw - allows the user to paste the raw data file from both the "Candidate"
and "Reference" gas generators. This version of the workbook compares two Tekran 3310
generators, where the "display" value is recorded in the 47th column (Excel column "AU").
Detector_Raw - allows the user to paste the raw data file from the detector, for this version a
Tekran 2537 with two parallel measurement channels ("A" and "B") operating.
AB-Split - parses the 2537 raw data into two separate datastreams based on the measurement
channel. This layer reserves additional columns for the user to structure the data in a way that allows
other layers to lookup timestamps for individual measurement intervals.
Point-1, Point-2, etc. - These are the "setpoint" layers, which may have different names but are
always behind the AB-Split layer. Each of these layers contains the formulae necessary to report the
average comparison ratio and related uncertainty components for that setpoint.
Bracketing Results - Re-assembles the data from other layers into a table summarizing the
bracketing procedure results and uncertainty components.
Uncertainty - calculates the overall combined uncertainty of the candidate gas concentrations,
based all of the component uncertainties (e.g., reference gas, bracketing, reproducibility).
Certification - Tabulates the certification and uncertainty data.
2. Step by Step
Step 1 -- Fill in the input cells (blue font) of the Test_Info layer.
Step 2 -- Populate the Cand_Raw and Ref_Raw layers with 3310 data (optional).
Step 3 -- Populate the Detector_Raw layer with 2537 data.
Step 4 -- On the AB_Split layer, click on the "Recalculate" button to populate columns A-B and G-H.
Step 5a -- Populate columns C-E and I-K of the AB_Split layer based on measurement intervals OR
Step 5b -- Fill in the "Detector Measurement Intervals" tables (columns M & O of the setpoint layers)
and execute the Ctrl-P macro (once on each setpoint layer) to populate the AB_Split layer
Step 6 -- Ensure that all header information on the setpoint layers is accurate.
Step 7 -- Inspect the "Detector Measurement Intervals" tables and correct any inaccuracies. Most
table entries can be made using the Ctrl-R macro, provided the AB_Split layer is populated
correctly, the cell pointer is in the cell to be filled, and all earlier timestamp entries are valid.
Step 8 -- Fill in the names of the spreadsheet layers in column K of the Bracketing_Results layer and
press F9 to recalculate all workbook formulae.
Step 9 -- Inspect all layers before saving the workbook with a new name.
3. Test Info
The blue font cells are the input ranges. The top-center box contains general test information. Any of
these cells can be linked to other workbooks, or to "_Raw" layers within this workbook (referencing
calculation layers may risk creating a "circular reference" problem). It is common practice, for
example, to use a simple pointing formula (e.g. "=Detector_Raw!A2") to populate the "Date/Time"
cells from one of the "_Raw" layers.
The bottom-center box receives information about the reference generator certifications. This table
may be populated by file links or, if the layout is the same, pasting the table directly from the source
file. The values in columns I and J are optional, included only to match the customary certification
output format. To either side of the reference generator certification data are the "A" and "B" channel
detector linearity data. Note that the response units are not specified, but these linearity check
responses need to be reported in the same units as used in the "Detector_Raw" layer.
4. Cand_Raw & Ref_Raw |  |  |  | (These layers are not protected)
In order to input the raw data from the "Candidate" and "Reference" Tekran 3310 calibrators, the user
must import the trending files created by the Tekran software. These files are usually stored in the
"C:\Program Files\Tekran\StackCem\Data\Trending" directory on the Tekran SCADA computer, and
are named "Trend yyyy-mm-dd.csv." Each file must be opened in Microsoft Excel, which will arrange
it as a workbook. Highlight all the data and click "Copy". Open the appropriate "_Raw" sheet and
"right" click on cell A1. Choose the "Paste" option. The raw data file should now be pasted into the
sheet. Please ensure that the "DATETIME" column is Column A, and "CALIB_CONCEPA" is in
Column AU. Any data in columns beyond Column AU do not pertain to the 3310 Calibrator, and can be
deleted to reduce file size. If the bracketing procedure spans more than one day, repeat this
procedure and paste the additional records (minus the header row) at the bottom of the same layer.
5. Detector_Raw |  |  | (This layer is not protected)
Depending on the experimental setup used during bracketing, the detector responses may be
recorded in the same raw data file as one of the generators. If this is the case, the "Detector_Raw"
layer is extraneous, and may be deleted after the named ranges are redefined (see below).
Otherwise, this data is input using the same procedure as the other two "_Raw" layers above. It
should be noted that if Tekran software other than the CEMS SCADA system is used to record
detector responses, the file format will be different. Depending on the filetype, Excel may display an
additional dialog box, prompting the user to open as "Fixed Width" or "Delimited". In this case, click on
the "Delimited" tab and hit "Next". At the next prompt, click on the "Space" button and then click
"Finish". The file should open up in a separate worksheet. Highlight all the data and use "Copy" and
"Paste" similarly to the SCADA software files.
Check whether the date and time are recorded in separate columns, or in a single column. If column A
does not contain both the date and time, it must be altered. The workbook includes a macro
for this purpose, which is executed using the Ctrl-T keystroke combination.
This layer contains three named ranges for the data columns, defined as follows:
| Range Name |  | Column |  | Example Range Specification
| Detector_Time |  | Time/Date |  | =OFFSET(Detector_Raw!$A$1,1,0,COUNTA(Detector_Raw!$A:$A)-1)
| Detector_AB |  | Channel: A or B |  | =OFFSET(Detector_Raw!$D$1,1,0,COUNTA(Detector_Raw!$A:$A)-1)
| Detector_Table |  | Entire Table |  | =OFFSET(Detector_Raw!$A$1,1,0,COUNTA(Detector_Raw!$A:$A)-1,13)
The example specifications essentially define named ranges whose heights extend to the last
occupied row of the layer.
6. AB-Split
This layer is used to process the data from the "Detector_Raw" layer into a form more useable by
the setpoint layers (i.e., Point-1, Point-2, etc.). The formulae used in this layer are as follows:
| =Test_Info!$G$8 |  | used in cells A2 and G2 (the first row of the two "Date & Time" columns)
$A4=IF($A3<Test_Info!$G$9,MIN(IF(Detector_AB="A",IF(Detector_Time>$A3,Detector_Time,""),"")),NA())
$G4=IF($G3<Test_Info!$G$9,MIN(IF(Detector_AB="B",IF(Detector_Time>$G3,Detector_Time,""),"")),NA())
These functions populate the two distinct columns of time/date stamps based on which detector
channel was active during each record of the raw data file.
B3=VLOOKUP(A3,Detector_Table,13)
H3=VLOOKUP(G3,Detector_Table,13)
These functions populate columns to the right of the time stamps with detector response values.
In the prototype spreadsheet, rows 5 through the bottom of the table have been converted into values
(i.e., no formulae) to reduce the workbook's ongoing computational complexity. To re-calculate or re-
populate this layer, the formulae in row 4 must be copied down beyond the end of the data, the
workbook recalculated, and the non-numeric rows deleted from the bottom of the layer. The workbook
includes a macro, executed by clicking the "Recalculate" button, that performs all of these steps and
converts rows 5 through the bottom into values.
Columns C-E and/or I-K must be populated by the user, by copying response values over into the
appropriate column during the measurement intervals of the bracketing procedure. The prototype
workbook contains several calculations in the setpoint layers that make use of the entries in columns
C-E and I-K, all of which assume that these columns contain only "stable" response data. Any
response representing "transitional" data should not be included, even if this excludes all data from one
of the two measurement channels.
As an alternative to populating the columns, the user may choose to enter the interval start and stop
times directly into the tables in the setpoint layers. The workbook includes a macro, executed with
the Ctrl-P keystroke combination, that will automatically populate the C-E and I-K columns of the
AB-Split layer based on the current layer's interval tables. This macro must be executed for each
setpoint layer after the interval start and stop times are entered in that layer.
The AB_Split layer contains twelve named ranges for the data columns, defined as follows:
| Range Name |  | Column |  | Range Specification
| A_Time |  | A |  | =OFFSET('AB-Split'!$A$2,1,0,COUNTA('AB-Split'!$A:$A)-1)
| A_Time1 |  | A |  | =OFFSET('AB-Split'!$A$2,0,0,COUNTA('AB-Split'!$A:$A)-1)
| A_Conc |  | B |  | =OFFSET('AB-Split'!$B$2,1,0,COUNTA('AB-Split'!$A:$A)-1)
| A_Ref |  | C |  | =OFFSET('AB-Split'!$C$2,1,0,COUNTA('AB-Split'!$A:$A)-1)
| A_Cand |  | D |  | =OFFSET('AB-Split'!$D$2,1,0,COUNTA('AB-Split'!$A:$A)-1)
| A_Zero |  | E |  | =OFFSET('AB-Split'!$E$2,1,0,COUNTA('AB-Split'!$A:$A)-1)
| B_Time |  | G |  | =OFFSET('AB-Split'!$G$2,1,0,COUNTA('AB-Split'!$A:$A)-1)
| B_Time1 |  | G |  | =OFFSET('AB-Split'!$G$2,0,0,COUNTA('AB-Split'!$A:$A)-1)
| B_Conc |  | H |  | =OFFSET('AB-Split'!$H$2,1,0,COUNTA('AB-Split'!$A:$A)-1)
| B_Ref |  | I |  | =OFFSET('AB-Split'!$I$2,1,0,COUNTA('AB-Split'!$A:$A)-1)
| B_Cand |  | J |  | =OFFSET('AB-Split'!$J$2,1,0,COUNTA('AB-Split'!$A:$A)-1)
| B_Zero |  | K |  | =OFFSET('AB-Split'!$K$2,1,0,COUNTA('AB-Split'!$A:$A)-1)
7. Setpoints: Point-1, Point-2, etc.
These layers perform, for each setpoint, all calculations related to bracketing ratio and uncertainty
components. The "input" area of this layer is to the right, where the interval start and stop times are
tabulated. The prototype spreadsheet has this table populated with formulae to calculate the times
based on the user's entries in columns C-E and I-K of the "AB-Split" layer (notwithstanding the option
to enter the times directly in these layers and populate the AB-Split columns using the Ctrl-P macro,
as mentioned above). It is important to inspect this table carefully, overwriting the formulae where
necessary to have the table accurately reflect the first and last record of each measurement interval.
Often, once the first "valid" timestamp is entered near the top of the table, the formulae will calculate
the remaining timestamps. Missing or overwritten formulae in this layer can be replaced by placing the
cell pointer in the affected cell and pressing the Ctrl-R keystroke combination.
Once all timestamps are entered correctly, most of the calculations should proceed to completion
with a simple press of the {Calc} (F9) key. If, while investigating abnormal results, it is found that a
particular measurement interval needs to be "dropped" from the set, do so by simply overwriting
text comments (or simply a space) into the associated timestamp cells
8. Bracketing Results
This layer is a simple lookup table that brings all of the data from the other layers into one table. The
only required input is the names of the setpoint layers (e.g., Point-1, Point-2, etc.) in the table to the
right.
This layer contains 8 named ranges for the data columns, as follows:
| Range Name |  |  | Column |  | Range Specification
| Bracketing_K |  |  | F |  | =OFFSET('Bracketing Results'!$F$2,0,0,20)
| Bracketing_Ratio |  |  | G |  | =OFFSET('Bracketing Results'!$G$2,0,0,20)
| Bracketing_S2 |  |  | D |  | =OFFSET('Bracketing Results'!$D$2,0,0,20)
| Bracketing_Uncertainty |  |  | E |  | =OFFSET('Bracketing Results'!$E$2,0,0,20)
| Ref_Cert_Uncertainty |  |  | C |  | =OFFSET('Bracketing Results'!$C$2,0,0,20)
| Reference_Cert |  |  | B |  | =OFFSET('Bracketing Results'!$B$2,0,0,20)
| Reference_ID |  |  | K |  | =OFFSET('Bracketing Results'!$K$2,0,0,20)
| Setpoint |  |  | A |  | =OFFSET('Bracketing Results'!$A$2,0,0,20)
9. Uncertainty
The uncertainty calculations are described in detail in the traceability protocol. This layer receives all
of its data input from the bracketing results summary table. There is only one unprotected cell in this
layer: the "Candidate Location" cell. The output of the this layer is the rightmost column: the combined
uncertainties of the candidate gas concentrations.
10. Certification
This layer summarizes the certification, including the reference generator data, certified concentrations
for the candidate, and overall expanded uncertainties.

## Test_Info
|  |  | Operator: |  | Chris Winterrowd
|  |  | Company/Lab: |  | EPA/RTP Laboratories
|  |  | Candidate ID: |  | Tekran 6
|  |  | Candidate S/N: |  | 3009
|  |  | Reference ID: |  | Tekran 3
|  |  | Reference S/N: |  | 3039
|  |  | Start Date/Time: |  | 2008-08-04 15:25
|  |  | End Date/Time: |  | 2008-08-05 18:10
|  |  | Barometric Pressure: |  | 29.6719 | in. Hg
"A-Channel" Linearity |  |  | Reference Generator Certification |  |  |  |  | "B-Channel" Linearity
Injected ug/scm | Response | Linearity Uncertainty Factor | Certification Values |  | Certification Uncertainties |  |  | Injected ug/scm | Response | Linearity Uncertainty Factor
|  |  | Setpoint | Certified | 1s | 2s | Relative
0 | 8.2208 |  | ug/scm | ug/scm | ug/scm | ug/scm | 2s, % | 0 | 8.6717
0.53524 | 590.5074 | Reserved | 0.5 | 0.53524 | 0.0061034 | 0.012207 | 2.281% | 0.53524 | 581.7008 | Reserved
1.6557 | 1815.1429 | Reserved | 1.5 | 1.6557 | 0.033035 | 0.06607 | 3.99% | 1.6557 | 1796.1905 | Reserved
2.8581 | 3174 | Reserved | 2.7 | 2.8581 | 0.032743 | 0.065486 | 2.291% | 2.8581 | 3144.5 | Reserved
6.0937 | 6712.4286 | Reserved | 5.7 | 6.0937 | 0.0297 | 0.059399 | 0.9748% | 6.0937 | 6654.5714 | Reserved
10.4051 | 10933.7143 | Reserved | 9.5 | 10.4051 | 0.062122 | 0.12424 | 1.194% | 10.4051 | 10836.381 | Reserved
20.4651 | 21164.1053 | Reserved | 19 | 20.4651 | 0.10021 | 0.20041 | 0.9793% | 20.4651 | 20978.0417 | Reserved
39.6771 | 40452.5 | Reserved | 38 | 39.6771 | 0.20566 | 0.41132 | 1.037% | 39.6771 | 40063.1905 | Reserved
Note on E15: Ed Brown: The "nominal" setpoint, used for identification, not for calculation

## Cand_Raw
DATETIME | CALIB_AUXOUTPUTSELECT | CALIB_CALIBAIRPREHEATER_PV | CALIB_CALIBAIRPREHEATER_SP | CALIB_CALIBFAULTLATCH | CALIB_CALIBHEATEDPLATE_PV | CALIB_CALIBHEATEDPLATE_SP | CALIB_CALPRESSURE | CALIB_CALPRESSURESENSORFAIL | CALIB_CALSOURCEACTUAL | CALIB_CALSOURCEADJUST | CALIB_CALSOURCECOMPARE | CALIB_CALSOURCECORRECTION | CALIB_CALSOURCEFACTOR | CALIB_CALSOURCEFLOW | CALIB_CALSOURCEREADBACKRAW | CALIB_CALSOURCESTABLE | CALIB_CASEPV | CALIB_CONCENTRATION | CALIB_DILUTECOMPARE | CALIB_DILUTECONCENTRATION | CALIB_DILUTEFLOW | CALIB_DILUTEFLOWACTUAL | CALIB_DILUTEFLOWCORRECTION | CALIB_DILUTEFLOWREADBACK | CALIB_DILUTIONGAS | CALIB_INJECTSELECT | CALIB_LED | CALIB_MERCURYSOURCESP | CALIB_MERCURYSOURCETEMP | CALIB_NISTHGTABLE | CALIB_PLCSTATUS | CALIB_REMOTECONTROL | CALIB_SOURCEACTIVATE | CALIB_SOURCEHEATERENABLE | CALIB_STATTEMP | CALIB_STATTEMPLOGIC | CALIB_STATUS | CALIB_STATUSER1 | CALIB_STATUSER2 | CALIB_STATUSER3 | CALIB_TOTALFLOW | CALIB_TOTALFLOWACTUAL | CALIB_TOTALFLOWALTERNATE | CALIB_TOTALFLOWX | CALIB_CALSOURCEREADBACK | CALIB_CONCEPA
2008-08-04 14:39 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.22 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 14:40 | 0 | 85 | 85 | 0 | 40.1 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.22 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46831 | 0
2008-08-04 14:41 | 0 | 84.96 | 85 | 0 | 39.92 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.19 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 14:42 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.2 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 14:43 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.15 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 14:44 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.12 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 14:45 | 0 | 85 | 85 | 0 | 39.81 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.11 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48001 | 0
2008-08-04 14:46 | 0 | 84.91 | 85 | 0 | 40 | 40 | 99.5 | 0 | 0 | 0 | 0 |  | 0.96782 | 0 | 0.50061 | 0 | 33.09 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 14:47 | 0 | 85.05 | 85 | 0 | 39.9 | 40 | 99.5 | 0 | 0 | 0 | 0 |  | 0.96782 | 0 | 0.50061 | 0 | 33.02 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48001 | 0
2008-08-04 14:48 | 0 | 85.03 | 85 | 0 | 39.93 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96638 | 0 | 0.50061 | 0 | 33.06 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46855 | 0
2008-08-04 14:49 | 0 | 85 | 85 | 0 | 40.04 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.24 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 14:50 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.37 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 14:51 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96736 | 0 | 0.50061 | 0 | 33.4 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48001 | 0
2008-08-04 14:52 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 99.5505 | 0 | 0 | 0 | 0 |  | 0.96831 | 0 | 0.50061 | 0 | 33.36 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 14:53 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96736 | 0 | 0.50061 | 0 | 33.31 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48001 | 0
2008-08-04 14:54 | 0 | 85 | 85 | 0 | 39.96 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.31 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 14:55 | 0 | 84.97 | 85 | 0 | 40 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.3 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48001 | 0
2008-08-04 14:56 | 0 | 85 | 85 | 0 | 39.92 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.24 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46902 | 0
2008-08-04 14:57 | 0 | 84.97 | 85 | 0 | 40 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.2 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46902 | 0
2008-08-04 14:58 | 0 | 85.03 | 85 | 0 | 40 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.4884 | 0 | 33.16 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46902 | 0
2008-08-04 14:59 | 0 | 85.05 | 85 | 0 | 40.03 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.22 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:00 | 0 | 84.92 | 85 | 0 | 39.94 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.22 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 15:01 | 0 | 85.02 | 85 | 0 | 40.03 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.2 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46902 | 0
2008-08-04 15:02 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.23 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:03 | 0 | 84.99 | 85 | 0 | 40 | 40 | 99.5 | 0 | 0 | 0 | 0 |  | 0.96782 | 0 | 0.50061 | 0 | 33.24 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48024 | 0
2008-08-04 15:04 | 0 | 84.92 | 85 | 0 | 39.97 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.22 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48073 | 0
2008-08-04 15:05 | 0 | 85 | 85 | 0 | 39.91 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96736 | 0 | 0.50061 | 0 | 33.22 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48024 | 0
2008-08-04 15:06 | 0 | 85 | 85 | 0 | 40 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.22 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 15:07 | 0 | 84.92 | 85 | 0 | 40.04 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.22 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48073 | 0
2008-08-04 15:08 | 0 | 85.03 | 85 | 0 | 39.98 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.23 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 15:09 | 0 | 84.92 | 85 | 0 | 39.96 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96638 | 0 | 0.50061 | 0 | 33.24 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:10 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.4884 | 0 | 33.24 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46902 | 0
2008-08-04 15:11 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.35 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 15:12 | 0 | 85.05 | 85 | 0 | 39.96 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.43 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:13 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.43 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46855 | 0
2008-08-04 15:14 | 0 | 85.05 | 85 | 0 | 39.98 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.41 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:15 | 0 | 85 | 85 | 0 | 39.93 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.39 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:16 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.38 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46902 | 0
2008-08-04 15:17 | 0 | 84.99 | 85 | 0 | 40.07 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.4884 | 0 | 33.3 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:18 | 0 | 84.92 | 85 | 0 | 40.08 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.4884 | 0 | 33.31 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46855 | 0
2008-08-04 15:19 | 0 | 85 | 85 | 0 | 39.88 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.3 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:20 | 0 | 84.92 | 85 | 0 | 40 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.27 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 15:21 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.38 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:22 | 0 | 84.92 | 85 | 0 | 40 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.43 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48026 | 0
2008-08-04 15:23 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.47 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:24 | 0 | 84.92 | 85 | 0 | 39.92 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.45 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46902 | 0
2008-08-04 15:25 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.4884 | 0 | 33.41 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:26 | 0 | 84.96 | 85 | 0 | 40.08 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.4884 | 0 | 33.37 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46902 | 0
2008-08-04 15:27 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.4884 | 0 | 33.38 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:28 | 0 | 85 | 85 | 0 | 39.93 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.4884 | 0 | 33.33 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:29 | 0 | 85.03 | 85 | 0 | 39.91 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.3 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:30 | 0 | 85.02 | 85 | 0 | 39.93 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.3 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:31 | 0 | 84.99 | 85 | 0 | 40 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.50061 | 0 | 33.31 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46855 | 0
2008-08-04 15:32 | 0 | 84.92 | 85 | 0 | 39.95 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.33 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:33 | 0 | 85 | 85 | 0 | 39.96 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.4884 | 0 | 33.36 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:34 | 0 | 84.94 | 85 | 0 | 40.03 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.4884 | 0 | 33.41 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46902 | 0
2008-08-04 15:35 | 0 | 84.97 | 85 | 0 | 40 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96736 | 0 | 0.50061 | 0 | 33.42 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:36 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96638 | 0 | 0.50061 | 0 | 33.38 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:37 | 0 | 85.02 | 85 | 0 | 40.07 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96638 | 0 | 0.50061 | 0 | 33.34 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46902 | 0
2008-08-04 15:38 | 0 | 84.96 | 85 | 0 | 40 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.4884 | 0 | 33.31 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46902 | 0
2008-08-04 15:39 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96736 | 0 | 0.4884 | 0 | 33.26 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46902 | 0
2008-08-04 15:40 | 0 | 84.99 | 85 | 0 | 40 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.50061 | 0 | 33.23 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48073 | 0
2008-08-04 15:41 | 0 | 84.96 | 85 | 0 | 39.92 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96736 | 0 | 0.50061 | 0 | 33.19 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48024 | 0
2008-08-04 15:42 | 0 | 85 | 85 | 0 | 40.1 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96638 | 0 | 0.4884 | 0 | 33.16 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46853 | 0
2008-08-04 15:43 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96733 | 0 | 0.4884 | 0 | 33.11 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:44 | 0 | 84.94 | 85 | 0 | 39.97 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96638 | 0 | 0.4884 | 0 | 33.1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.469 | 0
2008-08-04 15:45 | 0 | 84.91 | 85 | 0 | 40 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96638 | 0 | 0.50061 | 0 | 33.16 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48024 | 0
2008-08-04 15:46 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96638 | 0 | 0.50061 | 0 | 33.23 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48073 | 0
2008-08-04 15:47 | 0 | 84.96 | 85 | 0 | 40 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96638 | 0 | 0.4884 | 0 | 33.29 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48073 | 0
2008-08-04 15:48 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96638 | 0 | 0.50061 | 0 | 33.31 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.469 | 0
2008-08-04 15:49 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 99.4495 | 0 | 0 | 0 | 0 |  | 0.96736 | 0 | 0.50061 | 0 | 33.27 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48075 | 0
2008-08-04 15:50 | 0 | 85 | 85 | 0 | 39.92 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96634 | 0 | 0.4884 | 0 | 33.26 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46902 | 0
2008-08-04 15:51 | 0 | 84.94 | 85 | 0 | 39.96 | 40 | 99.197 | 0 | 4.2229 | 4.0428 | 1 |  | 0.96487 | 4.19 | 0.53724 | 0 | 33.19 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.1795 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 0.48026 | 0
2008-08-04 15:52 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 99.096 | 0 | 4.2211 | 4.0387 | 0 |  | 0.96389 | 4.19 | 4.2369 | 0 | 33.22 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2047 | 0
2008-08-04 15:53 | 0 | 84.96 | 85 | 0 | 40.07 | 40 | 99.096 | 0 | 4.2171 | 4.045 | 0 |  | 0.96392 | 4.19 | 4.2369 | 0 | 33.23 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2178 | 0
2008-08-04 15:54 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 99.197 | 0 | 4.221 | 4.0387 | 0 |  | 0.96487 | 4.19 | 4.2247 | 0 | 33.22 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.1982 | 0
2008-08-04 15:55 | 0 | 84.89 | 85 | 0 | 39.97 | 40 | 99.096 | 0 | 4.217 | 4.0387 | 0 |  | 0.96389 | 4.19 | 4.2369 | 0 | 33.22 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2242 | 0
2008-08-04 15:56 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 99.096 | 0 | 4.2171 | 4.0387 | 0 |  | 0.96389 | 4.19 | 4.2247 | 0 | 33.29 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2112 | 0
2008-08-04 15:57 | 0 | 84.89 | 85 | 0 | 39.98 | 40 | 99.096 | 0 | 4.2171 | 4.0388 | 0 |  | 0.96392 | 4.19 | 4.2369 | 0 | 33.39 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2113 | 0
2008-08-04 15:58 | 0 | 84.91 | 85 | 0 | 39.91 | 40 | 99.197 | 0 | 4.217 | 4.0387 | 0 |  | 0.9649 | 4.19 | 4.2491 | 0 | 33.53 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2244 | 0
2008-08-04 15:59 | 0 | 84.88 | 85 | 0 | 39.92 | 40 | 99.197 | 0 | 4.221 | 4.0387 | 0 |  | 0.96487 | 4.19 | 4.2369 | 0 | 33.56 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.207 | 0
2008-08-04 16:00 | 0 | 84.89 | 85 | 0 | 39.86 | 40 | 99.0454 | 0 | 4.215 | 4.0387 | 0 |  | 0.9634 | 4.19 | 4.2369 | 0 | 33.55 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2113 | 0
2008-08-04 16:01 | 0 | 84.94 | 85 | 0 | 40.01 | 40 | 99.0454 | 0 | 4.221 | 4.0428 | 0 |  | 0.9634 | 4.19 | 4.2369 | 0 | 33.5 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2113 | 0
2008-08-04 16:02 | 0 | 84.89 | 85 | 0 | 40.01 | 40 | 99.0454 | 0 | 4.215 | 4.0387 | 0 |  | 0.9634 | 4.19 | 4.2369 | 0 | 33.5 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2113 | 0
2008-08-04 16:03 | 0 | 84.88 | 85 | 0 | 40.03 | 40 | 99.0454 | 0 | 4.217 | 4.0387 | 0 |  | 0.9634 | 4.19 | 4.2369 | 0 | 33.47 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2113 | 0
2008-08-04 16:04 | 0 | 84.84 | 85 | 0 | 39.94 | 40 | 99.0454 | 0 | 4.217 | 4.0387 | 0 |  | 0.9634 | 4.19 | 4.2369 | 0 | 33.43 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2135 | 0
2008-08-04 16:05 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 99.096 | 0 | 4.215 | 4.0387 | 0 |  | 0.96389 | 4.19 | 4.2491 | 0 | 33.43 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.207 | 0
2008-08-04 16:06 | 0 | 84.91 | 85 | 0 | 39.84 | 40 | 106.3189 | 0 | 4.4566 | 4.3331 | 0 |  | 1.0341 | 4.19 | 4.2369 | 0 | 33.43 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 3.9252 | 0.50071
2008-08-04 16:07 | 0 | 84.97 | 85 | 0 | 39.89 | 40 | 106.3695 | 0 | 4.5045 | 4.3331 | 0 |  | 1.0346 | 4.19 | 4.5299 | 0 | 33.49 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2032 | 0.50071
2008-08-04 16:08 | 0 | 84.89 | 85 | 0 | 39.94 | 40 | 106.3189 | 0 | 4.5065 | 4.3351 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.55 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2153 | 0.50071
2008-08-04 16:09 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 106.3189 | 0 | 4.5045 | 4.3351 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.62 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2012 | 0.50071
2008-08-04 16:10 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 106.3695 | 0 | 4.5045 | 4.3351 | 0 |  | 1.0346 | 4.19 | 4.5299 | 0 | 33.6 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2153 | 0.50071
2008-08-04 16:11 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 106.3189 | 0 | 4.5025 | 4.3331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.6 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2153 | 0.50071
2008-08-04 16:12 | 0 | 84.88 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 4.5025 | 4.3351 | 0 |  | 1.0341 | 4.19 | 4.5177 | 0 | 33.64 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2032 | 0.50071
2008-08-04 16:13 | 0 | 84.96 | 85 | 0 | 40.15 | 40 | 106.3189 | 0 | 4.5045 | 4.3351 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.66 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2153 | 0.50071
2008-08-04 16:14 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 106.3189 | 0 | 4.5065 | 4.3372 | 0 |  | 1.0341 | 4.19 | 4.5177 | 0 | 33.64 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2012 | 0.50071
2008-08-04 16:15 | 0 | 84.97 | 85 | 0 | 40 | 40 | 106.3695 | 0 | 4.5045 | 4.3331 | 0 |  | 1.0346 | 4.19 | 4.5177 | 0 | 33.64 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2175 | 0.50071
2008-08-04 16:16 | 0 | 84.91 | 85 | 0 | 40.05 | 40 | 106.3189 | 0 | 4.5023 | 4.3329 | 0 |  | 1.0341 | 4.19 | 4.5177 | 0 | 33.63 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 16:17 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 106.3189 | 0 | 4.5045 | 4.3351 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.62 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 16:18 | 0 | 84.95 | 85 | 0 | 40.05 | 40 | 106.3189 | 0 | 4.5025 | 4.3351 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.64 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2153 | 0.50071
2008-08-04 16:19 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 106.3695 | 0 | 4.5025 | 4.3331 | 0 |  | 1.0346 | 4.19 | 4.5299 | 0 | 33.66 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 16:20 | 0 | 84.99 | 85 | 0 | 40.05 | 40 | 106.3695 | 0 | 4.5025 | 4.3351 | 0 |  | 1.0346 | 4.19 | 4.5177 | 0 | 33.73 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2175 | 0.50071
2008-08-04 16:21 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 106.2684 | 0 | 4.5003 | 4.3329 | 0 |  | 1.0336 | 4.19 | 4.5299 | 0 | 33.8 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2033 | 0.50071
2008-08-04 16:22 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 106.2684 | 0 | 4.5005 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.86 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 16:23 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 106.2684 | 0 | 4.5003 | 4.3308 | 0 |  | 1.0336 | 4.19 | 4.5299 | 0 | 33.88 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2052 | 0.50071
2008-08-04 16:24 | 0 | 85 | 85 | 0 | 39.9 | 40 | 106.2684 | 0 | 4.5023 | 4.3308 | 0 |  | 1.0336 | 4.19 | 4.5299 | 0 | 33.82 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2052 | 0.50071
2008-08-04 16:25 | 0 | 85.05 | 85 | 0 | 39.96 | 40 | 106.2684 | 0 | 4.5025 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.81 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2195 | 0.50071
2008-08-04 16:26 | 0 | 84.96 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 4.5023 | 4.3269 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.82 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2234 | 0.50071
2008-08-04 16:27 | 0 | 84.96 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.5005 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.82 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 16:28 | 0 | 84.94 | 85 | 0 | 40.07 | 40 | 106.2684 | 0 | 4.5005 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.76 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 16:29 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 106.2684 | 0 | 4.5005 | 4.3331 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.76 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 16:30 | 0 | 84.94 | 85 | 0 | 39.93 | 40 | 106.2684 | 0 | 4.5025 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.73 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 16:31 | 0 | 85 | 85 | 0 | 40.04 | 40 | 106.2684 | 0 | 4.5025 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.68 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 16:32 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 106.2684 | 0 | 4.5005 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.7 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 16:33 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 106.2684 | 0 | 4.5005 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.76 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 16:34 | 0 | 85 | 85 | 0 | 40.08 | 40 | 106.2684 | 0 | 4.5025 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.88 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 16:35 | 0 | 85 | 85 | 0 | 40.03 | 40 | 106.2684 | 0 | 4.5045 | 4.3331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.03 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2072 | 0.50071
2008-08-04 16:36 | 0 | 84.92 | 85 | 0 | 39.94 | 40 | 106.3189 | 0 | 4.5025 | 4.3331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 34.1 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 16:37 | 0 | 85 | 85 | 0 | 39.97 | 40 | 106.3189 | 0 | 4.5025 | 4.3331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 34.1 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2032 | 0.50071
2008-08-04 16:38 | 0 | 84.97 | 85 | 0 | 39.86 | 40 | 106.2684 | 0 | 4.5005 | 4.3351 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.04 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2153 | 0.50071
2008-08-04 16:39 | 0 | 84.92 | 85 | 0 | 39.91 | 40 | 106.3189 | 0 | 4.5025 | 4.331 | 0 |  | 1.0341 | 4.19 | 4.5177 | 0 | 34.01 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 16:40 | 0 | 84.95 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.5005 | 4.3331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.97 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2052 | 0.50071
2008-08-04 16:41 | 0 | 85 | 85 | 0 | 39.91 | 40 | 106.3189 | 0 | 4.5025 | 4.3351 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.95 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 16:42 | 0 | 85.02 | 85 | 0 | 40.03 | 40 | 106.3189 | 0 | 4.5025 | 4.331 | 0 |  | 1.0341 | 4.19 | 4.5177 | 0 | 33.93 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2072 | 0.50071
2008-08-04 16:43 | 0 | 85.03 | 85 | 0 | 39.94 | 40 | 106.3189 | 0 | 4.5045 | 4.3331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.89 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2153 | 0.50071
2008-08-04 16:44 | 0 | 85 | 85 | 0 | 39.89 | 40 | 106.3695 | 0 | 4.5025 | 4.3331 | 0 |  | 1.0346 | 4.19 | 4.5299 | 0 | 33.88 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 16:45 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 106.3695 | 0 | 4.5025 | 4.3331 | 0 |  | 1.0346 | 4.19 | 4.5299 | 0 | 33.88 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 16:46 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 106.3189 | 0 | 4.5025 | 4.3331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.8 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 16:47 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 106.3189 | 0 | 4.5005 | 4.3331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.8 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 16:48 | 0 | 84.97 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 4.5005 | 4.331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.77 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 16:49 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 106.3189 | 0 | 4.5005 | 4.3351 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.77 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 16:50 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 106.3695 | 0 | 4.5005 | 4.3331 | 0 |  | 1.0347 | 4.19 | 4.5299 | 0 | 33.74 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 16:51 | 0 | 85 | 85 | 0 | 40.01 | 40 | 106.2684 | 0 | 4.5025 | 4.327 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.75 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 16:52 | 0 | 84.96 | 85 | 0 | 39.9 | 40 | 106.2684 | 0 | 4.5005 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.93 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.207 | 0.50071
2008-08-04 16:53 | 0 | 85 | 85 | 0 | 39.94 | 40 | 106.1674 | 0 | 4.4965 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 34.09 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.207 | 0.50071
2008-08-04 16:54 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 106.1674 | 0 | 4.5025 | 4.331 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 34.2 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2072 | 0.50071
2008-08-04 16:55 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 106.1674 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 34.24 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2052 | 0.50071
2008-08-04 16:56 | 0 | 84.95 | 85 | 0 | 40.01 | 40 | 106.1674 | 0 | 4.5005 | 4.331 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 34.17 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 16:57 | 0 | 84.94 | 85 | 0 | 40.01 | 40 | 106.1674 | 0 | 4.5005 | 4.331 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 34.17 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.199 | 0.50071
2008-08-04 16:58 | 0 | 85.03 | 85 | 0 | 39.88 | 40 | 106.2684 | 0 | 4.5005 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 34.09 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.195 | 0.50071
2008-08-04 16:59 | 0 | 84.92 | 85 | 0 | 39.98 | 40 | 106.2684 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.06 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2381 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2072 | 0.50071
2008-08-04 17:00 | 0 | 85 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 34.05 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:01 | 0 | 85.02 | 85 | 0 | 39.94 | 40 | 106.1674 | 0 | 4.5025 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5055 | 0 | 34.07 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:02 | 0 | 84.94 | 85 | 0 | 39.92 | 40 | 106.1674 | 0 | 4.5005 | 4.3331 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 34.22 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 17:03 | 0 | 85.03 | 85 | 0 | 39.95 | 40 | 106.1674 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 34.31 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2072 | 0.50071
2008-08-04 17:04 | 0 | 84.92 | 85 | 0 | 40.03 | 40 | 106.2684 | 0 | 4.4965 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.34 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:05 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 106.3189 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0341 | 4.19 | 4.5055 | 0 | 34.3 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:06 | 0 | 84.94 | 85 | 0 | 39.98 | 40 | 106.2684 | 0 | 4.5005 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.28 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2072 | 0.50071
2008-08-04 17:07 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 106.2684 | 0 | 4.5005 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.24 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:08 | 0 | 84.96 | 85 | 0 | 40.22 | 40 | 106.3189 | 0 | 4.5025 | 4.3331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 34.21 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2072 | 0.50071
2008-08-04 17:09 | 0 | 84.94 | 85 | 0 | 40.05 | 40 | 106.3189 | 0 | 4.5005 | 4.331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 34.2 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2052 | 0.50071
2008-08-04 17:10 | 0 | 85 | 85 | 0 | 39.98 | 40 | 106.1674 | 0 | 4.5025 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5055 | 0 | 34.17 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.199 | 0.50071
2008-08-04 17:11 | 0 | 85 | 85 | 0 | 39.98 | 40 | 106.1674 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 34.1 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:12 | 0 | 84.96 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.5005 | 4.3331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.08 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:13 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 106.3189 | 0 | 4.5005 | 4.331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 34.09 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:14 | 0 | 85 | 85 | 0 | 39.96 | 40 | 106.3189 | 0 | 4.5005 | 4.331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 34.11 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:15 | 0 | 85 | 85 | 0 | 40.03 | 40 | 106.3189 | 0 | 4.5025 | 4.3331 | 0 |  | 1.0341 | 4.19 | 4.5177 | 0 | 34.22 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:16 | 0 | 85 | 85 | 0 | 39.9 | 40 | 106.3189 | 0 | 4.5005 | 4.3331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 34.37 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:17 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 106.3695 | 0 | 4.5005 | 4.3331 | 0 |  | 1.0346 | 4.19 | 4.5299 | 0 | 34.42 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2173 | 0.50071
2008-08-04 17:18 | 0 | 85.02 | 85 | 0 | 39.98 | 40 | 106.3695 | 0 | 4.5025 | 4.331 | 0 |  | 1.0346 | 4.19 | 4.5299 | 0 | 34.43 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:19 | 0 | 84.95 | 85 | 0 | 39.84 | 40 | 106.2684 | 0 | 4.5025 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.38 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:20 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 106.3189 | 0 | 4.5005 | 4.3331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 34.33 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2052 | 0.50071
2008-08-04 17:21 | 0 | 84.99 | 85 | 0 | 40 | 40 | 106.1674 | 0 | 4.5005 | 4.331 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 34.31 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2032 | 0.50071
2008-08-04 17:22 | 0 | 85 | 85 | 0 | 39.98 | 40 | 106.2684 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.28 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2234 | 0.50071
2008-08-04 17:23 | 0 | 84.91 | 85 | 0 | 39.97 | 40 | 106.2684 | 0 | 4.4965 | 4.3269 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 34.2 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.195 | 0.50071
2008-08-04 17:24 | 0 | 85 | 85 | 0 | 39.84 | 40 | 106.2684 | 0 | 4.4965 | 4.3269 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 34.17 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:25 | 0 | 85 | 85 | 0 | 39.95 | 40 | 106.1674 | 0 | 4.5005 | 4.331 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 34.14 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:26 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 106.1674 | 0 | 4.4965 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5055 | 0 | 34.11 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.199 | 0.50071
2008-08-04 17:27 | 0 | 85.02 | 85 | 0 | 39.98 | 40 | 106.1674 | 0 | 4.4965 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5055 | 0 | 34.07 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2232 | 0.50071
2008-08-04 17:28 | 0 | 84.96 | 85 | 0 | 40.07 | 40 | 106.1674 | 0 | 4.5006 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 34.05 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.1948 | 0.50071
2008-08-04 17:29 | 0 | 85.05 | 85 | 0 | 39.98 | 40 | 106.2684 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.08 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:30 | 0 | 84.92 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.4965 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 34.15 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2381 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.195 | 0.50071
2008-08-04 17:31 | 0 | 85 | 85 | 0 | 40 | 40 | 106.1674 | 0 | 4.4965 | 4.331 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 34.26 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.205 | 0.50071
2008-08-04 17:32 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 106.1674 | 0 | 4.4965 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5055 | 0 | 34.32 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:33 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 106.1674 | 0 | 4.4965 | 4.3312 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 34.32 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:34 | 0 | 85 | 85 | 0 | 39.83 | 40 | 106.2684 | 0 | 4.5025 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.28 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:35 | 0 | 84.97 | 85 | 0 | 40 | 40 | 106.1674 | 0 | 4.5005 | 4.331 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 34.25 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:36 | 0 | 85.05 | 85 | 0 | 39.88 | 40 | 106.1674 | 0 | 4.5025 | 4.331 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 34.21 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2234 | 0.50071
2008-08-04 17:37 | 0 | 84.92 | 85 | 0 | 39.87 | 40 | 106.2684 | 0 | 4.5025 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 34.17 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2234 | 0.50071
2008-08-04 17:38 | 0 | 85.03 | 85 | 0 | 40.01 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.13 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.207 | 0.50071
2008-08-04 17:39 | 0 | 84.92 | 85 | 0 | 39.87 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.08 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.1941 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:40 | 0 | 84.95 | 85 | 0 | 39.91 | 40 | 106.2684 | 0 | 4.5005 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.07 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:41 | 0 | 84.92 | 85 | 0 | 39.89 | 40 | 106.2684 | 0 | 4.5025 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.1 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:42 | 0 | 84.99 | 85 | 0 | 40.15 | 40 | 106.2684 | 0 | 4.5025 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.24 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:43 | 0 | 85 | 85 | 0 | 40.07 | 40 | 106.3189 | 0 | 4.5005 | 4.3331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 34.35 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.207 | 0.50071
2008-08-04 17:44 | 0 | 85.03 | 85 | 0 | 39.98 | 40 | 106.2684 | 0 | 4.5005 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.42 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2381 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.207 | 0.50071
2008-08-04 17:45 | 0 | 84.94 | 85 | 0 | 39.93 | 40 | 106.2684 | 0 | 4.5005 | 4.3332 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 34.42 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:46 | 0 | 84.92 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.4965 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.4 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:47 | 0 | 85 | 85 | 0 | 39.9 | 40 | 106.2684 | 0 | 4.5005 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.34 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:48 | 0 | 85.02 | 85 | 0 | 40.07 | 40 | 106.2684 | 0 | 4.4965 | 4.3269 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.31 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.1868 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 17:49 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 106.2684 | 0 | 4.5025 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.3 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2052 | 0.50071
2008-08-04 17:50 | 0 | 84.92 | 85 | 0 | 40.04 | 40 | 106.1674 | 0 | 4.5025 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 34.29 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2234 | 0.50071
2008-08-04 17:51 | 0 | 84.94 | 85 | 0 | 39.96 | 40 | 106.1674 | 0 | 4.4965 | 4.331 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 34.34 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.195 | 0.50071
2008-08-04 17:52 | 0 | 85 | 85 | 0 | 39.84 | 40 | 106.1674 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5055 | 0 | 34.47 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2072 | 0.50071
2008-08-04 17:53 | 0 | 85 | 85 | 0 | 40.01 | 40 | 106.2684 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.58 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:54 | 0 | 85 | 85 | 0 | 39.96 | 40 | 106.2684 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 34.6 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.199 | 0.50071
2008-08-04 17:55 | 0 | 84.96 | 85 | 0 | 40.01 | 40 | 106.1674 | 0 | 4.4965 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 34.57 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2527 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:56 | 0 | 85.02 | 85 | 0 | 40.05 | 40 | 106.1674 | 0 | 4.4945 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 34.55 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2381 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:57 | 0 | 84.92 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.4965 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 34.5 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 17:58 | 0 | 85 | 85 | 0 | 40.07 | 40 | 106.2684 | 0 | 4.4965 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 34.46 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.195 | 0.50071
2008-08-04 17:59 | 0 | 84.94 | 85 | 0 | 39.95 | 40 | 106.1674 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 34.43 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2072 | 0.50071
2008-08-04 18:00 | 0 | 85 | 85 | 0 | 39.94 | 40 | 106.1674 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 34.34 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 18:01 | 0 | 84.95 | 85 | 0 | 40.04 | 40 | 106.1674 | 0 | 4.4965 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 34.3 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 18:02 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 106.2684 | 0 | 4.5005 | 4.3269 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 34.3 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.199 | 0.50071
2008-08-04 18:03 | 0 | 84.96 | 85 | 0 | 39.89 | 40 | 106.1674 | 0 | 4.4965 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5055 | 0 | 34.27 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.1988 | 0.50071
2008-08-04 18:04 | 0 | 84.96 | 85 | 0 | 40.01 | 40 | 106.1674 | 0 | 4.4965 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5055 | 0 | 34.26 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.1988 | 0.50071
2008-08-04 18:05 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 106.1674 | 0 | 4.4966 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5055 | 0 | 34.24 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.211 | 0.50071
2008-08-04 18:06 | 0 | 85 | 85 | 0 | 39.89 | 40 | 106.2684 | 0 | 4.4966 | 4.325 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.21 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.213 | 0.50071
2008-08-04 18:07 | 0 | 84.96 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.21 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 18:08 | 0 | 84.91 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.5005 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.18 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:09 | 0 | 85.05 | 85 | 0 | 39.92 | 40 | 106.3189 | 0 | 4.5046 | 4.331 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 34.13 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.1948 | 0.50071
2008-08-04 18:10 | 0 | 85.02 | 85 | 0 | 40.07 | 40 | 106.3189 | 0 | 4.5006 | 4.3332 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 34.14 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2172 | 0.50071
2008-08-04 18:11 | 0 | 85 | 85 | 0 | 39.94 | 40 | 106.3189 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0342 | 4.19 | 4.5299 | 0 | 34.13 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:12 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 106.2684 | 0 | 4.5006 | 4.3332 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34.13 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2072 | 0.50071
2008-08-04 18:13 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.09 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:14 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 106.2684 | 0 | 4.5026 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 34.03 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.207 | 0.50071
2008-08-04 18:15 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.205 | 0.50071
2008-08-04 18:16 | 0 | 85 | 85 | 0 | 39.95 | 40 | 106.2684 | 0 | 4.5026 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 34 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.1868 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:17 | 0 | 84.91 | 85 | 0 | 40.03 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.96 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.207 | 0.50071
2008-08-04 18:18 | 0 | 84.92 | 85 | 0 | 39.93 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.96 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:19 | 0 | 85 | 85 | 0 | 39.94 | 40 | 106.3189 | 0 | 4.5046 | 4.3332 | 0 |  | 1.0342 | 4.19 | 4.5299 | 0 | 33.93 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.205 | 0.50071
2008-08-04 18:20 | 0 | 84.97 | 85 | 0 | 40.07 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.91 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:21 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 106.2684 | 0 | 4.5006 | 4.3332 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.88 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.207 | 0.50071
2008-08-04 18:22 | 0 | 85.07 | 85 | 0 | 40.05 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 33.91 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.1948 | 0.50071
2008-08-04 18:23 | 0 | 84.96 | 85 | 0 | 40.07 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 33.88 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.1988 | 0.50071
2008-08-04 18:24 | 0 | 84.92 | 85 | 0 | 39.95 | 40 | 106.1674 | 0 | 4.4966 | 4.3312 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 33.86 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.205 | 0.50071
2008-08-04 18:25 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 106.1674 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 33.88 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.1868 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.207 | 0.50071
2008-08-04 18:26 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.85 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:27 | 0 | 85 | 85 | 0 | 39.84 | 40 | 106.2684 | 0 | 4.5006 | 4.3332 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.8 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.1948 | 0.50071
2008-08-04 18:28 | 0 | 85 | 85 | 0 | 39.91 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.84 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.211 | 0.50071
2008-08-04 18:29 | 0 | 85.05 | 85 | 0 | 39.94 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.79 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2172 | 0.50071
2008-08-04 18:30 | 0 | 84.97 | 85 | 0 | 40.01 | 40 | 106.2684 | 0 | 4.5006 | 4.3332 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.75 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2172 | 0.50071
2008-08-04 18:31 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 106.3189 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0342 | 4.19 | 4.5177 | 0 | 33.76 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:32 | 0 | 84.94 | 85 | 0 | 40 | 40 | 106.1674 | 0 | 4.4966 | 4.327 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 33.75 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.205 | 0.50071
2008-08-04 18:33 | 0 | 85.03 | 85 | 0 | 40.01 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.75 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:34 | 0 | 85.05 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.4966 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.75 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.1868 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.207 | 0.50071
2008-08-04 18:35 | 0 | 85.03 | 85 | 0 | 39.95 | 40 | 106.2684 | 0 | 4.4966 | 4.327 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.71 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:36 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 106.1674 | 0 | 4.4966 | 4.327 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 33.7 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2232 | 0.50071
2008-08-04 18:37 | 0 | 85.02 | 85 | 0 | 39.92 | 40 | 106.3189 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0342 | 4.19 | 4.5177 | 0 | 33.7 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:38 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.68 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:39 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 106.2684 | 0 | 4.5006 | 4.3332 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.67 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2172 | 0.50071
2008-08-04 18:40 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.64 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.205 | 0.50071
2008-08-04 18:41 | 0 | 84.96 | 85 | 0 | 40.01 | 40 | 106.2684 | 0 | 4.5006 | 4.3332 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.63 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2454 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:42 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.62 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:43 | 0 | 85.05 | 85 | 0 | 39.79 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.63 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2232 | 0.50071
2008-08-04 18:44 | 0 | 84.92 | 85 | 0 | 40.01 | 40 | 106.3189 | 0 | 4.4966 | 4.3332 | 0 |  | 1.0342 | 4.19 | 4.5055 | 0 | 33.63 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:45 | 0 | 85.02 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.5026 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.6 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.205 | 0.50071
2008-08-04 18:46 | 0 | 85 | 85 | 0 | 40.03 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 33.55 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.1948 | 0.50071
2008-08-04 18:47 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 106.2684 | 0 | 4.4966 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.57 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:48 | 0 | 85 | 85 | 0 | 39.93 | 40 | 106.2684 | 0 | 4.5026 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.56 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:49 | 0 | 85.07 | 85 | 0 | 40.03 | 40 | 106.2684 | 0 | 4.4966 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.59 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2172 | 0.50071
2008-08-04 18:50 | 0 | 85 | 85 | 0 | 40.04 | 40 | 106.2684 | 0 | 4.5026 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.59 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:51 | 0 | 84.92 | 85 | 0 | 39.87 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.52 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:52 | 0 | 85.05 | 85 | 0 | 39.98 | 40 | 106.1674 | 0 | 4.4966 | 4.3312 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 33.52 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.205 | 0.50071
2008-08-04 18:53 | 0 | 84.97 | 85 | 0 | 39.89 | 40 | 106.1674 | 0 | 4.4966 | 4.327 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 33.53 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.1868 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 18:54 | 0 | 84.94 | 85 | 0 | 40.05 | 40 | 106.1674 | 0 | 4.4966 | 4.3312 | 0 |  | 1.0327 | 4.19 | 4.5055 | 0 | 33.53 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.1948 | 0.50071
2008-08-04 18:55 | 0 | 85.03 | 85 | 0 | 39.97 | 40 | 106.1674 | 0 | 4.5026 | 4.327 | 0 |  | 1.0327 | 4.19 | 4.5055 | 0 | 33.51 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2454 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.207 | 0.50071
2008-08-04 18:56 | 0 | 85.02 | 85 | 0 | 40.04 | 40 | 106.1674 | 0 | 4.4966 | 4.327 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 33.53 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2232 | 0.50071
2008-08-04 18:57 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 106.1674 | 0 | 4.5006 | 4.327 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 33.47 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2232 | 0.50071
2008-08-04 18:58 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 106.2684 | 0 | 4.5026 | 4.327 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.46 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.211 | 0.50071
2008-08-04 18:59 | 0 | 84.99 | 85 | 0 | 40 | 40 | 106.1674 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 33.46 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.211 | 0.50071
2008-08-04 19:00 | 0 | 85 | 85 | 0 | 39.94 | 40 | 106.1674 | 0 | 4.5006 | 4.327 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 33.49 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.211 | 0.50071
2008-08-04 19:01 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 106.1674 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 33.5 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 19:02 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 106.2684 | 0 | 4.5026 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.47 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.207 | 0.50071
2008-08-04 19:03 | 0 | 85 | 85 | 0 | 39.95 | 40 | 106.1674 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 33.44 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.211 | 0.50071
2008-08-04 19:04 | 0 | 85 | 85 | 0 | 39.93 | 40 | 106.1674 | 0 | 4.4965 | 4.3269 | 0 |  | 1.0327 | 4.19 | 4.5299 | 0 | 33.46 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.211 | 0.50071
2008-08-04 19:05 | 0 | 84.94 | 85 | 0 | 40.03 | 40 | 106.1674 | 0 | 4.4966 | 4.327 | 0 |  | 1.0327 | 4.19 | 4.5177 | 0 | 33.43 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2112 | 0.50071
2008-08-04 19:06 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 106.2684 | 0 | 4.5006 | 4.327 | 0 |  | 1.0337 | 4.19 | 4.5055 | 0 | 33.4 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.1988 | 0.50071
2008-08-04 19:07 | 0 | 84.99 | 85 | 0 | 39.9 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.43 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 19:08 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.41 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 19:09 | 0 | 84.99 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.4966 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.41 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.207 | 0.50071
2008-08-04 19:10 | 0 | 85.02 | 85 | 0 | 39.91 | 40 | 106.2684 | 0 | 4.5006 | 4.3332 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.4 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2172 | 0.50071
2008-08-04 19:11 | 0 | 84.94 | 85 | 0 | 39.89 | 40 | 106.3189 | 0 | 4.5006 | 4.3332 | 0 |  | 1.0341 | 4.19 | 4.5299 | 0 | 33.41 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 19:12 | 0 | 85.02 | 85 | 0 | 39.88 | 40 | 106.3189 | 0 | 4.5005 | 4.3312 | 0 |  | 1.0342 | 4.19 | 4.5177 | 0 | 33.41 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.205 | 0.50071
2008-08-04 19:13 | 0 | 84.92 | 85 | 0 | 39.96 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.42 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2172 | 0.50071
2008-08-04 19:14 | 0 | 85 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.5005 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.41 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2381 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 19:15 | 0 | 85 | 85 | 0 | 40.03 | 40 | 106.3189 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0342 | 4.19 | 4.5299 | 0 | 33.42 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 19:16 | 0 | 84.97 | 85 | 0 | 39.84 | 40 | 106.2684 | 0 | 4.5006 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.43 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2172 | 0.50071
2008-08-04 19:17 | 0 | 85 | 85 | 0 | 39.87 | 40 | 106.2684 | 0 | 4.4965 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.4 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 19:18 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 106.2684 | 0 | 4.5006 | 4.3312 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.42 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 19:19 | 0 | 85.02 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.5026 | 4.3332 | 0 |  | 1.0337 | 4.19 | 4.5299 | 0 | 33.4 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2194 | 0.50071
2008-08-04 19:20 | 0 | 85.03 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 4.5006 | 4.331 | 0 |  | 1.0337 | 4.19 | 4.5177 | 0 | 33.37 | 0.53725 | 0 | 0.022432 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.257 | -1 | 25.0042 | 4.2192 | 0.50071
2008-08-04 19:21 | 0 | 85 | 85 | 0 | 39.96 | 40 | 106.3695 | 0 | 13.3557 | 12.9955 | 1 |  | 1.0347 | 12.56 | 12.7473 | 0 | 33.37 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 4.207 | 1.5004
2008-08-04 19:22 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 106.3695 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.35 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 19:23 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 106.3695 | 0 | 13.3557 | 12.9893 | 0 |  | 1.0347 | 12.56 | 13.37 | 0 | 33.37 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 19:24 | 0 | 85 | 85 | 0 | 39.94 | 40 | 106.3695 | 0 | 13.3557 | 13.0016 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.39 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 19:25 | 0 | 84.95 | 85 | 0 | 40.04 | 40 | 106.3695 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.38 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 19:26 | 0 | 84.99 | 85 | 0 | 40 | 40 | 106.3695 | 0 | 13.3557 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.34 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 19:27 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 106.3189 | 0 | 13.3553 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.35 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.562 | 1.5004
2008-08-04 19:28 | 0 | 85 | 85 | 0 | 39.96 | 40 | 106.3189 | 0 | 13.3614 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.35 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 19:29 | 0 | 84.92 | 85 | 0 | 39.93 | 40 | 106.3695 | 0 | 13.3675 | 13.0012 | 0 |  | 1.0346 | 12.56 | 13.37 | 0 | 33.31 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2454 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 19:30 | 0 | 85 | 85 | 0 | 40.01 | 40 | 106.3189 | 0 | 13.3614 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.31 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 19:31 | 0 | 85 | 85 | 0 | 39.96 | 40 | 106.3189 | 0 | 13.3675 | 12.9888 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.34 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2454 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5739 | 1.5004
2008-08-04 19:32 | 0 | 85 | 85 | 0 | 39.9 | 40 | 106.3695 | 0 | 13.3675 | 12.9888 | 0 |  | 1.0346 | 12.56 | 13.3822 | 0 | 33.33 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 19:33 | 0 | 85.05 | 85 | 0 | 40 | 40 | 106.3695 | 0 | 13.368 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.31 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 19:34 | 0 | 84.99 | 85 | 0 | 39.89 | 40 | 106.3695 | 0 | 13.3619 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.31 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 19:35 | 0 | 85 | 85 | 0 | 40.03 | 40 | 106.3189 | 0 | 13.3614 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.3822 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 19:36 | 0 | 84.97 | 85 | 0 | 39.89 | 40 | 106.3189 | 0 | 13.3553 | 12.9955 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.27 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 19:37 | 0 | 84.92 | 85 | 0 | 39.94 | 40 | 106.3189 | 0 | 13.3557 | 12.9955 | 0 |  | 1.0342 | 12.56 | 13.37 | 0 | 33.3 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 19:38 | 0 | 85 | 85 | 0 | 39.95 | 40 | 106.3189 | 0 | 13.3619 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.29 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 19:39 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 106.3189 | 0 | 13.3619 | 12.9893 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.31 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 19:40 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.3 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 19:41 | 0 | 85 | 85 | 0 | 39.97 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.29 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.1795 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 19:42 | 0 | 84.97 | 85 | 0 | 39.84 | 40 | 106.3695 | 0 | 13.3614 | 12.9893 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.31 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.562 | 1.5004
2008-08-04 19:43 | 0 | 85 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 13.3553 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.3 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.562 | 1.5004
2008-08-04 19:44 | 0 | 85 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 13.3614 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.29 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.562 | 1.5004
2008-08-04 19:45 | 0 | 84.96 | 85 | 0 | 39.87 | 40 | 106.3189 | 0 | 13.3557 | 12.9893 | 0 |  | 1.0342 | 12.56 | 13.37 | 0 | 33.29 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.562 | 1.5004
2008-08-04 19:46 | 0 | 85 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 13.3619 | 12.9893 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.27 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.562 | 1.5004
2008-08-04 19:47 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.3 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 19:48 | 0 | 85.03 | 85 | 0 | 39.94 | 40 | 106.3189 | 0 | 13.3557 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.562 | 1.5004
2008-08-04 19:49 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 106.3189 | 0 | 13.3619 | 12.9893 | 0 |  | 1.0342 | 12.56 | 13.37 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.562 | 1.5004
2008-08-04 19:50 | 0 | 84.97 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 13.3619 | 12.9893 | 0 |  | 1.0342 | 12.56 | 13.3455 | 0 | 33.3 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.55 | 1.5004
2008-08-04 19:51 | 0 | 84.95 | 85 | 0 | 39.9 | 40 | 106.3695 | 0 | 13.3553 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 19:52 | 0 | 84.97 | 85 | 0 | 39.89 | 40 | 106.3695 | 0 | 13.3614 | 12.9888 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 19:53 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 106.3189 | 0 | 13.3614 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 19:54 | 0 | 85.02 | 85 | 0 | 40.01 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.24 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 19:55 | 0 | 85.02 | 85 | 0 | 40.01 | 40 | 106.42 | 0 | 13.368 | 12.9893 | 0 |  | 1.0352 | 12.56 | 13.37 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 19:56 | 0 | 84.97 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 13.3557 | 12.9955 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.24 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 19:57 | 0 | 84.96 | 85 | 0 | 40.07 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 19:58 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 106.3695 | 0 | 13.3557 | 12.9888 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.27 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.1868 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 19:59 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 106.3695 | 0 | 13.3619 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 20:00 | 0 | 84.94 | 85 | 0 | 40.05 | 40 | 106.3189 | 0 | 13.3614 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 20:01 | 0 | 84.92 | 85 | 0 | 40.03 | 40 | 106.3189 | 0 | 13.3614 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.27 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 20:02 | 0 | 84.96 | 85 | 0 | 40.13 | 40 | 106.3189 | 0 | 13.368 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.24 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 20:03 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 106.3189 | 0 | 13.3553 | 12.9893 | 0 |  | 1.0342 | 12.56 | 13.37 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 20:04 | 0 | 84.97 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 13.3553 | 13.0012 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 20:05 | 0 | 85 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 13.3553 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.24 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5739 | 1.5004
2008-08-04 20:06 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 106.3189 | 0 | 13.3675 | 12.9893 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.23 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.562 | 1.5004
2008-08-04 20:07 | 0 | 85.02 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 13.3614 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.27 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 20:08 | 0 | 85 | 85 | 0 | 40.03 | 40 | 106.3695 | 0 | 13.3492 | 12.9827 | 0 |  | 1.0346 | 12.56 | 13.3455 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 20:09 | 0 | 84.97 | 85 | 0 | 39.91 | 40 | 106.2684 | 0 | 13.3496 | 12.9831 | 0 |  | 1.0337 | 12.56 | 13.3578 | 0 | 33.23 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.562 | 1.5004
2008-08-04 20:10 | 0 | 84.94 | 85 | 0 | 39.84 | 40 | 106.3189 | 0 | 13.3557 | 12.9893 | 0 |  | 1.0342 | 12.56 | 13.37 | 0 | 33.27 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5739 | 1.5004
2008-08-04 20:11 | 0 | 85 | 85 | 0 | 39.96 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.27 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 20:12 | 0 | 84.94 | 85 | 0 | 40.04 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.562 | 1.5004
2008-08-04 20:13 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0342 | 12.56 | 13.37 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 20:14 | 0 | 84.97 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 13.3614 | 12.9827 | 0 |  | 1.0337 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 20:15 | 0 | 84.97 | 85 | 0 | 40.1 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 20:16 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 106.2684 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0337 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.55 | 1.5004
2008-08-04 20:17 | 0 | 84.94 | 85 | 0 | 39.9 | 40 | 106.2684 | 0 | 13.3557 | 12.9831 | 0 |  | 1.0337 | 12.56 | 13.3455 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 20:18 | 0 | 85.18 | 85 | 0 | 39.89 | 40 | 106.3189 | 0 | 13.3557 | 12.9831 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 20:19 | 0 | 84.91 | 85 | 0 | 40.03 | 40 | 106.3189 | 0 | 13.3492 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.26 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5565 | 1.5004
2008-08-04 20:20 | 0 | 84.97 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 13.3492 | 12.9827 | 0 |  | 1.0337 | 12.56 | 13.3578 | 0 | 33.2 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5803 | 1.5004
2008-08-04 20:21 | 0 | 85.03 | 85 | 0 | 39.9 | 40 | 106.3695 | 0 | 13.3492 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.3455 | 0 | 33.23 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 20:22 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 106.3695 | 0 | 13.3619 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.37 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 20:23 | 0 | 84.96 | 85 | 0 | 40.13 | 40 | 106.3695 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.24 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.562 | 1.5004
2008-08-04 20:24 | 0 | 84.92 | 85 | 0 | 39.89 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0342 | 12.56 | 13.37 | 0 | 33.24 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 20:25 | 0 | 85.03 | 85 | 0 | 39.89 | 40 | 106.3189 | 0 | 13.3557 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.2 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 20:26 | 0 | 84.94 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.19 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 20:27 | 0 | 85.05 | 85 | 0 | 39.93 | 40 | 106.3189 | 0 | 13.3553 | 12.9893 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 20:28 | 0 | 84.84 | 85 | 0 | 40.07 | 40 | 106.3189 | 0 | 13.3553 | 12.9955 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.24 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 20:29 | 0 | 84.94 | 85 | 0 | 39.92 | 40 | 106.3695 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0346 | 12.56 | 13.37 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 20:30 | 0 | 85.03 | 85 | 0 | 40.08 | 40 | 106.3189 | 0 | 13.3557 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.23 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 20:31 | 0 | 85.05 | 85 | 0 | 40.14 | 40 | 106.3695 | 0 | 13.3553 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.3822 | 0 | 33.24 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 20:32 | 0 | 85 | 85 | 0 | 40.07 | 40 | 106.3189 | 0 | 13.3619 | 12.9893 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 20:33 | 0 | 85.02 | 85 | 0 | 40.03 | 40 | 106.3189 | 0 | 13.3614 | 12.9831 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.24 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 20:34 | 0 | 84.99 | 85 | 0 | 40 | 40 | 106.3695 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.19 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 20:35 | 0 | 84.97 | 85 | 0 | 40.05 | 40 | 106.3189 | 0 | 13.3557 | 12.9955 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 20:36 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 106.3189 | 0 | 13.3614 | 12.9955 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 20:37 | 0 | 85.19 | 85 | 0 | 39.73 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0342 | 12.56 | 13.3455 | 0 | 33.23 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 20:38 | 0 | 85 | 85 | 0 | 39.95 | 40 | 106.3189 | 0 | 13.3492 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3455 | 0 | 33.2 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5505 | 1.5004
2008-08-04 20:39 | 0 | 84.91 | 85 | 0 | 39.97 | 40 | 106.3189 | 0 | 13.3557 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.2 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5505 | 1.5004
2008-08-04 20:40 | 0 | 84.94 | 85 | 0 | 39.91 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.2 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 20:41 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 106.3189 | 0 | 13.3496 | 12.9893 | 0 |  | 1.0342 | 12.56 | 13.37 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 20:42 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 106.2684 | 0 | 13.3614 | 12.995 | 0 |  | 1.0337 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 20:43 | 0 | 84.96 | 85 | 0 | 40.05 | 40 | 106.2684 | 0 | 13.3614 | 12.995 | 0 |  | 1.0337 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5441 | 1.5004
2008-08-04 20:44 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 106.3189 | 0 | 13.3557 | 12.9831 | 0 |  | 1.0342 | 12.56 | 13.3455 | 0 | 33.19 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 20:45 | 0 | 84.99 | 85 | 0 | 39.9 | 40 | 106.3189 | 0 | 13.3619 | 12.9893 | 0 |  | 1.0342 | 12.56 | 13.37 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 20:46 | 0 | 84.95 | 85 | 0 | 40 | 40 | 106.3695 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0346 | 12.56 | 13.37 | 0 | 33.24 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 20:47 | 0 | 84.92 | 85 | 0 | 39.94 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0342 | 12.56 | 13.37 | 0 | 33.2 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 20:48 | 0 | 84.89 | 85 | 0 | 39.86 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 20:49 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 20:50 | 0 | 85 | 85 | 0 | 40.05 | 40 | 106.3189 | 0 | 13.3557 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.2 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 20:51 | 0 | 85 | 85 | 0 | 39.98 | 40 | 106.3695 | 0 | 13.3557 | 12.9893 | 0 |  | 1.0347 | 12.56 | 13.3455 | 0 | 33.2 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 20:52 | 0 | 84.94 | 85 | 0 | 40 | 40 | 106.3695 | 0 | 13.3553 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.3822 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 20:53 | 0 | 85 | 85 | 0 | 39.96 | 40 | 106.3695 | 0 | 13.3557 | 13.0016 | 0 |  | 1.0347 | 12.56 | 13.37 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 20:54 | 0 | 84.96 | 85 | 0 | 40 | 40 | 106.42 | 0 | 13.368 | 12.995 | 0 |  | 1.0351 | 12.56 | 13.3578 | 0 | 33.2 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 20:55 | 0 | 84.96 | 85 | 0 | 40 | 40 | 106.3695 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.562 | 1.5004
2008-08-04 20:56 | 0 | 85 | 85 | 0 | 39.96 | 40 | 106.3189 | 0 | 13.3557 | 12.9955 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5505 | 1.5004
2008-08-04 20:57 | 0 | 84.97 | 85 | 0 | 39.89 | 40 | 106.3695 | 0 | 13.3614 | 12.9888 | 0 |  | 1.0346 | 12.56 | 13.37 | 0 | 33.2 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 20:58 | 0 | 85 | 85 | 0 | 40.03 | 40 | 106.42 | 0 | 13.3675 | 12.9955 | 0 |  | 1.0351 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 20:59 | 0 | 84.92 | 85 | 0 | 39.87 | 40 | 106.3695 | 0 | 13.3553 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.19 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.1868 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 21:00 | 0 | 85 | 85 | 0 | 40.08 | 40 | 106.3189 | 0 | 13.3614 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.19 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 21:01 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 106.3189 | 0 | 13.3619 | 12.9955 | 0 |  | 1.0342 | 12.56 | 13.37 | 0 | 33.15 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 21:02 | 0 | 84.97 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 13.3553 | 12.995 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.18 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 21:03 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 106.3695 | 0 | 13.3553 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.2 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5619 | 1.5004
2008-08-04 21:04 | 0 | 84.96 | 85 | 0 | 40 | 40 | 106.3695 | 0 | 13.3614 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.37 | 0 | 33.19 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 21:05 | 0 | 85 | 85 | 0 | 40.05 | 40 | 106.3189 | 0 | 13.3614 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.18 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 21:06 | 0 | 84.97 | 85 | 0 | 40.01 | 40 | 106.3695 | 0 | 13.3614 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.37 | 0 | 33.16 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 21:07 | 0 | 85 | 85 | 0 | 40 | 40 | 106.2684 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0337 | 12.56 | 13.3578 | 0 | 33.18 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 21:08 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 106.3189 | 0 | 13.3553 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.19 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2381 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 21:09 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 106.3189 | 0 | 13.3614 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.19 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 21:10 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.17 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 21:11 | 0 | 84.96 | 85 | 0 | 40.07 | 40 | 106.3695 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.2 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 21:12 | 0 | 85 | 85 | 0 | 40.05 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.19 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 21:13 | 0 | 85 | 85 | 0 | 39.93 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.16 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 21:14 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 106.3189 | 0 | 13.3553 | 12.9827 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.18 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 21:15 | 0 | 84.91 | 85 | 0 | 40.07 | 40 | 106.3695 | 0 | 13.3557 | 12.9893 | 0 |  | 1.0346 | 12.56 | 13.37 | 0 | 33.14 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 21:16 | 0 | 84.92 | 85 | 0 | 39.94 | 40 | 106.3189 | 0 | 13.3557 | 12.9955 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.16 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 21:17 | 0 | 85.02 | 85 | 0 | 39.94 | 40 | 106.3695 | 0 | 13.3557 | 12.9893 | 0 |  | 1.0347 | 12.56 | 13.37 | 0 | 33.17 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 21:18 | 0 | 85 | 85 | 0 | 39.95 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.17 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 21:19 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 106.3189 | 0 | 13.3553 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.18 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 21:20 | 0 | 84.87 | 85 | 0 | 39.88 | 40 | 106.2684 | 0 | 13.3557 | 12.9893 | 0 |  | 1.0337 | 12.56 | 13.37 | 0 | 33.17 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 21:21 | 0 | 84.91 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 13.3619 | 12.9955 | 0 |  | 1.0342 | 12.56 | 13.3578 | 0 | 33.17 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 21:22 | 0 | 85 | 85 | 0 | 39.88 | 40 | 106.3695 | 0 | 13.3675 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.19 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 21:23 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 106.3695 | 0 | 13.3614 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.37 | 0 | 33.19 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 21:24 | 0 | 85 | 85 | 0 | 39.91 | 40 | 106.3695 | 0 | 13.3619 | 13.0016 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.22 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5505 | 1.5004
2008-08-04 21:25 | 0 | 84.95 | 85 | 0 | 40.08 | 40 | 106.42 | 0 | 13.3675 | 13.0012 | 0 |  | 1.0351 | 12.56 | 13.3822 | 0 | 33.18 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 21:26 | 0 | 85 | 85 | 0 | 39.92 | 40 | 106.3695 | 0 | 13.3553 | 13.0012 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.17 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5505 | 1.5004
2008-08-04 21:27 | 0 | 84.92 | 85 | 0 | 40 | 40 | 106.3695 | 0 | 13.368 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.14 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 21:28 | 0 | 84.95 | 85 | 0 | 39.88 | 40 | 106.3695 | 0 | 13.368 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.37 | 0 | 33.15 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5505 | 1.5004
2008-08-04 21:29 | 0 | 84.95 | 85 | 0 | 40.05 | 40 | 106.42 | 0 | 13.3614 | 13.0012 | 0 |  | 1.0351 | 12.56 | 13.3578 | 0 | 33.17 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 21:30 | 0 | 85.03 | 85 | 0 | 39.95 | 40 | 106.3695 | 0 | 13.3614 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.37 | 0 | 33.2 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 21:31 | 0 | 85.02 | 85 | 0 | 39.86 | 40 | 106.3695 | 0 | 13.3619 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.37 | 0 | 33.16 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 21:32 | 0 | 84.95 | 85 | 0 | 40.04 | 40 | 106.3695 | 0 | 13.3619 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.37 | 0 | 33.18 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2381 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 21:33 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 106.3695 | 0 | 13.3675 | 12.995 | 0 |  | 1.0347 | 12.56 | 13.37 | 0 | 33.17 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 21:34 | 0 | 85 | 85 | 0 | 39.98 | 40 | 106.42 | 0 | 13.3675 | 13.0012 | 0 |  | 1.0351 | 12.56 | 13.37 | 0 | 33.16 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 21:35 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 106.42 | 0 | 13.3675 | 13.0012 | 0 |  | 1.0351 | 12.56 | 13.37 | 0 | 33.17 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 21:36 | 0 | 84.99 | 85 | 0 | 40.04 | 40 | 106.3695 | 0 | 13.3614 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.37 | 0 | 33.16 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 21:37 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 106.3695 | 0 | 13.3619 | 12.9893 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.16 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5739 | 1.5004
2008-08-04 21:38 | 0 | 84.99 | 85 | 0 | 40.07 | 40 | 106.3695 | 0 | 13.3553 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.37 | 0 | 33.17 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 21:39 | 0 | 85.07 | 85 | 0 | 40.07 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.17 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 21:40 | 0 | 84.92 | 85 | 0 | 40 | 40 | 106.3189 | 0 | 13.3614 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.37 | 0 | 33.19 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 21:41 | 0 | 84.92 | 85 | 0 | 39.96 | 40 | 106.3695 | 0 | 13.368 | 12.9955 | 0 |  | 1.0346 | 12.56 | 13.3822 | 0 | 33.16 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5619 | 1.5004
2008-08-04 21:42 | 0 | 85 | 85 | 0 | 40.11 | 40 | 106.3695 | 0 | 13.3553 | 13.0012 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.18 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2454 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5505 | 1.5004
2008-08-04 21:43 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 106.3695 | 0 | 13.3619 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.14 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 21:44 | 0 | 85 | 85 | 0 | 40.04 | 40 | 106.3695 | 0 | 13.3553 | 13.0016 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.16 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 21:45 | 0 | 85 | 85 | 0 | 39.92 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.13 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 21:46 | 0 | 84.99 | 85 | 0 | 39.89 | 40 | 106.3695 | 0 | 13.3557 | 12.9893 | 0 |  | 1.0347 | 12.56 | 13.37 | 0 | 33.13 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5739 | 1.5004
2008-08-04 21:47 | 0 | 84.94 | 85 | 0 | 40.04 | 40 | 106.3695 | 0 | 13.368 | 13.0016 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.13 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5505 | 1.5004
2008-08-04 21:48 | 0 | 84.92 | 85 | 0 | 39.97 | 40 | 106.3189 | 0 | 13.3675 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.12 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 21:49 | 0 | 85 | 85 | 0 | 39.87 | 40 | 106.3189 | 0 | 13.3553 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.14 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 21:50 | 0 | 85.03 | 85 | 0 | 39.98 | 40 | 106.42 | 0 | 13.3614 | 12.995 | 0 |  | 1.0351 | 12.56 | 13.3578 | 0 | 33.1 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 21:51 | 0 | 85.03 | 85 | 0 | 39.88 | 40 | 106.3189 | 0 | 13.3614 | 12.9888 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.12 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 21:52 | 0 | 84.97 | 85 | 0 | 39.84 | 40 | 106.3695 | 0 | 13.3614 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.37 | 0 | 33.14 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 21:53 | 0 | 84.95 | 85 | 0 | 39.91 | 40 | 106.521 | 0 | 13.368 | 13.0016 | 0 |  | 1.0361 | 12.56 | 13.37 | 0 | 33.17 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5619 | 1.5004
2008-08-04 21:54 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 106.3695 | 0 | 13.3619 | 13.0016 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.15 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.55 | 1.5004
2008-08-04 21:55 | 0 | 84.89 | 85 | 0 | 39.98 | 40 | 106.3695 | 0 | 13.3675 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.15 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 21:56 | 0 | 85 | 85 | 0 | 39.86 | 40 | 106.3695 | 0 | 13.3675 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.16 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5505 | 1.5004
2008-08-04 21:57 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 106.42 | 0 | 13.3802 | 13.0016 | 0 |  | 1.0352 | 12.56 | 13.37 | 0 | 33.15 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5803 | 1.5004
2008-08-04 21:58 | 0 | 85 | 85 | 0 | 39.86 | 40 | 106.42 | 0 | 13.3675 | 13.0016 | 0 |  | 1.0352 | 12.56 | 13.3822 | 0 | 33.16 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5739 | 1.5004
2008-08-04 21:59 | 0 | 84.94 | 85 | 0 | 39.9 | 40 | 106.42 | 0 | 13.3614 | 12.995 | 0 |  | 1.0352 | 12.56 | 13.37 | 0 | 33.16 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 22:00 | 0 | 84.92 | 85 | 0 | 40.03 | 40 | 106.3695 | 0 | 13.3614 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.15 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.55 | 1.5004
2008-08-04 22:01 | 0 | 85.05 | 85 | 0 | 40.03 | 40 | 106.521 | 0 | 13.3675 | 13.0012 | 0 |  | 1.0361 | 12.56 | 13.37 | 0 | 33.14 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 22:02 | 0 | 84.99 | 85 | 0 | 40.13 | 40 | 106.42 | 0 | 13.368 | 13.0016 | 0 |  | 1.0352 | 12.56 | 13.37 | 0 | 33.15 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5619 | 1.5004
2008-08-04 22:03 | 0 | 84.97 | 85 | 0 | 40.07 | 40 | 106.3695 | 0 | 13.3675 | 13.0012 | 0 |  | 1.0346 | 12.56 | 13.37 | 0 | 33.13 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5619 | 1.5004
2008-08-04 22:04 | 0 | 84.89 | 85 | 0 | 39.98 | 40 | 106.42 | 0 | 13.3614 | 12.995 | 0 |  | 1.0351 | 12.56 | 13.3822 | 0 | 33.16 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.55 | 1.5004
2008-08-04 22:05 | 0 | 85.03 | 85 | 0 | 39.9 | 40 | 106.521 | 0 | 13.3675 | 13.0012 | 0 |  | 1.0361 | 12.56 | 13.3822 | 0 | 33.12 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 22:06 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 106.42 | 0 | 13.3675 | 13.0012 | 0 |  | 1.0351 | 12.56 | 13.3822 | 0 | 33.13 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 22:07 | 0 | 85 | 85 | 0 | 40.01 | 40 | 106.3695 | 0 | 13.368 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.14 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 22:08 | 0 | 84.91 | 85 | 0 | 40.15 | 40 | 106.3695 | 0 | 13.3614 | 13.0012 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.13 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 22:09 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 106.3695 | 0 | 13.3614 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.13 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 22:10 | 0 | 85.07 | 85 | 0 | 39.93 | 40 | 106.42 | 0 | 13.3675 | 13.0012 | 0 |  | 1.0351 | 12.56 | 13.3578 | 0 | 33.14 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5505 | 1.5004
2008-08-04 22:11 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 106.3695 | 0 | 13.3614 | 13.0012 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.14 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5619 | 1.5004
2008-08-04 22:12 | 0 | 84.87 | 85 | 0 | 39.84 | 40 | 106.3695 | 0 | 13.3675 | 12.995 | 0 |  | 1.0347 | 12.56 | 13.37 | 0 | 33.13 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5505 | 1.5004
2008-08-04 22:13 | 0 | 85 | 85 | 0 | 40.19 | 40 | 106.3695 | 0 | 13.3614 | 13.0012 | 0 |  | 1.0346 | 12.56 | 13.37 | 0 | 33.1 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 22:14 | 0 | 84.95 | 85 | 0 | 40 | 40 | 106.42 | 0 | 13.3614 | 13.0012 | 0 |  | 1.0351 | 12.56 | 13.37 | 0 | 33.11 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 22:15 | 0 | 84.91 | 85 | 0 | 39.95 | 40 | 106.3695 | 0 | 13.368 | 12.9955 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.13 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 22:16 | 0 | 85 | 85 | 0 | 39.83 | 40 | 106.42 | 0 | 13.3619 | 13.0016 | 0 |  | 1.0352 | 12.56 | 13.37 | 0 | 33.12 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 22:17 | 0 | 84.95 | 85 | 0 | 40 | 40 | 106.3695 | 0 | 13.3619 | 12.9955 | 0 |  | 1.0347 | 12.56 | 13.37 | 0 | 33.13 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 22:18 | 0 | 84.95 | 85 | 0 | 39.92 | 40 | 106.3695 | 0 | 13.3614 | 12.995 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.11 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2381 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 22:19 | 0 | 84.92 | 85 | 0 | 40.05 | 40 | 106.3189 | 0 | 13.3614 | 12.995 | 0 |  | 1.0341 | 12.56 | 13.3578 | 0 | 33.1 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 22:20 | 0 | 85.03 | 85 | 0 | 40.13 | 40 | 106.42 | 0 | 13.3614 | 12.995 | 0 |  | 1.0351 | 12.56 | 13.37 | 0 | 33.11 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 22:21 | 0 | 85.05 | 85 | 0 | 39.91 | 40 | 106.3695 | 0 | 13.3619 | 13.0016 | 0 |  | 1.0347 | 12.56 | 13.3578 | 0 | 33.08 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.1868 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.556 | 1.5004
2008-08-04 22:22 | 0 | 85.02 | 85 | 0 | 39.91 | 40 | 106.42 | 0 | 13.368 | 13.0135 | 0 |  | 1.0351 | 12.56 | 13.3578 | 0 | 33.09 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5386 | 1.5004
2008-08-04 22:23 | 0 | 85.05 | 85 | 0 | 40.1 | 40 | 106.42 | 0 | 13.3614 | 13.0012 | 0 |  | 1.0351 | 12.56 | 13.3822 | 0 | 33.13 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.1868 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5858 | 1.5004
2008-08-04 22:24 | 0 | 85 | 85 | 0 | 39.95 | 40 | 106.42 | 0 | 13.368 | 12.995 | 0 |  | 1.0351 | 12.56 | 13.37 | 0 | 33.09 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5505 | 1.5004
2008-08-04 22:25 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 106.42 | 0 | 13.368 | 12.9955 | 0 |  | 1.0352 | 12.56 | 13.3822 | 0 | 33.09 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2381 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5739 | 1.5004
2008-08-04 22:26 | 0 | 85 | 85 | 0 | 39.93 | 40 | 106.42 | 0 | 13.3675 | 13.0012 | 0 |  | 1.0351 | 12.56 | 13.3822 | 0 | 33.1 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 22:27 | 0 | 85 | 85 | 0 | 40.11 | 40 | 106.42 | 0 | 13.368 | 13.0016 | 0 |  | 1.0351 | 12.56 | 13.3578 | 0 | 33.08 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 22:28 | 0 | 84.99 | 85 | 0 | 40.14 | 40 | 106.42 | 0 | 13.3675 | 12.9955 | 0 |  | 1.0352 | 12.56 | 13.37 | 0 | 33.12 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5684 | 1.5004
2008-08-04 22:29 | 0 | 85.03 | 85 | 0 | 39.98 | 40 | 106.42 | 0 | 13.3614 | 13.0012 | 0 |  | 1.0351 | 12.56 | 13.3822 | 0 | 33.1 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5624 | 1.5004
2008-08-04 22:30 | 0 | 84.94 | 85 | 0 | 39.93 | 40 | 106.42 | 0 | 13.3675 | 12.995 | 0 |  | 1.0351 | 12.56 | 13.3822 | 0 | 33.09 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5803 | 1.5004
2008-08-04 22:31 | 0 | 84.96 | 85 | 0 | 39.9 | 40 | 106.42 | 0 | 13.3614 | 12.995 | 0 |  | 1.0351 | 12.56 | 13.3822 | 0 | 33.1 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5743 | 1.5004
2008-08-04 22:32 | 0 | 84.92 | 85 | 0 | 39.98 | 40 | 106.42 | 0 | 13.3614 | 13.0012 | 0 |  | 1.0351 | 12.56 | 13.37 | 0 | 33.1 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5619 | 1.5004
2008-08-04 22:33 | 0 | 85.05 | 85 | 0 | 39.89 | 40 | 106.42 | 0 | 13.3675 | 12.995 | 0 |  | 1.0351 | 12.56 | 13.3578 | 0 | 33.07 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5679 | 1.5004
2008-08-04 22:34 | 0 | 85 | 85 | 0 | 40.03 | 40 | 106.3695 | 0 | 13.3619 | 13.0012 | 0 |  | 1.0346 | 12.56 | 13.3578 | 0 | 33.09 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5505 | 1.5004
2008-08-04 22:35 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 106.42 | 0 | 13.368 | 13.0016 | 0 |  | 1.0352 | 12.56 | 13.3578 | 0 | 33.11 | 1.6099 | 0 | 0.067218 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 25.2659 | -1 | 25.0126 | 12.5564 | 1.5004
2008-08-04 22:36 | 0 | 84.51 | 85 | 0 | 39.93 | 40 | 99.4495 | 0 | 7.2269 | 6.5115 | 1 |  | 0.96764 | 6.73 | 13.3578 | 0 | 32.69 | 0 | 0 | 0 | 20 | 20.021 |  | 21.2967 | 0 | 0 | 0 | 15 | 3.91 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0282 | -1 | 20.0067 | 12.5624 | 0
2008-08-04 22:37 | 0 | 85.23 | 85 | 0 | 39.93 | 40 | 99.4495 | 0 | 6.6068 | 6.3738 | 0 |  | 0.94639 | 6.73 | 6.6178 | 0 | 32.22 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 10.13 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0276 | -1 | 20.0067 | 6.7581 | 0
2008-08-04 22:38 | 0 | 84.91 | 85 | 0 | 40.03 | 40 | 99.4495 | 0 | 6.5321 | 6.3035 | 0 |  | 0.93605 | 6.73 | 6.5446 | 0 | 31.25 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 13.26 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0275 | -1 | 20.0067 | 6.7418 | 0
2008-08-04 22:39 | 0 | 84.92 | 85 | 0 | 39.97 | 40 | 99.6011 | 0 | 6.514 | 6.2851 | 0 |  | 0.93386 | 6.73 | 6.5324 | 0 | 30.16 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.37 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0275 | -1 | 20.0067 | 6.7529 | 0
2008-08-04 22:40 | 0 | 85.03 | 85 | 0 | 39.96 | 40 | 99.5505 | 0 | 6.5015 | 6.2736 | 0 |  | 0.93215 | 6.73 | 6.5079 | 0 | 29.62 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.75 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0275 | -1 | 20.0067 | 6.7518 | 0
2008-08-04 22:41 | 0 | 85.02 | 85 | 0 | 39.94 | 40 | 99.5 | 0 | 6.4949 | 6.264 | 0 |  | 0.93123 | 6.73 | 6.5079 | 0 | 29.38 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.89 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0275 | -1 | 20.0067 | 6.7433 | 0
2008-08-04 22:42 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 99.3485 | 0 | 6.4898 | 6.2625 | 0 |  | 0.92958 | 6.73 | 6.5079 | 0 | 29.38 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0275 | -1 | 20.0067 | 6.7486 | 0
2008-08-04 22:43 | 0 | 85 | 85 | 0 | 39.93 | 40 | 99.3485 | 0 | 6.4886 | 6.2614 | 0 |  | 0.92942 | 6.73 | 6.4957 | 0 | 29.42 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0275 | -1 | 20.0067 | 6.7564 | 0
2008-08-04 22:44 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 99.4495 | 0 | 6.4815 | 6.2546 | 0 |  | 0.9303 | 6.73 | 6.5079 | 0 | 29.54 | 0 | 0 | 0 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15.03 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7446 | 0
2008-08-04 22:45 | 0 | 84.97 | 85 | 0 | 40 | 40 | 99.4495 | 0 | 6.4877 | 6.2605 | 0 |  | 0.93024 | 6.73 | 6.4957 | 0 | 29.64 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.05 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0275 | -1 | 20.0067 | 6.7451 | 0
2008-08-04 22:46 | 0 | 85 | 85 | 0 | 39.95 | 40 | 99.3485 | 0 | 6.4808 | 6.2539 | 0 |  | 0.92926 | 6.73 | 6.5079 | 0 | 29.78 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.06 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7578 | 0
2008-08-04 22:47 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 99.3485 | 0 | 6.4806 | 6.2601 | 0 |  | 0.92923 | 6.73 | 6.4957 | 0 | 29.9 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.749 | 0
2008-08-04 22:48 | 0 | 84.99 | 85 | 0 | 40 | 40 | 99.3485 | 0 | 6.4806 | 6.2537 | 0 |  | 0.92923 | 6.73 | 6.4957 | 0 | 30.01 | 0 | 0 | 0 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.758 | 0
2008-08-04 22:49 | 0 | 85 | 85 | 0 | 39.97 | 40 | 99.3485 | 0 | 6.4772 | 6.2537 | 0 |  | 0.92923 | 6.73 | 6.5079 | 0 | 30.07 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7455 | 0
2008-08-04 22:50 | 0 | 85.05 | 85 | 0 | 39.97 | 40 | 99.3485 | 0 | 6.4806 | 6.2537 | 0 |  | 0.92923 | 6.73 | 6.5079 | 0 | 30.16 | 0 | 0 | 0 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.758 | 0
2008-08-04 22:51 | 0 | 84.97 | 85 | 0 | 39.86 | 40 | 99.4495 | 0 | 6.4806 | 6.2537 | 0 |  | 0.93017 | 6.73 | 6.5079 | 0 | 30.21 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.758 | 0
2008-08-04 22:52 | 0 | 85 | 85 | 0 | 39.93 | 40 | 99.3485 | 0 | 6.4806 | 6.2601 | 0 |  | 0.92923 | 6.73 | 6.4957 | 0 | 30.28 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7455 | 0
2008-08-04 22:53 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 99.3485 | 0 | 6.4808 | 6.2539 | 0 |  | 0.92926 | 6.73 | 6.5079 | 0 | 30.33 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.06 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7453 | 0
2008-08-04 22:54 | 0 | 85 | 85 | 0 | 39.95 | 40 | 99.3485 | 0 | 6.4808 | 6.2507 | 0 |  | 0.92926 | 6.73 | 6.5079 | 0 | 30.36 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.06 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7612 | 0
2008-08-04 22:55 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 99.298 | 0 | 6.4875 | 6.2539 | 0 |  | 0.92879 | 6.73 | 6.4957 | 0 | 30.4 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.06 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0275 | -1 | 20.0067 | 6.7487 | 0
2008-08-04 22:56 | 0 | 85 | 85 | 0 | 39.9 | 40 | 99.3485 | 0 | 6.4774 | 6.2603 | 0 |  | 0.92926 | 6.73 | 6.5079 | 0 | 30.44 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.06 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7451 | 0
2008-08-04 22:57 | 0 | 85 | 85 | 0 | 39.92 | 40 | 99.3485 | 0 | 6.4777 | 6.2541 | 0 |  | 0.92929 | 6.73 | 6.5079 | 0 | 30.46 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.05 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7451 | 0
2008-08-04 22:58 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 99.3485 | 0 | 6.481 | 6.251 | 0 |  | 0.92929 | 6.73 | 6.5079 | 0 | 30.47 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.05 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.761 | 0
2008-08-04 22:59 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 99.3485 | 0 | 6.481 | 6.2541 | 0 |  | 0.92929 | 6.73 | 6.4957 | 0 | 30.53 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.05 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7448 | 0
2008-08-04 23:00 | 0 | 84.89 | 85 | 0 | 40.05 | 40 | 99.298 | 0 | 6.4812 | 6.2544 | 0 |  | 0.92885 | 6.73 | 6.5079 | 0 | 30.54 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.04 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7448 | 0
2008-08-04 23:01 | 0 | 85.02 | 85 | 0 | 39.93 | 40 | 99.298 | 0 | 6.4779 | 6.2544 | 0 |  | 0.92885 | 6.73 | 6.4957 | 0 | 30.55 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.04 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7448 | 0
2008-08-04 23:02 | 0 | 84.97 | 85 | 0 | 40 | 40 | 99.298 | 0 | 6.4812 | 6.2544 | 0 |  | 0.92885 | 6.73 | 6.5079 | 0 | 30.57 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15.04 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7573 | 0
2008-08-04 23:03 | 0 | 84.96 | 85 | 0 | 39.89 | 40 | 99.3485 | 0 | 6.4815 | 6.2546 | 0 |  | 0.92936 | 6.73 | 6.4957 | 0 | 30.58 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.03 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7608 | 0
2008-08-04 23:04 | 0 | 85 | 85 | 0 | 39.91 | 40 | 99.3485 | 0 | 6.4781 | 6.2514 | 0 |  | 0.92936 | 6.73 | 6.4957 | 0 | 30.58 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.03 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7446 | 0
2008-08-04 23:05 | 0 | 84.92 | 85 | 0 | 40.07 | 40 | 99.3485 | 0 | 6.4817 | 6.2548 | 0 |  | 0.92939 | 6.73 | 6.4957 | 0 | 30.59 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7444 | 0
2008-08-04 23:06 | 0 | 84.97 | 85 | 0 | 40 | 40 | 99.298 | 0 | 6.4817 | 6.2548 | 0 |  | 0.92892 | 6.73 | 6.4957 | 0 | 30.62 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7444 | 0
2008-08-04 23:07 | 0 | 85 | 85 | 0 | 40.03 | 40 | 99.3485 | 0 | 6.4817 | 6.2516 | 0 |  | 0.92939 | 6.73 | 6.5079 | 0 | 30.61 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7444 | 0
2008-08-04 23:08 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 99.298 | 0 | 6.4783 | 6.2548 | 0 |  | 0.92892 | 6.73 | 6.4957 | 0 | 30.65 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7568 | 0
2008-08-04 23:09 | 0 | 84.95 | 85 | 0 | 40.1 | 40 | 99.3485 | 0 | 6.4817 | 6.2548 | 0 |  | 0.92939 | 6.73 | 6.4957 | 0 | 30.65 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7478 | 0
2008-08-04 23:10 | 0 | 85 | 85 | 0 | 40 | 40 | 99.3485 | 0 | 6.4783 | 6.2516 | 0 |  | 0.92939 | 6.73 | 6.4957 | 0 | 30.66 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7478 | 0
2008-08-04 23:11 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 99.3485 | 0 | 6.4819 | 6.255 | 0 |  | 0.92942 | 6.73 | 6.5079 | 0 | 30.66 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7568 | 0
2008-08-04 23:12 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 99.4495 | 0 | 6.4884 | 6.2548 | 0 |  | 0.93034 | 6.73 | 6.4957 | 0 | 30.66 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0275 | -1 | 20.0067 | 6.7441 | 0
2008-08-04 23:13 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 99.3485 | 0 | 6.4819 | 6.255 | 0 |  | 0.92942 | 6.73 | 6.4957 | 0 | 30.68 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7441 | 0
2008-08-04 23:14 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 99.3485 | 0 | 6.4819 | 6.255 | 0 |  | 0.92942 | 6.73 | 6.5079 | 0 | 30.65 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7566 | 0
2008-08-04 23:15 | 0 | 84.92 | 85 | 0 | 39.96 | 40 | 99.3485 | 0 | 6.4819 | 6.255 | 0 |  | 0.92942 | 6.73 | 6.4957 | 0 | 30.66 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7566 | 0
2008-08-04 23:16 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 99.3485 | 0 | 6.4819 | 6.255 | 0 |  | 0.92942 | 6.73 | 6.4957 | 0 | 30.62 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7566 | 0
2008-08-04 23:17 | 0 | 84.99 | 85 | 0 | 40 | 40 | 99.3485 | 0 | 6.4819 | 6.255 | 0 |  | 0.92942 | 6.73 | 6.4957 | 0 | 30.66 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7441 | 0
2008-08-04 23:18 | 0 | 84.92 | 85 | 0 | 39.98 | 40 | 99.3485 | 0 | 6.4819 | 6.255 | 0 |  | 0.92942 | 6.73 | 6.4957 | 0 | 30.66 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7441 | 0
2008-08-04 23:19 | 0 | 84.96 | 85 | 0 | 40 | 40 | 99.298 | 0 | 6.4786 | 6.255 | 0 |  | 0.92895 | 6.73 | 6.4957 | 0 | 30.69 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7566 | 0
2008-08-04 23:20 | 0 | 84.92 | 85 | 0 | 39.94 | 40 | 99.298 | 0 | 6.4822 | 6.2552 | 0 |  | 0.92898 | 6.73 | 6.4957 | 0 | 30.68 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7476 | 0
2008-08-04 23:21 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 104.147 | 0 | 6.4788 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.5079 | 0 | 30.66 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0274 | -1 | 20.0067 | 6.7439 | 2.7021
2008-08-04 23:22 | 0 | 85.05 | 85 | 0 | 39.94 | 40 | 104.147 | 0 | 6.8066 | 6.5635 | 0 |  | 0.97431 | 6.73 | 6.8132 | 0 | 30.71 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7367 | 2.7021
2008-08-04 23:23 | 0 | 85 | 85 | 0 | 40.04 | 40 | 104.248 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 30.73 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-04 23:24 | 0 | 85.02 | 85 | 0 | 40.03 | 40 | 104.147 | 0 | 6.8064 | 6.5635 | 0 |  | 0.97431 | 6.73 | 6.8132 | 0 | 30.74 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7483 | 2.7021
2008-08-04 23:25 | 0 | 85 | 85 | 0 | 39.89 | 40 | 104.147 | 0 | 6.8066 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 30.73 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7364 | 2.7021
2008-08-04 23:26 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 6.7997 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 30.75 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-04 23:27 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 6.8066 | 6.5637 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 30.75 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-04 23:28 | 0 | 85 | 85 | 0 | 39.94 | 40 | 104.147 | 0 | 6.7997 | 6.5571 | 0 |  | 0.97431 | 6.73 | 6.8254 | 0 | 30.8 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-04 23:29 | 0 | 85.05 | 85 | 0 | 40.03 | 40 | 104.248 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 30.84 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-04 23:30 | 0 | 85 | 85 | 0 | 40 | 40 | 104.248 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97529 | 6.73 | 6.8254 | 0 | 30.87 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7483 | 2.7021
2008-08-04 23:31 | 0 | 84.92 | 85 | 0 | 39.92 | 40 | 104.0965 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97384 | 6.73 | 6.8254 | 0 | 30.87 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7367 | 2.7021
2008-08-04 23:32 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 104.0965 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97384 | 6.73 | 6.801 | 0 | 30.89 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7432 | 2.7021
2008-08-04 23:33 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 6.7963 | 6.5571 | 0 |  | 0.97431 | 6.73 | 6.8132 | 0 | 30.91 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7465 | 2.7021
2008-08-04 23:34 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.0965 | 0 | 6.7997 | 6.5574 | 0 |  | 0.97384 | 6.73 | 6.8132 | 0 | 30.95 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-04 23:35 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 104.0965 | 0 | 6.7997 | 6.5571 | 0 |  | 0.97384 | 6.73 | 6.8132 | 0 | 30.95 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7465 | 2.7021
2008-08-04 23:36 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 104.0965 | 0 | 6.7997 | 6.5539 | 0 |  | 0.97384 | 6.73 | 6.8132 | 0 | 30.95 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7465 | 2.7021
2008-08-04 23:37 | 0 | 85.03 | 85 | 0 | 39.93 | 40 | 104.0965 | 0 | 6.7997 | 6.5571 | 0 |  | 0.97384 | 6.73 | 6.8132 | 0 | 30.95 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7465 | 2.7021
2008-08-04 23:38 | 0 | 85 | 85 | 0 | 40.05 | 40 | 104.046 | 0 | 6.7963 | 6.5571 | 0 |  | 0.97337 | 6.73 | 6.8132 | 0 | 30.98 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.956 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7432 | 2.7021
2008-08-04 23:39 | 0 | 85.08 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 6.7963 | 6.5571 | 0 |  | 0.97431 | 6.73 | 6.8254 | 0 | 31.01 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7465 | 2.7021
2008-08-04 23:40 | 0 | 85 | 85 | 0 | 40 | 40 | 104.0965 | 0 | 6.7999 | 6.5539 | 0 |  | 0.97384 | 6.73 | 6.8254 | 0 | 31.01 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-04 23:41 | 0 | 84.92 | 85 | 0 | 40 | 40 | 104.0965 | 0 | 6.7966 | 6.5571 | 0 |  | 0.97384 | 6.73 | 6.8132 | 0 | 31.03 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7465 | 2.7021
2008-08-04 23:42 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 104.046 | 0 | 6.7963 | 6.5539 | 0 |  | 0.97337 | 6.73 | 6.8132 | 0 | 31.03 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7465 | 2.7021
2008-08-04 23:43 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 6.7963 | 6.5571 | 0 |  | 0.97431 | 6.73 | 6.8132 | 0 | 31.03 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-04 23:44 | 0 | 85 | 85 | 0 | 40.07 | 40 | 104.0965 | 0 | 6.7997 | 6.5539 | 0 |  | 0.97384 | 6.73 | 6.8254 | 0 | 31.05 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-04 23:45 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.07 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-04 23:46 | 0 | 85.03 | 85 | 0 | 39.9 | 40 | 104.0965 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.09 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-04 23:47 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 104.147 | 0 | 6.7963 | 6.5571 | 0 |  | 0.97431 | 6.73 | 6.8132 | 0 | 31.12 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-04 23:48 | 0 | 85 | 85 | 0 | 39.86 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.1 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-04 23:49 | 0 | 84.99 | 85 | 0 | 40 | 40 | 104.147 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.13 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7364 | 2.7021
2008-08-04 23:50 | 0 | 84.97 | 85 | 0 | 39.89 | 40 | 104.248 | 0 | 6.8066 | 6.5542 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 31.14 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-04 23:51 | 0 | 84.91 | 85 | 0 | 40 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.14 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-04 23:52 | 0 | 85 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.15 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-04 23:53 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.18 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7483 | 2.7021
2008-08-04 23:54 | 0 | 85.05 | 85 | 0 | 40.04 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.15 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-04 23:55 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 104.0965 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.17 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-04 23:56 | 0 | 85.02 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.17 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-04 23:57 | 0 | 84.96 | 85 | 0 | 39.91 | 40 | 104.0965 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8254 | 0 | 31.21 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7432 | 2.7021
2008-08-04 23:58 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 104.0965 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.16 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-04 23:59 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.18 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 104.147 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.2 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 00:01 | 0 | 84.96 | 85 | 0 | 40 | 40 | 104.147 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.18 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:02 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 104.0965 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.2 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:03 | 0 | 84.99 | 85 | 0 | 40.07 | 40 | 104.0965 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8254 | 0 | 31.2 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:04 | 0 | 85 | 85 | 0 | 39.89 | 40 | 104.147 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.21 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 00:05 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.21 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:06 | 0 | 85.05 | 85 | 0 | 39.98 | 40 | 104.0965 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:07 | 0 | 84.92 | 85 | 0 | 39.98 | 40 | 104.0965 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.2 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:08 | 0 | 84.94 | 85 | 0 | 39.93 | 40 | 104.046 | 0 | 6.7999 | 6.5542 | 0 |  | 0.9734 | 6.73 | 6.8254 | 0 | 31.22 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:09 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 104.147 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.22 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:10 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 6.7966 | 6.5576 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.2 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7427 | 2.7021
2008-08-05 00:11 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 104.0965 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8254 | 0 | 31.22 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:12 | 0 | 85 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.2 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 00:13 | 0 | 85.03 | 85 | 0 | 39.97 | 40 | 104.0965 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8254 | 0 | 31.22 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:14 | 0 | 85 | 85 | 0 | 39.93 | 40 | 104.0965 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:15 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 104.147 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.21 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 00:16 | 0 | 85.03 | 85 | 0 | 40.1 | 40 | 104.0965 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:17 | 0 | 85 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:18 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.248 | 0 | 6.8001 | 6.5637 | 0 |  | 0.97529 | 6.73 | 6.8254 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7483 | 2.7021
2008-08-05 00:19 | 0 | 84.92 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 00:20 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 104.248 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97529 | 6.73 | 6.8254 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:21 | 0 | 85 | 85 | 0 | 40.13 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 00:22 | 0 | 84.92 | 85 | 0 | 40 | 40 | 104.147 | 0 | 6.8066 | 6.5637 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7364 | 2.7021
2008-08-05 00:23 | 0 | 84.92 | 85 | 0 | 39.89 | 40 | 104.147 | 0 | 6.8001 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.26 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:24 | 0 | 85 | 85 | 0 | 40 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.22 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7427 | 2.7021
2008-08-05 00:25 | 0 | 85 | 85 | 0 | 39.92 | 40 | 104.147 | 0 | 6.8001 | 6.5637 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.22 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7546 | 2.7021
2008-08-05 00:26 | 0 | 85 | 85 | 0 | 39.93 | 40 | 104.147 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:27 | 0 | 85.02 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 6.8001 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.26 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:28 | 0 | 85.05 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 6.8066 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:29 | 0 | 85 | 85 | 0 | 40 | 40 | 104.147 | 0 | 6.8001 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7427 | 2.7021
2008-08-05 00:30 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 104.248 | 0 | 6.8066 | 6.5574 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:31 | 0 | 85.02 | 85 | 0 | 40.13 | 40 | 104.147 | 0 | 6.8066 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7546 | 2.7021
2008-08-05 00:32 | 0 | 84.94 | 85 | 0 | 39.98 | 40 | 104.0965 | 0 | 6.8001 | 6.5576 | 0 |  | 0.97391 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:33 | 0 | 85 | 85 | 0 | 40 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:34 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:35 | 0 | 84.95 | 85 | 0 | 39.89 | 40 | 104.0965 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8254 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:36 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97438 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:37 | 0 | 84.94 | 85 | 0 | 39.91 | 40 | 104.0965 | 0 | 6.7968 | 6.5576 | 0 |  | 0.97391 | 6.73 | 6.8254 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:38 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.147 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 00:39 | 0 | 85.03 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 6.8001 | 6.5576 | 0 |  | 0.97438 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:40 | 0 | 84.95 | 85 | 0 | 40.04 | 40 | 104.0965 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:41 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.0965 | 0 | 6.7999 | 6.5576 | 0 |  | 0.97391 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:42 | 0 | 85.05 | 85 | 0 | 39.9 | 40 | 104.0965 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:43 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 104.0965 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.22 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7483 | 2.7021
2008-08-05 00:44 | 0 | 84.88 | 85 | 0 | 40 | 40 | 104.147 | 0 | 6.7968 | 6.5576 | 0 |  | 0.97438 | 6.73 | 6.8254 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 00:45 | 0 | 84.92 | 85 | 0 | 39.94 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 00:46 | 0 | 84.92 | 85 | 0 | 40.01 | 40 | 104.0965 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:47 | 0 | 85 | 85 | 0 | 39.9 | 40 | 104.147 | 0 | 6.8068 | 6.5639 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.26 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:48 | 0 | 85 | 85 | 0 | 39.94 | 40 | 104.248 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:49 | 0 | 84.92 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 6.8066 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 00:50 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 104.248 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:51 | 0 | 84.88 | 85 | 0 | 39.94 | 40 | 104.147 | 0 | 6.8066 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7427 | 2.7021
2008-08-05 00:52 | 0 | 85.05 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7427 | 2.7021
2008-08-05 00:53 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.0965 | 0 | 6.8066 | 6.5637 | 0 |  | 0.97387 | 6.73 | 6.8254 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7364 | 2.7021
2008-08-05 00:54 | 0 | 85 | 85 | 0 | 39.94 | 40 | 104.248 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:55 | 0 | 84.96 | 85 | 0 | 40.01 | 40 | 104.0965 | 0 | 6.8001 | 6.5576 | 0 |  | 0.97391 | 6.73 | 6.8132 | 0 | 31.26 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7483 | 2.7021
2008-08-05 00:56 | 0 | 85 | 85 | 0 | 39.96 | 40 | 104.0965 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8254 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7427 | 2.7021
2008-08-05 00:57 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 00:58 | 0 | 85 | 85 | 0 | 39.86 | 40 | 104.147 | 0 | 6.8066 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.26 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 00:59 | 0 | 85.03 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 6.8066 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:00 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 104.248 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97529 | 6.73 | 6.8254 | 0 | 31.26 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:01 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.248 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97529 | 6.73 | 6.8254 | 0 | 31.26 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 01:02 | 0 | 84.99 | 85 | 0 | 40 | 40 | 104.0965 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 01:03 | 0 | 85.02 | 85 | 0 | 39.91 | 40 | 104.147 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.746 | 2.7021
2008-08-05 01:04 | 0 | 84.95 | 85 | 0 | 40.04 | 40 | 104.0965 | 0 | 6.7966 | 6.5637 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:05 | 0 | 84.97 | 85 | 0 | 40.23 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.26 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:06 | 0 | 85.03 | 85 | 0 | 39.9 | 40 | 104.147 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 01:07 | 0 | 85 | 85 | 0 | 40.05 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 01:08 | 0 | 84.96 | 85 | 0 | 40.17 | 40 | 104.147 | 0 | 6.8001 | 6.5576 | 0 |  | 0.97438 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:09 | 0 | 84.94 | 85 | 0 | 39.83 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.26 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:10 | 0 | 85.02 | 85 | 0 | 39.93 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9634 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 01:11 | 0 | 84.92 | 85 | 0 | 39.83 | 40 | 104.147 | 0 | 6.8001 | 6.5576 | 0 |  | 0.97438 | 6.73 | 6.8132 | 0 | 31.26 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:12 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 104.147 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:13 | 0 | 84.96 | 85 | 0 | 40.07 | 40 | 104.0965 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 01:14 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 104.0965 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:15 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.26 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:16 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 104.0965 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 01:17 | 0 | 85 | 85 | 0 | 39.94 | 40 | 104.147 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:18 | 0 | 84.92 | 85 | 0 | 40.04 | 40 | 104.248 | 0 | 6.8066 | 6.5574 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:19 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 01:20 | 0 | 84.91 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 6.7999 | 6.5576 | 0 |  | 0.97438 | 6.73 | 6.8254 | 0 | 31.26 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:21 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 104.147 | 0 | 6.8001 | 6.5576 | 0 |  | 0.97438 | 6.73 | 6.8132 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7546 | 2.7021
2008-08-05 01:22 | 0 | 85 | 85 | 0 | 39.97 | 40 | 104.248 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97529 | 6.73 | 6.8254 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 01:23 | 0 | 85 | 85 | 0 | 40 | 40 | 104.248 | 0 | 6.8066 | 6.5574 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7427 | 2.7021
2008-08-05 01:24 | 0 | 84.97 | 85 | 0 | 40.07 | 40 | 104.147 | 0 | 6.8066 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.31 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7364 | 2.7021
2008-08-05 01:25 | 0 | 85.02 | 85 | 0 | 40.04 | 40 | 104.147 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:26 | 0 | 85 | 85 | 0 | 39.91 | 40 | 104.0965 | 0 | 6.8066 | 6.5637 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7364 | 2.7021
2008-08-05 01:27 | 0 | 84.97 | 85 | 0 | 39.84 | 40 | 104.147 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.24 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 01:28 | 0 | 84.92 | 85 | 0 | 39.93 | 40 | 104.147 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 01:29 | 0 | 84.96 | 85 | 0 | 39.91 | 40 | 104.147 | 0 | 6.8066 | 6.5576 | 0 |  | 0.97438 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 01:30 | 0 | 84.97 | 85 | 0 | 39.88 | 40 | 104.248 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97529 | 6.73 | 6.8254 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 01:31 | 0 | 84.92 | 85 | 0 | 39.95 | 40 | 104.248 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 31.31 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 01:32 | 0 | 85.05 | 85 | 0 | 40.01 | 40 | 104.0965 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 01:33 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 01:34 | 0 | 84.95 | 85 | 0 | 40 | 40 | 104.0965 | 0 | 6.7935 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:35 | 0 | 84.97 | 85 | 0 | 40.07 | 40 | 104.147 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:36 | 0 | 84.97 | 85 | 0 | 40.15 | 40 | 104.0965 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.746 | 2.7021
2008-08-05 01:37 | 0 | 85.03 | 85 | 0 | 39.93 | 40 | 104.0965 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.26 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:38 | 0 | 85.03 | 85 | 0 | 39.95 | 40 | 104.0965 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8254 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 01:39 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 104.0965 | 0 | 6.7968 | 6.5576 | 0 |  | 0.97391 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:40 | 0 | 84.95 | 85 | 0 | 40.01 | 40 | 104.0965 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8254 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 01:41 | 0 | 84.92 | 85 | 0 | 39.93 | 40 | 104.0965 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8254 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 01:42 | 0 | 85.02 | 85 | 0 | 39.93 | 40 | 104.0965 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 01:43 | 0 | 85.07 | 85 | 0 | 39.9 | 40 | 104.0965 | 0 | 6.7966 | 6.5544 | 0 |  | 0.97391 | 6.73 | 6.8132 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 01:44 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.31 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 01:45 | 0 | 84.94 | 85 | 0 | 40.08 | 40 | 104.147 | 0 | 6.7968 | 6.5576 | 0 |  | 0.97438 | 6.73 | 6.8132 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7579 | 2.7021
2008-08-05 01:46 | 0 | 85 | 85 | 0 | 39.89 | 40 | 104.0965 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 01:47 | 0 | 84.94 | 85 | 0 | 40.03 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.31 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7427 | 2.7021
2008-08-05 01:48 | 0 | 85.05 | 85 | 0 | 39.9 | 40 | 104.147 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.31 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7481 | 2.7021
2008-08-05 01:49 | 0 | 85 | 85 | 0 | 39.96 | 40 | 104.0965 | 0 | 6.7999 | 6.5576 | 0 |  | 0.97391 | 6.73 | 6.8254 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:50 | 0 | 84.97 | 85 | 0 | 39.91 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:51 | 0 | 85.02 | 85 | 0 | 40.03 | 40 | 104.147 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.31 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 01:52 | 0 | 84.96 | 85 | 0 | 40 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 01:53 | 0 | 84.97 | 85 | 0 | 40 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97438 | 6.73 | 6.8254 | 0 | 31.31 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:54 | 0 | 85.02 | 85 | 0 | 40.05 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.25 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:55 | 0 | 84.94 | 85 | 0 | 40.05 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:56 | 0 | 84.92 | 85 | 0 | 39.84 | 40 | 104.0965 | 0 | 6.7999 | 6.5544 | 0 |  | 0.97391 | 6.73 | 6.8132 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 01:57 | 0 | 84.92 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 6.8066 | 6.5637 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 01:58 | 0 | 84.95 | 85 | 0 | 39.84 | 40 | 104.0965 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 01:59 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 02:00 | 0 | 84.94 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 02:01 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 02:02 | 0 | 84.97 | 85 | 0 | 40.07 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 02:03 | 0 | 85.02 | 85 | 0 | 39.94 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:04 | 0 | 85 | 85 | 0 | 39.84 | 40 | 104.147 | 0 | 6.8066 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7483 | 2.7021
2008-08-05 02:05 | 0 | 84.95 | 85 | 0 | 40.01 | 40 | 104.248 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 02:06 | 0 | 84.96 | 85 | 0 | 40.08 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 02:07 | 0 | 85 | 85 | 0 | 39.93 | 40 | 104.147 | 0 | 6.8066 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.31 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 02:08 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.248 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97529 | 6.73 | 6.8254 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:09 | 0 | 84.99 | 85 | 0 | 40.05 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:10 | 0 | 84.92 | 85 | 0 | 39.92 | 40 | 104.147 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.31 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:11 | 0 | 84.92 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:12 | 0 | 85 | 85 | 0 | 40.13 | 40 | 104.147 | 0 | 6.8066 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:13 | 0 | 84.95 | 85 | 0 | 39.89 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 02:14 | 0 | 84.92 | 85 | 0 | 40.07 | 40 | 104.0965 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.33 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9634 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7364 | 2.7021
2008-08-05 02:15 | 0 | 85.05 | 85 | 0 | 39.88 | 40 | 104.248 | 0 | 6.7966 | 6.5637 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7364 | 2.7021
2008-08-05 02:16 | 0 | 85 | 85 | 0 | 39.93 | 40 | 104.147 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7362 | 2.7021
2008-08-05 02:17 | 0 | 85.03 | 85 | 0 | 39.97 | 40 | 104.147 | 0 | 6.7968 | 6.5576 | 0 |  | 0.97438 | 6.73 | 6.8254 | 0 | 31.28 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.746 | 2.7021
2008-08-05 02:18 | 0 | 85 | 85 | 0 | 40.07 | 40 | 104.248 | 0 | 6.7966 | 6.5542 | 0 |  | 0.97529 | 6.73 | 6.8254 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 02:19 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 104.248 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7483 | 2.7021
2008-08-05 02:20 | 0 | 84.99 | 85 | 0 | 40 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:21 | 0 | 85.02 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 6.7999 | 6.5542 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:22 | 0 | 84.99 | 85 | 0 | 40.05 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 02:23 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.31 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:24 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7364 | 2.7021
2008-08-05 02:25 | 0 | 85 | 85 | 0 | 39.97 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.31 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:26 | 0 | 84.94 | 85 | 0 | 40.01 | 40 | 104.248 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97529 | 6.73 | 6.8254 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 02:27 | 0 | 84.92 | 85 | 0 | 39.97 | 40 | 104.147 | 0 | 6.8066 | 6.5637 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7581 | 2.7021
2008-08-05 02:28 | 0 | 85.03 | 85 | 0 | 40.03 | 40 | 104.147 | 0 | 6.7999 | 6.5637 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.33 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7462 | 2.7021
2008-08-05 02:29 | 0 | 84.94 | 85 | 0 | 40.07 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.35 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:30 | 0 | 84.89 | 85 | 0 | 39.93 | 40 | 104.248 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97529 | 6.73 | 6.8132 | 0 | 31.35 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:31 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 104.0965 | 0 | 6.7966 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:32 | 0 | 84.94 | 85 | 0 | 39.9 | 40 | 104.0965 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8132 | 0 | 31.33 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7364 | 2.7021
2008-08-05 02:33 | 0 | 85.03 | 85 | 0 | 40.05 | 40 | 104.0965 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97387 | 6.73 | 6.8254 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7549 | 2.7021
2008-08-05 02:34 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8132 | 0 | 31.29 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.743 | 2.7021
2008-08-05 02:35 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 104.147 | 0 | 6.7999 | 6.5574 | 0 |  | 0.97435 | 6.73 | 6.8254 | 0 | 31.32 | 2.8992 | 0 | 0.12105 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 20.0278 | -1 | 20.0067 | 6.7579 | 2.7021
2008-08-05 02:36 | 0 | 84.96 | 85 | 0 | 39.91 | 40 | 104.147 | 0 | 14.1973 | 13.8459 | 1 |  | 0.97438 | 14.21 | 6.8132 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 6.7462 | 5.7032
2008-08-05 02:37 | 0 | 85.09 | 85 | 0 | 39.91 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.232 | 5.7032
2008-08-05 02:38 | 0 | 84.95 | 85 | 0 | 40.01 | 40 | 104.248 | 0 | 14.2167 | 13.8656 | 0 |  | 0.97533 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1987 | 5.7032
2008-08-05 02:39 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 02:40 | 0 | 85 | 85 | 0 | 39.92 | 40 | 104.248 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 02:41 | 0 | 85.07 | 85 | 0 | 39.97 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 02:42 | 0 | 85 | 85 | 0 | 39.86 | 40 | 104.248 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 02:43 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2189 | 5.7032
2008-08-05 02:44 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 104.147 | 0 | 14.2034 | 13.8594 | 0 |  | 0.97438 | 14.21 | 14.2125 | 0 | 31.28 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 02:45 | 0 | 85 | 85 | 0 | 39.95 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.26 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 02:46 | 0 | 84.92 | 85 | 0 | 39.97 | 40 | 104.248 | 0 | 14.2172 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 02:47 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 104.2985 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97576 | 14.21 | 14.2125 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 02:48 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 02:49 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.28 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 02:50 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 104.147 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 02:51 | 0 | 84.92 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 02:52 | 0 | 85.02 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 02:53 | 0 | 84.96 | 85 | 0 | 40.08 | 40 | 104.147 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 02:54 | 0 | 84.97 | 85 | 0 | 40.07 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 02:55 | 0 | 85 | 85 | 0 | 39.91 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 02:56 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 104.248 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2189 | 5.7032
2008-08-05 02:57 | 0 | 85.05 | 85 | 0 | 39.92 | 40 | 104.248 | 0 | 14.2034 | 13.8388 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 02:58 | 0 | 84.94 | 85 | 0 | 40.01 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 02:59 | 0 | 85 | 85 | 0 | 40.14 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 03:00 | 0 | 84.89 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 14.2172 | 13.8594 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2062 | 5.7032
2008-08-05 03:01 | 0 | 85.05 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 14.2172 | 13.8459 | 0 |  | 0.97438 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 03:02 | 0 | 85 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 03:03 | 0 | 84.97 | 85 | 0 | 40 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 03:04 | 0 | 84.95 | 85 | 0 | 40.07 | 40 | 104.0965 | 0 | 14.2034 | 13.8459 | 0 |  | 0.97391 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:05 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:06 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 104.248 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 03:07 | 0 | 85.05 | 85 | 0 | 39.98 | 40 | 104.248 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 03:08 | 0 | 85 | 85 | 0 | 40 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 03:09 | 0 | 84.92 | 85 | 0 | 40.07 | 40 | 104.248 | 0 | 14.2172 | 13.8459 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2062 | 5.7032
2008-08-05 03:10 | 0 | 85.02 | 85 | 0 | 40.03 | 40 | 104.248 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.28 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 03:11 | 0 | 85 | 85 | 0 | 39.83 | 40 | 104.248 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:12 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 14.2039 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9634 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2189 | 5.7032
2008-08-05 03:13 | 0 | 84.97 | 85 | 0 | 39.84 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:14 | 0 | 85.05 | 85 | 0 | 40.08 | 40 | 104.248 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 03:15 | 0 | 85 | 85 | 0 | 39.84 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 03:16 | 0 | 85 | 85 | 0 | 39.94 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.232 | 5.7032
2008-08-05 03:17 | 0 | 84.96 | 85 | 0 | 39.87 | 40 | 104.248 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97533 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:18 | 0 | 84.95 | 85 | 0 | 40 | 40 | 104.147 | 0 | 14.2039 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 03:19 | 0 | 85 | 85 | 0 | 40.11 | 40 | 104.248 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.232 | 5.7032
2008-08-05 03:20 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 03:21 | 0 | 85.02 | 85 | 0 | 39.95 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:22 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:23 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 03:24 | 0 | 85 | 85 | 0 | 39.97 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:25 | 0 | 85 | 85 | 0 | 39.8 | 40 | 104.0965 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97387 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 03:26 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 14.1973 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.36 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 03:27 | 0 | 85.02 | 85 | 0 | 40.01 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:28 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 03:29 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 14.1968 | 13.8392 | 0 |  | 0.97438 | 14.21 | 14.2002 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2263 | 5.7032
2008-08-05 03:30 | 0 | 85.02 | 85 | 0 | 39.93 | 40 | 104.147 | 0 | 14.2034 | 13.8388 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.36 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2263 | 5.7032
2008-08-05 03:31 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 104.0965 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97387 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2136 | 5.7032
2008-08-05 03:32 | 0 | 85.02 | 85 | 0 | 39.92 | 40 | 104.0965 | 0 | 14.2039 | 13.8589 | 0 |  | 0.97387 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:33 | 0 | 85.05 | 85 | 0 | 39.93 | 40 | 104.0965 | 0 | 14.2034 | 13.8388 | 0 |  | 0.97387 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 03:34 | 0 | 84.94 | 85 | 0 | 39.96 | 40 | 104.0965 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97387 | 14.21 | 14.2002 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 03:35 | 0 | 84.99 | 85 | 0 | 39.91 | 40 | 104.0965 | 0 | 14.2167 | 13.8388 | 0 |  | 0.97387 | 14.21 | 14.2125 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:36 | 0 | 85.05 | 85 | 0 | 39.91 | 40 | 104.147 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 03:37 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 104.147 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 03:38 | 0 | 84.92 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 14.1973 | 13.8594 | 0 |  | 0.97438 | 14.21 | 14.2002 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1924 | 5.7032
2008-08-05 03:39 | 0 | 85.02 | 85 | 0 | 39.9 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 03:40 | 0 | 85.02 | 85 | 0 | 39.84 | 40 | 104.248 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 03:41 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:42 | 0 | 85 | 85 | 0 | 39.97 | 40 | 104.248 | 0 | 14.2172 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 03:43 | 0 | 84.99 | 85 | 0 | 40.1 | 40 | 104.147 | 0 | 14.2039 | 13.8459 | 0 |  | 0.97438 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 03:44 | 0 | 84.94 | 85 | 0 | 39.97 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:45 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2247 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 03:46 | 0 | 84.94 | 85 | 0 | 39.91 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:47 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:48 | 0 | 84.99 | 85 | 0 | 39.91 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.36 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 03:49 | 0 | 84.92 | 85 | 0 | 40.03 | 40 | 104.248 | 0 | 14.1968 | 13.8388 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.36 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:50 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 104.248 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:51 | 0 | 84.92 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 03:52 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2136 | 5.7032
2008-08-05 03:53 | 0 | 84.92 | 85 | 0 | 39.96 | 40 | 104.0965 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97387 | 14.21 | 14.2125 | 0 | 31.36 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 03:54 | 0 | 85.05 | 85 | 0 | 40 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:55 | 0 | 85.05 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:56 | 0 | 84.94 | 85 | 0 | 39.94 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 03:57 | 0 | 84.99 | 85 | 0 | 40.05 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2263 | 5.7032
2008-08-05 03:58 | 0 | 85 | 85 | 0 | 40.05 | 40 | 104.147 | 0 | 14.1968 | 13.8392 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2258 | 5.7032
2008-08-05 03:59 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:00 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 104.0965 | 0 | 14.1968 | 13.8388 | 0 |  | 0.97387 | 14.21 | 14.2002 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:01 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 104.0965 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97387 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2136 | 5.7032
2008-08-05 04:02 | 0 | 84.96 | 85 | 0 | 40.05 | 40 | 104.0965 | 0 | 14.1968 | 13.8388 | 0 |  | 0.97387 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2136 | 5.7032
2008-08-05 04:03 | 0 | 84.94 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:04 | 0 | 84.91 | 85 | 0 | 40.13 | 40 | 104.0965 | 0 | 14.2039 | 13.8455 | 0 |  | 0.97387 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:05 | 0 | 85.05 | 85 | 0 | 40.05 | 40 | 104.0965 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97387 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:06 | 0 | 84.92 | 85 | 0 | 39.94 | 40 | 104.0965 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97387 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:07 | 0 | 84.94 | 85 | 0 | 39.98 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:08 | 0 | 85.05 | 85 | 0 | 40.03 | 40 | 104.0965 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97387 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 04:09 | 0 | 84.97 | 85 | 0 | 39.88 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2263 | 5.7032
2008-08-05 04:10 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:11 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 04:12 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.232 | 5.7032
2008-08-05 04:13 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.36 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:14 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:15 | 0 | 84.92 | 85 | 0 | 39.84 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:16 | 0 | 84.95 | 85 | 0 | 39.87 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:17 | 0 | 84.92 | 85 | 0 | 40.01 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:18 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 04:19 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 04:20 | 0 | 84.99 | 85 | 0 | 39.89 | 40 | 104.147 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:21 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 04:22 | 0 | 85.02 | 85 | 0 | 39.9 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.28 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:23 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.28 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:24 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 14.2034 | 13.8388 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:25 | 0 | 84.89 | 85 | 0 | 40.1 | 40 | 104.147 | 0 | 14.2034 | 13.8388 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.28 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:26 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 104.0965 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97387 | 14.21 | 14.2002 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:27 | 0 | 85.08 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:28 | 0 | 84.96 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:29 | 0 | 84.95 | 85 | 0 | 40.05 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:30 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 104.0965 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97387 | 14.21 | 14.2002 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:31 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:32 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 104.0965 | 0 | 14.2034 | 13.8388 | 0 |  | 0.97387 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2136 | 5.7032
2008-08-05 04:33 | 0 | 85.03 | 85 | 0 | 39.96 | 40 | 104.248 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:34 | 0 | 85.09 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 14.2034 | 13.8388 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2136 | 5.7032
2008-08-05 04:35 | 0 | 84.94 | 85 | 0 | 40.07 | 40 | 104.248 | 0 | 14.2172 | 13.8459 | 0 |  | 0.97533 | 14.21 | 14.2125 | 0 | 31.28 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 04:36 | 0 | 85.07 | 85 | 0 | 39.86 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2263 | 5.7032
2008-08-05 04:37 | 0 | 84.99 | 85 | 0 | 40.08 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2247 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 04:38 | 0 | 84.92 | 85 | 0 | 39.96 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2247 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.232 | 5.7032
2008-08-05 04:39 | 0 | 84.94 | 85 | 0 | 39.96 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:40 | 0 | 84.92 | 85 | 0 | 39.84 | 40 | 104.248 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.28 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 04:41 | 0 | 85.05 | 85 | 0 | 39.87 | 40 | 104.147 | 0 | 14.2172 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2189 | 5.7032
2008-08-05 04:42 | 0 | 84.97 | 85 | 0 | 40 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.28 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 04:43 | 0 | 85.07 | 85 | 0 | 39.98 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.28 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:44 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 04:45 | 0 | 84.95 | 85 | 0 | 39.91 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:46 | 0 | 84.91 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 04:47 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:48 | 0 | 84.96 | 85 | 0 | 40 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:49 | 0 | 85 | 85 | 0 | 39.93 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:50 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 14.2039 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:51 | 0 | 84.95 | 85 | 0 | 39.88 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:52 | 0 | 84.91 | 85 | 0 | 39.87 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:53 | 0 | 84.99 | 85 | 0 | 40.04 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:54 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 104.248 | 0 | 14.2172 | 13.8594 | 0 |  | 0.97533 | 14.21 | 14.2125 | 0 | 31.28 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2051 | 5.7032
2008-08-05 04:55 | 0 | 84.99 | 85 | 0 | 40 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2258 | 5.7032
2008-08-05 04:56 | 0 | 85 | 85 | 0 | 40.07 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 04:57 | 0 | 85.03 | 85 | 0 | 39.97 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 04:58 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 04:59 | 0 | 84.94 | 85 | 0 | 39.88 | 40 | 104.147 | 0 | 14.1973 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:00 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2136 | 5.7032
2008-08-05 05:01 | 0 | 84.94 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:02 | 0 | 85 | 85 | 0 | 39.92 | 40 | 104.248 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 05:03 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 05:04 | 0 | 84.96 | 85 | 0 | 40.07 | 40 | 104.147 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.28 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:05 | 0 | 84.99 | 85 | 0 | 40 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.26 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 05:06 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 104.147 | 0 | 14.1968 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 05:07 | 0 | 84.97 | 85 | 0 | 40 | 40 | 104.248 | 0 | 14.2034 | 13.8594 | 0 |  | 0.97533 | 14.21 | 14.2125 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:08 | 0 | 84.99 | 85 | 0 | 40.05 | 40 | 104.248 | 0 | 14.2234 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:09 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 05:10 | 0 | 85.03 | 85 | 0 | 40.07 | 40 | 104.147 | 0 | 14.2034 | 13.8594 | 0 |  | 0.97438 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2051 | 5.7032
2008-08-05 05:11 | 0 | 84.92 | 85 | 0 | 40.05 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:12 | 0 | 84.94 | 85 | 0 | 40 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 05:13 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 104.248 | 0 | 14.2039 | 13.8594 | 0 |  | 0.97533 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2189 | 5.7032
2008-08-05 05:14 | 0 | 84.89 | 85 | 0 | 40.03 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.36 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 05:15 | 0 | 84.94 | 85 | 0 | 39.96 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2189 | 5.7032
2008-08-05 05:16 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 104.248 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2182 | 5.7032
2008-08-05 05:17 | 0 | 85 | 85 | 0 | 40.15 | 40 | 104.248 | 0 | 14.2239 | 13.8656 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2051 | 5.7032
2008-08-05 05:18 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 104.248 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:19 | 0 | 85 | 85 | 0 | 40.07 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 05:20 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 104.0965 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97387 | 14.21 | 14.2125 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:21 | 0 | 84.94 | 85 | 0 | 39.97 | 40 | 104.147 | 0 | 14.2167 | 13.8656 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 05:22 | 0 | 84.92 | 85 | 0 | 40 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 05:23 | 0 | 84.94 | 85 | 0 | 39.88 | 40 | 104.147 | 0 | 14.1968 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 05:24 | 0 | 84.92 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 14.2039 | 13.8459 | 0 |  | 0.97438 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.232 | 5.7032
2008-08-05 05:25 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 104.248 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 05:26 | 0 | 84.95 | 85 | 0 | 39.87 | 40 | 104.147 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 05:27 | 0 | 85 | 85 | 0 | 39.89 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2067 | 5.7032
2008-08-05 05:28 | 0 | 84.94 | 85 | 0 | 39.94 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2263 | 5.7032
2008-08-05 05:29 | 0 | 85 | 85 | 0 | 40 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2136 | 5.7032
2008-08-05 05:30 | 0 | 85.02 | 85 | 0 | 40 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 05:31 | 0 | 84.91 | 85 | 0 | 40.07 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 05:32 | 0 | 84.97 | 85 | 0 | 39.87 | 40 | 104.248 | 0 | 14.2172 | 13.8594 | 0 |  | 0.97533 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:33 | 0 | 85.02 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 14.2039 | 13.8459 | 0 |  | 0.97438 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1924 | 5.7032
2008-08-05 05:34 | 0 | 84.95 | 85 | 0 | 40 | 40 | 104.248 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 05:35 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.147 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2136 | 5.7032
2008-08-05 05:36 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.147 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:37 | 0 | 85.02 | 85 | 0 | 40.07 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2002 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:38 | 0 | 84.99 | 85 | 0 | 40 | 40 | 104.248 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 05:39 | 0 | 84.94 | 85 | 0 | 39.96 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2247 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1929 | 5.7032
2008-08-05 05:40 | 0 | 85 | 85 | 0 | 40.04 | 40 | 104.248 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2002 | 0 | 31.36 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2182 | 5.7032
2008-08-05 05:41 | 0 | 85 | 85 | 0 | 40 | 40 | 104.248 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.36 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 05:42 | 0 | 84.99 | 85 | 0 | 39.87 | 40 | 104.147 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:43 | 0 | 85.02 | 85 | 0 | 40.01 | 40 | 104.147 | 0 | 14.2167 | 13.8589 | 0 |  | 0.97435 | 14.21 | 14.2125 | 0 | 31.29 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.1861 | 5.7032
2008-08-05 05:44 | 0 | 84.94 | 85 | 0 | 40 | 40 | 104.248 | 0 | 14.2172 | 13.8459 | 0 |  | 0.97533 | 14.21 | 14.2247 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:45 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 104.248 | 0 | 14.2034 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.35 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:46 | 0 | 84.99 | 85 | 0 | 40 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.32 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:47 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 104.248 | 0 | 14.2034 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 05:48 | 0 | 85.03 | 85 | 0 | 40.07 | 40 | 104.248 | 0 | 14.2167 | 13.8455 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2194 | 5.7032
2008-08-05 05:49 | 0 | 85.05 | 85 | 0 | 39.96 | 40 | 104.2985 | 0 | 14.2167 | 13.8656 | 0 |  | 0.97576 | 14.21 | 14.2125 | 0 | 31.31 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 05:50 | 0 | 84.96 | 85 | 0 | 40 | 40 | 104.248 | 0 | 14.2234 | 13.8589 | 0 |  | 0.97529 | 14.21 | 14.2125 | 0 | 31.33 | 6.1193 | 0 | 0.2555 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 20.0352 | -1 | 20.0142 | 14.2056 | 5.7032
2008-08-05 05:51 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 104.248 | 0 | 14.2167 | 23.1047 | 0 |  | 0.97529 | 23.69 | 14.2247 | 0 | 31.33 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0352 | -1 | 20.0237 | 14.2182 | 9.5036
2008-08-05 05:52 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 104.349 | 0 | 23.5242 | 23.1279 | 0 |  | 0.97627 | 23.69 | 23.5165 | 0 | 31.35 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 05:53 | 0 | 84.92 | 85 | 0 | 40 | 40 | 104.3995 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 31.36 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 05:54 | 0 | 84.94 | 85 | 0 | 39.97 | 40 | 104.349 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 31.36 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 05:55 | 0 | 84.97 | 85 | 0 | 40.1 | 40 | 104.349 | 0 | 23.5121 | 23.1167 | 0 |  | 0.97627 | 23.69 | 23.5043 | 0 | 31.39 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 05:56 | 0 | 84.99 | 85 | 0 | 40.08 | 40 | 104.2985 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97576 | 23.69 | 23.4921 | 0 | 31.36 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 05:57 | 0 | 84.92 | 85 | 0 | 39.93 | 40 | 104.2985 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 31.36 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 05:58 | 0 | 85.03 | 85 | 0 | 39.89 | 40 | 104.349 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 31.38 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 05:59 | 0 | 84.92 | 85 | 0 | 40 | 40 | 104.349 | 0 | 23.5121 | 23.1279 | 0 |  | 0.97627 | 23.69 | 23.5043 | 0 | 31.36 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 06:00 | 0 | 84.97 | 85 | 0 | 40 | 40 | 104.2985 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 31.35 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6584 | 9.5035
2008-08-05 06:01 | 0 | 85.03 | 85 | 0 | 39.95 | 40 | 104.349 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 31.32 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 06:02 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 104.3995 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5043 | 0 | 31.38 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 06:03 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 31.36 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 06:04 | 0 | 85 | 85 | 0 | 39.95 | 40 | 104.2985 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 31.33 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 06:05 | 0 | 85.02 | 85 | 0 | 39.94 | 40 | 104.2985 | 0 | 23.5234 | 23.1279 | 0 |  | 0.9758 | 23.69 | 23.5287 | 0 | 31.48 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 06:06 | 0 | 84.94 | 85 | 0 | 39.87 | 40 | 104.2985 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 31.72 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 06:07 | 0 | 85.03 | 85 | 0 | 39.93 | 40 | 104.349 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 31.94 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 06:08 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.349 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.13 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 06:09 | 0 | 85 | 85 | 0 | 40.01 | 40 | 104.2985 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.22 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 06:10 | 0 | 85 | 85 | 0 | 39.95 | 40 | 104.349 | 0 | 23.5121 | 23.1151 | 0 |  | 0.9762 | 23.69 | 23.5043 | 0 | 32.28 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5036
2008-08-05 06:11 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 104.349 | 0 | 23.5226 | 23.1263 | 0 |  | 0.9762 | 23.69 | 23.5165 | 0 | 32.25 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 06:12 | 0 | 84.95 | 85 | 0 | 39.91 | 40 | 104.2985 | 0 | 23.4999 | 23.1151 | 0 |  | 0.97573 | 23.69 | 23.5043 | 0 | 32.25 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5036
2008-08-05 06:13 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 104.2985 | 0 | 23.4999 | 23.1263 | 0 |  | 0.97573 | 23.69 | 23.5043 | 0 | 32.2 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5035
2008-08-05 06:14 | 0 | 85.07 | 85 | 0 | 39.95 | 40 | 104.349 | 0 | 23.5112 | 23.1039 | 0 |  | 0.9762 | 23.69 | 23.5043 | 0 | 32.21 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5036
2008-08-05 06:15 | 0 | 84.95 | 85 | 0 | 39.84 | 40 | 104.2985 | 0 | 23.5112 | 23.1039 | 0 |  | 0.97573 | 23.69 | 23.5043 | 0 | 32.17 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5036
2008-08-05 06:16 | 0 | 84.94 | 85 | 0 | 39.98 | 40 | 104.248 | 0 | 23.5112 | 23.1263 | 0 |  | 0.97526 | 23.69 | 23.5043 | 0 | 32.16 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6821 | 9.5035
2008-08-05 06:17 | 0 | 85.02 | 85 | 0 | 39.94 | 40 | 104.2985 | 0 | 23.5112 | 23.1151 | 0 |  | 0.97573 | 23.69 | 23.4921 | 0 | 32.21 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5036
2008-08-05 06:18 | 0 | 85.05 | 85 | 0 | 39.92 | 40 | 104.2985 | 0 | 23.5226 | 23.1151 | 0 |  | 0.97573 | 23.69 | 23.5043 | 0 | 32.31 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5036
2008-08-05 06:19 | 0 | 85 | 85 | 0 | 39.91 | 40 | 104.2985 | 0 | 23.4999 | 23.1263 | 0 |  | 0.97573 | 23.69 | 23.5043 | 0 | 32.45 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6936 | 9.5035
2008-08-05 06:20 | 0 | 84.92 | 85 | 0 | 39.92 | 40 | 104.2985 | 0 | 23.4999 | 23.1039 | 0 |  | 0.97573 | 23.69 | 23.4921 | 0 | 32.53 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5036
2008-08-05 06:21 | 0 | 85 | 85 | 0 | 39.88 | 40 | 104.2985 | 0 | 23.5226 | 23.1151 | 0 |  | 0.97573 | 23.69 | 23.5165 | 0 | 32.57 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5036
2008-08-05 06:22 | 0 | 84.95 | 85 | 0 | 39.92 | 40 | 104.2985 | 0 | 23.5226 | 23.1263 | 0 |  | 0.97573 | 23.69 | 23.5043 | 0 | 32.56 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 06:23 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 104.349 | 0 | 23.5226 | 23.1151 | 0 |  | 0.9762 | 23.69 | 23.5043 | 0 | 32.53 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6724 | 9.5036
2008-08-05 06:24 | 0 | 84.92 | 85 | 0 | 39.95 | 40 | 104.349 | 0 | 23.5112 | 23.1263 | 0 |  | 0.9762 | 23.69 | 23.5165 | 0 | 32.49 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 06:25 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.2985 | 0 | 23.5112 | 23.1263 | 0 |  | 0.97573 | 23.69 | 23.5043 | 0 | 32.45 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 06:26 | 0 | 84.94 | 85 | 0 | 39.8 | 40 | 104.349 | 0 | 23.5234 | 23.1151 | 0 |  | 0.9762 | 23.69 | 23.5165 | 0 | 32.43 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5036
2008-08-05 06:27 | 0 | 85.05 | 85 | 0 | 39.98 | 40 | 104.349 | 0 | 23.5112 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.38 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 06:28 | 0 | 85 | 85 | 0 | 39.96 | 40 | 104.2985 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97573 | 23.69 | 23.5043 | 0 | 32.32 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 06:29 | 0 | 84.95 | 85 | 0 | 40 | 40 | 104.349 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.4 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6592 | 9.5035
2008-08-05 06:30 | 0 | 84.91 | 85 | 0 | 40.03 | 40 | 104.2985 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.52 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 06:31 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 104.3995 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.68 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 06:32 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 104.2985 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.69 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 06:33 | 0 | 84.97 | 85 | 0 | 40.11 | 40 | 104.2985 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.69 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 06:34 | 0 | 84.96 | 85 | 0 | 40 | 40 | 104.2985 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 32.64 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 06:35 | 0 | 84.94 | 85 | 0 | 39.96 | 40 | 104.349 | 0 | 23.5112 | 23.1159 | 0 |  | 0.9762 | 23.69 | 23.5165 | 0 | 32.67 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 06:36 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 104.2985 | 0 | 23.5234 | 23.1151 | 0 |  | 0.97573 | 23.69 | 23.5043 | 0 | 32.6 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5036
2008-08-05 06:37 | 0 | 84.94 | 85 | 0 | 39.94 | 40 | 104.2985 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.56 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 06:38 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 104.2985 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.4921 | 0 | 32.51 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 06:39 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 104.2985 | 0 | 23.5007 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.48 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 06:40 | 0 | 84.97 | 85 | 0 | 39.91 | 40 | 104.2985 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 32.49 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 06:41 | 0 | 84.97 | 85 | 0 | 39.88 | 40 | 104.2985 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 32.64 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 06:42 | 0 | 84.95 | 85 | 0 | 40.01 | 40 | 104.2985 | 0 | 23.5234 | 23.1047 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 32.79 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6936 | 9.5036
2008-08-05 06:43 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.349 | 0 | 23.5234 | 23.1047 | 0 |  | 0.97529 | 23.69 | 23.5165 | 0 | 32.93 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5036
2008-08-05 06:44 | 0 | 84.89 | 85 | 0 | 40.11 | 40 | 104.349 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 32.95 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.706 | 9.5035
2008-08-05 06:45 | 0 | 84.99 | 85 | 0 | 40 | 40 | 104.349 | 0 | 23.5226 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 32.93 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5035
2008-08-05 06:46 | 0 | 84.99 | 85 | 0 | 39.91 | 40 | 104.349 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.4921 | 0 | 32.87 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6698 | 9.5035
2008-08-05 06:47 | 0 | 85 | 85 | 0 | 39.94 | 40 | 104.349 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.84 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 06:48 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 104.349 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.8 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 06:49 | 0 | 85.05 | 85 | 0 | 39.97 | 40 | 104.2985 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 32.76 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 06:50 | 0 | 85 | 85 | 0 | 40 | 40 | 104.349 | 0 | 23.5226 | 23.1151 | 0 |  | 0.9762 | 23.69 | 23.5165 | 0 | 32.73 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5036
2008-08-05 06:51 | 0 | 84.99 | 85 | 0 | 39.9 | 40 | 104.349 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 32.68 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 06:52 | 0 | 85.05 | 85 | 0 | 39.93 | 40 | 104.349 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.61 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 06:53 | 0 | 84.86 | 85 | 0 | 40.03 | 40 | 104.2985 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 32.62 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 06:54 | 0 | 84.94 | 85 | 0 | 39.94 | 40 | 104.349 | 0 | 23.5348 | 23.1279 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.67 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 06:55 | 0 | 85.07 | 85 | 0 | 39.97 | 40 | 104.349 | 0 | 23.5234 | 23.1382 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 32.72 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 06:56 | 0 | 85.03 | 85 | 0 | 39.97 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5409 | 0 | 32.76 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 06:57 | 0 | 85 | 85 | 0 | 39.91 | 40 | 104.3995 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.81 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 06:58 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 104.349 | 0 | 23.5234 | 23.1382 | 0 |  | 0.97627 | 23.69 | 23.5287 | 0 | 32.75 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 06:59 | 0 | 84.99 | 85 | 0 | 40 | 40 | 104.2985 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.72 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 07:00 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 104.3995 | 0 | 23.5348 | 23.1167 | 0 |  | 0.97674 | 23.69 | 23.5165 | 0 | 32.71 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 07:01 | 0 | 84.92 | 85 | 0 | 40.01 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.67 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 07:02 | 0 | 84.99 | 85 | 0 | 40 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 32.6 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 07:03 | 0 | 84.97 | 85 | 0 | 39.91 | 40 | 104.3995 | 0 | 23.5348 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5043 | 0 | 32.56 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 07:04 | 0 | 84.95 | 85 | 0 | 40.07 | 40 | 104.349 | 0 | 23.5234 | 23.1382 | 0 |  | 0.97624 | 23.69 | 23.5409 | 0 | 32.6 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 07:05 | 0 | 85 | 85 | 0 | 39.96 | 40 | 104.3995 | 0 | 23.5242 | 23.1391 | 0 |  | 0.97674 | 23.69 | 23.5165 | 0 | 32.64 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5035
2008-08-05 07:06 | 0 | 84.96 | 85 | 0 | 40.07 | 40 | 104.2985 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.72 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 07:07 | 0 | 84.97 | 85 | 0 | 40.11 | 40 | 104.2985 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.78 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 07:08 | 0 | 85 | 85 | 0 | 39.96 | 40 | 104.349 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 32.76 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 07:09 | 0 | 84.97 | 85 | 0 | 40.07 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.75 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 07:10 | 0 | 84.92 | 85 | 0 | 40 | 40 | 104.349 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.71 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 07:11 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 104.2985 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.67 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 07:12 | 0 | 84.97 | 85 | 0 | 39.91 | 40 | 104.2985 | 0 | 23.5121 | 23.1382 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 32.6 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 07:13 | 0 | 84.99 | 85 | 0 | 39.9 | 40 | 104.2985 | 0 | 23.5242 | 23.1382 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.55 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 07:14 | 0 | 85.05 | 85 | 0 | 39.98 | 40 | 104.3995 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.55 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 07:15 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.58 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 07:16 | 0 | 85.05 | 85 | 0 | 39.98 | 40 | 104.2985 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.64 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6936 | 9.5035
2008-08-05 07:17 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 104.2985 | 0 | 23.5242 | 23.1167 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.74 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 07:18 | 0 | 85 | 85 | 0 | 39.97 | 40 | 104.2985 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.74 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 07:19 | 0 | 84.97 | 85 | 0 | 40.01 | 40 | 104.3995 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.76 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 07:20 | 0 | 85.05 | 85 | 0 | 39.97 | 40 | 104.2985 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.73 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 07:21 | 0 | 85.02 | 85 | 0 | 39.91 | 40 | 104.349 | 0 | 23.5242 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.68 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 07:22 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 104.349 | 0 | 23.5348 | 23.1391 | 0 |  | 0.97627 | 23.69 | 23.5287 | 0 | 32.62 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 07:23 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 104.349 | 0 | 23.5348 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.61 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 07:24 | 0 | 84.92 | 85 | 0 | 40 | 40 | 104.3995 | 0 | 23.5242 | 23.1391 | 0 |  | 0.97674 | 23.69 | 23.5287 | 0 | 32.53 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 07:25 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 104.349 | 0 | 23.5234 | 23.1382 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.51 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 07:26 | 0 | 84.95 | 85 | 0 | 39.86 | 40 | 104.3995 | 0 | 23.5129 | 23.1279 | 0 |  | 0.97674 | 23.69 | 23.5165 | 0 | 32.47 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 07:27 | 0 | 84.92 | 85 | 0 | 39.96 | 40 | 104.3995 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.49 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 07:28 | 0 | 84.95 | 85 | 0 | 39.9 | 40 | 104.3995 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.54 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 07:29 | 0 | 85.03 | 85 | 0 | 39.92 | 40 | 104.3995 | 0 | 23.5584 | 23.1391 | 0 |  | 0.97674 | 23.69 | 23.5287 | 0 | 32.64 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 07:30 | 0 | 84.96 | 85 | 0 | 40 | 40 | 104.3995 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5409 | 0 | 32.77 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.7076 | 9.5035
2008-08-05 07:31 | 0 | 85 | 85 | 0 | 40 | 40 | 104.3995 | 0 | 23.5356 | 23.1391 | 0 |  | 0.97674 | 23.69 | 23.5165 | 0 | 32.77 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 07:32 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 104.3995 | 0 | 23.5348 | 23.1606 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.78 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 07:33 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 104.3995 | 0 | 23.5348 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.75 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 07:34 | 0 | 84.95 | 85 | 0 | 39.88 | 40 | 104.3995 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.73 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 07:35 | 0 | 84.94 | 85 | 0 | 40 | 40 | 104.3995 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.67 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6961 | 9.5035
2008-08-05 07:36 | 0 | 84.94 | 85 | 0 | 39.95 | 40 | 104.3995 | 0 | 23.5242 | 23.1279 | 0 |  | 0.97674 | 23.69 | 23.5165 | 0 | 32.64 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 07:37 | 0 | 84.94 | 85 | 0 | 39.88 | 40 | 104.3995 | 0 | 23.5348 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.58 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 07:38 | 0 | 85 | 85 | 0 | 40.04 | 40 | 104.349 | 0 | 23.5348 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.55 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.7068 | 9.5035
2008-08-05 07:39 | 0 | 84.94 | 85 | 0 | 39.91 | 40 | 104.2985 | 0 | 23.5234 | 23.1391 | 0 |  | 0.97576 | 23.69 | 23.5287 | 0 | 32.48 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 07:40 | 0 | 85 | 85 | 0 | 40.08 | 40 | 104.3995 | 0 | 23.5242 | 23.1391 | 0 |  | 0.97674 | 23.69 | 23.5165 | 0 | 32.46 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 07:41 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 104.2985 | 0 | 23.5348 | 23.1271 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 32.46 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 07:42 | 0 | 85.02 | 85 | 0 | 40.04 | 40 | 104.349 | 0 | 23.5121 | 23.1279 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.41 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 07:43 | 0 | 85 | 85 | 0 | 39.97 | 40 | 104.2985 | 0 | 23.5356 | 23.1279 | 0 |  | 0.9758 | 23.69 | 23.5287 | 0 | 32.36 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 07:44 | 0 | 84.92 | 85 | 0 | 40 | 40 | 104.349 | 0 | 23.5356 | 23.1391 | 0 |  | 0.97627 | 23.69 | 23.5043 | 0 | 32.42 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 07:45 | 0 | 84.95 | 85 | 0 | 40.05 | 40 | 104.349 | 0 | 23.5242 | 23.1279 | 0 |  | 0.97627 | 23.69 | 23.5165 | 0 | 32.57 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 07:46 | 0 | 85 | 85 | 0 | 40.04 | 40 | 104.2985 | 0 | 23.5242 | 23.1279 | 0 |  | 0.9758 | 23.69 | 23.5165 | 0 | 32.69 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 07:47 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.3995 | 0 | 23.5121 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.71 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 07:48 | 0 | 85 | 85 | 0 | 40 | 40 | 104.2985 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.71 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5035
2008-08-05 07:49 | 0 | 84.95 | 85 | 0 | 39.91 | 40 | 104.3995 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.67 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 07:50 | 0 | 85 | 85 | 0 | 40.04 | 40 | 104.349 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.6 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.7068 | 9.5035
2008-08-05 07:51 | 0 | 84.96 | 85 | 0 | 39.9 | 40 | 104.349 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.57 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 07:52 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 104.3995 | 0 | 23.5356 | 23.1391 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.56 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 07:53 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 104.3995 | 0 | 23.5234 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.53 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 07:54 | 0 | 84.92 | 85 | 0 | 40.03 | 40 | 104.349 | 0 | 23.5234 | 23.1382 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.55 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 07:55 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 104.3995 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.64 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 07:56 | 0 | 84.92 | 85 | 0 | 40 | 40 | 104.3995 | 0 | 23.5348 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.76 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 07:57 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 104.3995 | 0 | 23.5348 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.83 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 07:58 | 0 | 85 | 85 | 0 | 40 | 40 | 104.3995 | 0 | 23.5234 | 23.1606 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.88 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6609 | 9.5035
2008-08-05 07:59 | 0 | 85 | 85 | 0 | 39.93 | 40 | 104.3995 | 0 | 23.5576 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.86 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:00 | 0 | 84.92 | 85 | 0 | 40 | 40 | 104.3995 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.81 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 08:01 | 0 | 85 | 85 | 0 | 39.96 | 40 | 104.3995 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.78 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5035
2008-08-05 08:02 | 0 | 85 | 85 | 0 | 39.92 | 40 | 104.349 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.73 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 08:03 | 0 | 85 | 85 | 0 | 40.04 | 40 | 104.3995 | 0 | 23.5234 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.69 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 08:04 | 0 | 84.97 | 85 | 0 | 40 | 40 | 104.3995 | 0 | 23.5234 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.71 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 08:05 | 0 | 84.96 | 85 | 0 | 40 | 40 | 104.349 | 0 | 23.5348 | 23.1614 | 0 |  | 0.97627 | 23.69 | 23.5165 | 0 | 32.72 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:06 | 0 | 84.94 | 85 | 0 | 39.92 | 40 | 104.3995 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.76 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 08:07 | 0 | 84.96 | 85 | 0 | 40.07 | 40 | 104.3995 | 0 | 23.5121 | 23.1279 | 0 |  | 0.97674 | 23.69 | 23.5287 | 0 | 32.85 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 08:08 | 0 | 84.99 | 85 | 0 | 39.87 | 40 | 104.349 | 0 | 23.5348 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.85 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:09 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.3995 | 0 | 23.5234 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.82 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:10 | 0 | 84.97 | 85 | 0 | 40 | 40 | 104.349 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.79 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 08:11 | 0 | 85.02 | 85 | 0 | 40.03 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.72 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 08:12 | 0 | 84.97 | 85 | 0 | 39.91 | 40 | 104.349 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.72 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:13 | 0 | 84.92 | 85 | 0 | 40 | 40 | 104.3995 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.68 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6707 | 9.5035
2008-08-05 08:14 | 0 | 84.94 | 85 | 0 | 39.92 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.61 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5035
2008-08-05 08:15 | 0 | 85.07 | 85 | 0 | 39.96 | 40 | 104.349 | 0 | 23.5356 | 23.1167 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.61 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 08:16 | 0 | 84.94 | 85 | 0 | 39.98 | 40 | 104.3995 | 0 | 23.5234 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.68 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 08:17 | 0 | 84.99 | 85 | 0 | 39.9 | 40 | 104.349 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.75 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5035
2008-08-05 08:18 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.78 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:19 | 0 | 84.88 | 85 | 0 | 40.03 | 40 | 104.3995 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.79 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 08:20 | 0 | 85 | 85 | 0 | 39.86 | 40 | 104.2985 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 32.77 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.022 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5035
2008-08-05 08:21 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 104.3995 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.73 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6584 | 9.5035
2008-08-05 08:22 | 0 | 85 | 85 | 0 | 40.04 | 40 | 104.3995 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.69 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 08:23 | 0 | 85.02 | 85 | 0 | 40.08 | 40 | 104.349 | 0 | 23.5242 | 23.1391 | 0 |  | 0.97627 | 23.69 | 23.5165 | 0 | 32.62 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5035
2008-08-05 08:24 | 0 | 84.99 | 85 | 0 | 39.84 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 32.64 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 08:25 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 104.3995 | 0 | 23.5356 | 23.1391 | 0 |  | 0.97674 | 23.69 | 23.5165 | 0 | 32.67 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 08:26 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.3995 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.74 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:27 | 0 | 84.92 | 85 | 0 | 39.84 | 40 | 104.3995 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.83 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 08:28 | 0 | 84.97 | 85 | 0 | 39.84 | 40 | 104.349 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.84 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 08:29 | 0 | 85 | 85 | 0 | 39.96 | 40 | 104.349 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.81 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 08:30 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 104.349 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.82 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 08:31 | 0 | 84.94 | 85 | 0 | 39.9 | 40 | 104.3995 | 0 | 23.5348 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.78 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:32 | 0 | 84.91 | 85 | 0 | 39.94 | 40 | 104.3995 | 0 | 23.5356 | 23.1279 | 0 |  | 0.97674 | 23.69 | 23.5287 | 0 | 32.74 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:33 | 0 | 85 | 85 | 0 | 39.87 | 40 | 104.3995 | 0 | 23.5348 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5165 | 0 | 32.68 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.7076 | 9.5035
2008-08-05 08:34 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 104.349 | 0 | 23.5234 | 23.1382 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.64 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 08:35 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 104.3995 | 0 | 23.5576 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.65 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 08:36 | 0 | 85 | 85 | 0 | 39.89 | 40 | 104.349 | 0 | 23.5348 | 23.1606 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.73 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 08:37 | 0 | 84.97 | 85 | 0 | 40.13 | 40 | 104.349 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.81 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:38 | 0 | 84.95 | 85 | 0 | 39.84 | 40 | 104.3995 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.91 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 08:39 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 104.349 | 0 | 23.5121 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5287 | 0 | 32.91 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 08:40 | 0 | 85.05 | 85 | 0 | 39.94 | 40 | 104.3995 | 0 | 23.5121 | 23.1382 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.9 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 08:41 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 104.349 | 0 | 23.5129 | 23.1391 | 0 |  | 0.97627 | 23.69 | 23.5043 | 0 | 32.85 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 08:42 | 0 | 85 | 85 | 0 | 39.95 | 40 | 104.3995 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 32.83 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 08:43 | 0 | 85.02 | 85 | 0 | 39.91 | 40 | 104.349 | 0 | 23.5129 | 23.1167 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 32.79 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:44 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 104.349 | 0 | 23.5348 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 32.71 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5035
2008-08-05 08:45 | 0 | 85.05 | 85 | 0 | 39.95 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 32.71 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 08:46 | 0 | 84.97 | 85 | 0 | 40.01 | 40 | 104.349 | 0 | 23.5129 | 23.1279 | 0 |  | 0.97627 | 23.69 | 23.5043 | 0 | 32.64 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6699 | 9.5035
2008-08-05 08:47 | 0 | 85.05 | 85 | 0 | 40.03 | 40 | 104.2985 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97576 | 23.69 | 23.5043 | 0 | 32.64 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.7068 | 9.5035
2008-08-05 08:48 | 0 | 84.94 | 85 | 0 | 39.94 | 40 | 104.349 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.65 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 08:49 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 104.2985 | 0 | 23.5121 | 23.1159 | 0 |  | 0.97576 | 23.69 | 23.4921 | 0 | 32.65 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 08:50 | 0 | 85.05 | 85 | 0 | 39.93 | 40 | 104.3995 | 0 | 23.5129 | 23.1167 | 0 |  | 0.97674 | 23.69 | 23.5043 | 0 | 32.71 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6813 | 9.5035
2008-08-05 08:51 | 0 | 84.96 | 85 | 0 | 39.9 | 40 | 104.2985 | 0 | 23.5356 | 23.1279 | 0 |  | 0.9758 | 23.69 | 23.5287 | 0 | 32.77 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6945 | 9.5035
2008-08-05 08:52 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5043 | 0 | 32.79 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:53 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5409 | 0 | 32.76 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 08:54 | 0 | 84.97 | 85 | 0 | 39.84 | 40 | 104.3995 | 0 | 23.5348 | 23.1271 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.75 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:55 | 0 | 84.97 | 85 | 0 | 40 | 40 | 104.3995 | 0 | 23.5356 | 23.1279 | 0 |  | 0.97671 | 23.69 | 23.5287 | 0 | 32.71 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 08:56 | 0 | 85 | 85 | 0 | 40 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.68 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 08:57 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 104.3995 | 0 | 23.5234 | 23.1159 | 0 |  | 0.97671 | 23.69 | 23.5165 | 0 | 32.61 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:58 | 0 | 84.91 | 85 | 0 | 39.91 | 40 | 104.349 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.58 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 08:59 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 104.349 | 0 | 23.5234 | 23.1382 | 0 |  | 0.97624 | 23.69 | 23.5165 | 0 | 32.62 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.683 | 9.5035
2008-08-05 09:00 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 104.3995 | 0 | 23.5348 | 23.1279 | 0 |  | 0.97674 | 23.69 | 23.5165 | 0 | 32.65 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6715 | 9.5035
2008-08-05 09:01 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 104.3995 | 0 | 23.5234 | 23.1271 | 0 |  | 0.97674 | 23.69 | 23.5287 | 0 | 32.72 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 09:02 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 104.349 | 0 | 23.5356 | 23.1279 | 0 |  | 0.97627 | 23.69 | 23.5165 | 0 | 32.71 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6822 | 9.5035
2008-08-05 09:03 | 0 | 84.97 | 85 | 0 | 40 | 40 | 104.349 | 0 | 23.5348 | 23.1271 | 0 |  | 0.97627 | 23.69 | 23.5287 | 0 | 32.72 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5035
2008-08-05 09:04 | 0 | 84.95 | 85 | 0 | 39.91 | 40 | 104.3995 | 0 | 23.5356 | 23.1167 | 0 |  | 0.97674 | 23.69 | 23.5165 | 0 | 32.71 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6838 | 9.5035
2008-08-05 09:05 | 0 | 85 | 85 | 0 | 40.07 | 40 | 104.3995 | 0 | 23.5121 | 23.1382 | 0 |  | 0.97674 | 23.69 | 23.5165 | 0 | 32.64 | 10.1969 | 0 | 0.42575 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 20.0445 | -1 | 20.0237 | 23.6953 | 9.5035
2008-08-05 09:06 | 0 | 85 | 85 | 0 | 40.08 | 40 | 104.2985 | 0 | 23.5348 | 30.2878 | 0 |  | 0.97576 | 31.01 | 23.5043 | 0 | 32.6 | 20.3986 | 0 | 0.8517 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0445 | -1 | 20.031 | 23.6838 | 19.0115
2008-08-05 09:07 | 0 | 85.05 | 85 | 0 | 39.97 | 40 | 99.8031 | 0 | 28.9858 | 28.5549 | 0 |  | 0.92077 | 31.01 | 28.9621 | 0 | 31.88 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 19.05 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.05 | -1 | 20.031 | 31.0085 | 0
2008-08-05 09:08 | 0 | 85.07 | 85 | 0 | 40.03 | 40 | 99.8536 | 0 | 28.9407 | 28.5108 | 0 |  | 0.91938 | 31.01 | 28.9133 | 0 | 31.21 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 19.64 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0499 | -1 | 20.031 | 30.9559 | 0
2008-08-05 09:09 | 0 | 84.94 | 85 | 0 | 39.95 | 40 | 99.7526 | 0 | 28.8923 | 28.4789 | 0 |  | 0.91791 | 31.01 | 28.8767 | 0 | 30.83 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 19.81 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0499 | -1 | 20.031 | 31.0343 | 0
2008-08-05 09:10 | 0 | 85.05 | 85 | 0 | 39.94 | 40 | 99.8031 | 0 | 28.8835 | 28.4557 | 0 |  | 0.9181 | 31.01 | 28.8645 | 0 | 30.68 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 19.9 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9893 | 0
2008-08-05 09:11 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 99.7526 | 0 | 28.8796 | 28.4375 | 0 |  | 0.91751 | 31.01 | 28.8523 | 0 | 30.62 | 0 | 0 | 0 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 19.94 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 31.0263 | 0
2008-08-05 09:12 | 0 | 84.99 | 85 | 0 | 40 | 40 | 99.7526 | 0 | 28.8632 | 28.4489 | 0 |  | 0.91741 | 31.01 | 28.84 | 0 | 30.61 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 19.97 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9834 | 0
2008-08-05 09:13 | 0 | 84.96 | 85 | 0 | 40 | 40 | 99.7021 | 0 | 28.8748 | 28.4326 | 0 |  | 0.91688 | 31.01 | 28.84 | 0 | 30.62 | 0 | 0 | 0 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9712 | 0
2008-08-05 09:14 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 99.7526 | 0 | 28.8603 | 28.447 | 0 |  | 0.91732 | 31.01 | 28.84 | 0 | 30.66 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9746 | 0
2008-08-05 09:15 | 0 | 85 | 85 | 0 | 39.97 | 40 | 99.6011 | 0 | 28.8593 | 28.4307 | 0 |  | 0.91589 | 31.01 | 28.8278 | 0 | 30.62 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.989 | 0
2008-08-05 09:16 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 99.7021 | 0 | 28.8584 | 28.4451 | 0 |  | 0.91682 | 31.01 | 28.8523 | 0 | 30.62 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 31.0033 | 0
2008-08-05 09:17 | 0 | 84.92 | 85 | 0 | 40.04 | 40 | 99.7021 | 0 | 28.8584 | 28.4441 | 0 |  | 0.91679 | 31.01 | 28.8523 | 0 | 30.69 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9768 | 0
2008-08-05 09:18 | 0 | 85 | 85 | 0 | 40.03 | 40 | 99.7526 | 0 | 28.8872 | 28.4307 | 0 |  | 0.91729 | 31.01 | 28.84 | 0 | 30.75 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0499 | -1 | 20.031 | 30.9454 | 0
2008-08-05 09:19 | 0 | 84.95 | 85 | 0 | 40.01 | 40 | 99.8031 | 0 | 28.8718 | 28.4297 | 0 |  | 0.91772 | 31.01 | 28.84 | 0 | 30.8 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 31.0092 | 0
2008-08-05 09:20 | 0 | 84.91 | 85 | 0 | 40.04 | 40 | 99.7021 | 0 | 28.8574 | 28.4441 | 0 |  | 0.91679 | 31.01 | 28.8278 | 0 | 30.87 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 31.0009 | 0
2008-08-05 09:21 | 0 | 84.99 | 85 | 0 | 40 | 40 | 99.7526 | 0 | 28.8574 | 28.4297 | 0 |  | 0.91726 | 31.01 | 28.8278 | 0 | 30.94 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9778 | 0
2008-08-05 09:22 | 0 | 84.92 | 85 | 0 | 40.11 | 40 | 99.8031 | 0 | 28.8718 | 28.4441 | 0 |  | 0.91772 | 31.01 | 28.8523 | 0 | 30.96 | 0 | 0 | 0 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9621 | 0
2008-08-05 09:23 | 0 | 85 | 85 | 0 | 39.96 | 40 | 99.7021 | 0 | 28.8863 | 28.4297 | 0 |  | 0.91679 | 31.01 | 28.8523 | 0 | 30.91 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0499 | -1 | 20.031 | 31.0044 | 0
2008-08-05 09:24 | 0 | 84.99 | 85 | 0 | 40 | 40 | 99.7021 | 0 | 28.8574 | 28.4451 | 0 |  | 0.91682 | 31.01 | 28.8278 | 0 | 30.91 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9887 | 0
2008-08-05 09:25 | 0 | 84.96 | 85 | 0 | 39.92 | 40 | 99.7021 | 0 | 28.8574 | 28.4441 | 0 |  | 0.91679 | 31.01 | 28.8523 | 0 | 30.88 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 31.0033 | 0
2008-08-05 09:26 | 0 | 84.94 | 85 | 0 | 39.91 | 40 | 99.7021 | 0 | 28.8728 | 28.4307 | 0 |  | 0.91682 | 31.01 | 28.84 | 0 | 30.85 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9744 | 0
2008-08-05 09:27 | 0 | 85.03 | 85 | 0 | 40 | 40 | 99.7526 | 0 | 28.8728 | 28.4441 | 0 |  | 0.91726 | 31.01 | 28.8523 | 0 | 30.83 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9611 | 0
2008-08-05 09:28 | 0 | 84.92 | 85 | 0 | 39.95 | 40 | 99.7526 | 0 | 28.8584 | 28.4307 | 0 |  | 0.91729 | 31.01 | 28.84 | 0 | 30.81 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9587 | 0
2008-08-05 09:29 | 0 | 85 | 85 | 0 | 39.95 | 40 | 99.7021 | 0 | 28.8584 | 28.4307 | 0 |  | 0.91682 | 31.01 | 28.84 | 0 | 30.87 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.99 | 0
2008-08-05 09:30 | 0 | 85.03 | 85 | 0 | 39.96 | 40 | 99.6011 | 0 | 28.8728 | 28.4019 | 0 |  | 0.91589 | 31.01 | 28.8278 | 0 | 31.02 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9478 | 0
2008-08-05 09:31 | 0 | 84.94 | 85 | 0 | 39.95 | 40 | 99.7021 | 0 | 28.8584 | 28.4451 | 0 |  | 0.91682 | 31.01 | 28.8278 | 0 | 31.15 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.99 | 0
2008-08-05 09:32 | 0 | 84.94 | 85 | 0 | 39.93 | 40 | 99.7021 | 0 | 28.8728 | 28.446 | 0 |  | 0.91685 | 31.01 | 28.8523 | 0 | 31.2 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9744 | 0
2008-08-05 09:33 | 0 | 85 | 85 | 0 | 40.04 | 40 | 99.7526 | 0 | 28.8584 | 28.4451 | 0 |  | 0.91729 | 31.01 | 28.8523 | 0 | 31.17 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9876 | 0
2008-08-05 09:34 | 0 | 84.97 | 85 | 0 | 39.84 | 40 | 99.7526 | 0 | 28.8584 | 28.4307 | 0 |  | 0.91729 | 31.01 | 28.8523 | 0 | 31.15 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 31.0033 | 0
2008-08-05 09:35 | 0 | 84.95 | 85 | 0 | 39.88 | 40 | 99.7021 | 0 | 28.8728 | 28.4451 | 0 |  | 0.91682 | 31.01 | 28.8278 | 0 | 31.11 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9744 | 0
2008-08-05 09:36 | 0 | 84.97 | 85 | 0 | 40.01 | 40 | 99.7021 | 0 | 28.8584 | 28.4307 | 0 |  | 0.91682 | 31.01 | 28.8034 | 0 | 31.09 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 31.0033 | 0
2008-08-05 09:37 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 99.7021 | 0 | 28.8584 | 28.4451 | 0 |  | 0.91682 | 31.01 | 28.8278 | 0 | 31.02 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.99 | 0
2008-08-05 09:38 | 0 | 85.05 | 85 | 0 | 40.03 | 40 | 99.7526 | 0 | 28.8882 | 28.4316 | 0 |  | 0.91732 | 31.01 | 28.8156 | 0 | 31.01 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0499 | -1 | 20.031 | 30.9709 | 0
2008-08-05 09:39 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 99.7021 | 0 | 28.8593 | 28.4316 | 0 |  | 0.91685 | 31.01 | 28.84 | 0 | 31.03 | 0 | 0 | 0 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9733 | 0
2008-08-05 09:40 | 0 | 84.94 | 85 | 0 | 39.93 | 40 | 99.7021 | 0 | 28.8593 | 28.4316 | 0 |  | 0.91685 | 31.01 | 28.84 | 0 | 31.11 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9866 | 0
2008-08-05 09:41 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 99.6011 | 0 | 28.8593 | 28.4316 | 0 |  | 0.91592 | 31.01 | 28.8278 | 0 | 31.25 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.989 | 0
2008-08-05 09:42 | 0 | 85.03 | 85 | 0 | 39.98 | 40 | 99.7526 | 0 | 28.8738 | 28.4451 | 0 |  | 0.91732 | 31.01 | 28.8523 | 0 | 31.26 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9733 | 0
2008-08-05 09:43 | 0 | 84.94 | 85 | 0 | 39.95 | 40 | 99.7021 | 0 | 28.8593 | 28.446 | 0 |  | 0.91685 | 31.01 | 28.84 | 0 | 31.28 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9733 | 0
2008-08-05 09:44 | 0 | 85.07 | 85 | 0 | 40 | 40 | 99.7021 | 0 | 28.8738 | 28.4316 | 0 |  | 0.91685 | 31.01 | 28.8523 | 0 | 31.25 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9866 | 0
2008-08-05 09:45 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 99.7526 | 0 | 28.8593 | 28.4316 | 0 |  | 0.91732 | 31.01 | 28.8523 | 0 | 31.2 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9733 | 0
2008-08-05 09:46 | 0 | 84.94 | 85 | 0 | 40 | 40 | 99.7021 | 0 | 28.8593 | 28.4316 | 0 |  | 0.91685 | 31.01 | 28.84 | 0 | 31.16 | 0 | 0 | 0 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.989 | 0
2008-08-05 09:47 | 0 | 85 | 85 | 0 | 39.97 | 40 | 99.7021 | 0 | 28.8593 | 28.4316 | 0 |  | 0.91685 | 31.01 | 28.84 | 0 | 31.11 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9866 | 0
2008-08-05 09:48 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 99.7526 | 0 | 28.8738 | 28.4316 | 0 |  | 0.91732 | 31.01 | 28.8523 | 0 | 31.07 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.989 | 0
2008-08-05 09:49 | 0 | 85 | 85 | 0 | 39.9 | 40 | 99.7021 | 0 | 28.8882 | 28.446 | 0 |  | 0.91685 | 31.01 | 28.84 | 0 | 31.05 | 0 | 0 | 0 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0499 | -1 | 20.031 | 30.9866 | 0
2008-08-05 09:50 | 0 | 84.94 | 85 | 0 | 40.03 | 40 | 99.7021 | 0 | 28.8738 | 28.446 | 0 |  | 0.91685 | 31.01 | 28.84 | 0 | 31.01 | 0 | 0 | 0 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9866 | 0
2008-08-05 09:51 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 99.7021 | 0 | 28.8738 | 28.4316 | 0 |  | 0.91685 | 31.01 | 28.8523 | 0 | 30.96 | 0 | 0 | 0 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0498 | -1 | 20.031 | 30.9866 | 0
2008-08-05 09:52 | 0 | 84.95 | 85 | 0 | 40.05 | 40 | 104.5006 | 0 | 30.2504 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.2076 | 0 | 30.99 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 09:53 | 0 | 85 | 85 | 0 | 39.94 | 40 | 104.5511 | 0 | 30.2504 | 29.8 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.1 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 09:54 | 0 | 84.95 | 85 | 0 | 40 | 40 | 104.5006 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.21 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 09:55 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 104.5511 | 0 | 30.2504 | 29.8 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.25 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 09:56 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.24 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 09:57 | 0 | 84.92 | 85 | 0 | 39.83 | 40 | 104.5006 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.22 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9815 | 19.0119
2008-08-05 09:58 | 0 | 85.02 | 85 | 0 | 39.93 | 40 | 104.5006 | 0 | 30.2504 | 29.8288 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.21 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 09:59 | 0 | 85.03 | 85 | 0 | 39.98 | 40 | 104.5006 | 0 | 30.2504 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.1954 | 0 | 31.2 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:00 | 0 | 85.02 | 85 | 0 | 40.04 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.2 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:01 | 0 | 84.97 | 85 | 0 | 39.84 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.1954 | 0 | 31.15 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:02 | 0 | 84.92 | 85 | 0 | 39.9 | 40 | 104.5006 | 0 | 30.2504 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.2198 | 0 | 31.21 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:03 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.29 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.0119
2008-08-05 10:04 | 0 | 85 | 85 | 0 | 39.9 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.43 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 10:05 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 104.6016 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.2076 | 0 | 31.52 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 10:06 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 104.5511 | 0 | 30.2494 | 29.8134 | 0 |  | 0.96141 | 31.01 | 30.2198 | 0 | 31.58 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9676 | 19.0119
2008-08-05 10:07 | 0 | 85.05 | 85 | 0 | 40 | 40 | 104.6016 | 0 | 30.2789 | 29.8134 | 0 |  | 0.96188 | 31.01 | 30.2442 | 0 | 31.57 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:08 | 0 | 85 | 85 | 0 | 39.91 | 40 | 104.6016 | 0 | 30.2494 | 29.8278 | 0 |  | 0.96188 | 31.01 | 30.2198 | 0 | 31.53 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.965 | 19.0119
2008-08-05 10:09 | 0 | 85.03 | 85 | 0 | 39.95 | 40 | 104.5511 | 0 | 30.2494 | 29.8278 | 0 |  | 0.96141 | 31.01 | 30.2442 | 0 | 31.49 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.98 | 19.0119
2008-08-05 10:10 | 0 | 84.96 | 85 | 0 | 39.91 | 40 | 104.5511 | 0 | 30.2494 | 29.8278 | 0 |  | 0.96141 | 31.01 | 30.232 | 0 | 31.46 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.98 | 19.0119
2008-08-05 10:11 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 104.5511 | 0 | 30.2641 | 29.8278 | 0 |  | 0.96141 | 31.01 | 30.2198 | 0 | 31.44 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.965 | 19.0119
2008-08-05 10:12 | 0 | 85.02 | 85 | 0 | 40 | 40 | 104.6016 | 0 | 30.2346 | 29.8134 | 0 |  | 0.96188 | 31.01 | 30.2198 | 0 | 31.44 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9924 | 19.0119
2008-08-05 10:13 | 0 | 85 | 85 | 0 | 39.94 | 40 | 104.6016 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.51 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 10:14 | 0 | 85 | 85 | 0 | 40.05 | 40 | 104.6016 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.61 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9887 | 19.0119
2008-08-05 10:15 | 0 | 85.03 | 85 | 0 | 40.01 | 40 | 104.5511 | 0 | 30.2494 | 29.8134 | 0 |  | 0.96141 | 31.01 | 30.2198 | 0 | 31.65 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.0119
2008-08-05 10:16 | 0 | 85.05 | 85 | 0 | 40.03 | 40 | 104.5511 | 0 | 30.2652 | 29.8278 | 0 |  | 0.96141 | 31.01 | 30.232 | 0 | 31.69 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9526 | 19.0119
2008-08-05 10:17 | 0 | 84.94 | 85 | 0 | 40.07 | 40 | 104.6016 | 0 | 30.2504 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.65 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9789 | 19.0119
2008-08-05 10:18 | 0 | 84.92 | 85 | 0 | 39.9 | 40 | 104.5511 | 0 | 30.2346 | 29.8288 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.66 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9949 | 19.0119
2008-08-05 10:19 | 0 | 85.02 | 85 | 0 | 39.95 | 40 | 104.6016 | 0 | 30.2494 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.2442 | 0 | 31.61 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:20 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 104.6016 | 0 | 30.2504 | 29.799 | 0 |  | 0.96188 | 31.01 | 30.232 | 0 | 31.58 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9898 | 19.0119
2008-08-05 10:21 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.6016 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.2442 | 0 | 31.55 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.98 | 19.0119
2008-08-05 10:22 | 0 | 85 | 85 | 0 | 40 | 40 | 104.5006 | 0 | 30.2357 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.53 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 10:23 | 0 | 84.94 | 85 | 0 | 40 | 40 | 104.5006 | 0 | 30.2494 | 29.799 | 0 |  | 0.96095 | 31.01 | 30.232 | 0 | 31.52 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 10:24 | 0 | 85.05 | 85 | 0 | 40.1 | 40 | 104.5511 | 0 | 30.2357 | 29.8 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.52 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:25 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.59 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9887 | 19.0119
2008-08-05 10:26 | 0 | 84.96 | 85 | 0 | 40.01 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.63 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.0119
2008-08-05 10:27 | 0 | 85.03 | 85 | 0 | 40.01 | 40 | 104.5006 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.65 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 10:28 | 0 | 84.92 | 85 | 0 | 39.93 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.61 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:29 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 104.6016 | 0 | 30.2504 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.58 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 10:30 | 0 | 84.96 | 85 | 0 | 40 | 40 | 104.5511 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.57 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 10:31 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.52 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0187 | 19.0119
2008-08-05 10:32 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 104.6016 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.2442 | 0 | 31.48 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 10:33 | 0 | 84.94 | 85 | 0 | 39.86 | 40 | 104.6016 | 0 | 30.2504 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.43 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 10:34 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.42 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:35 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 104.6016 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.2076 | 0 | 31.42 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:36 | 0 | 84.88 | 85 | 0 | 39.83 | 40 | 104.6016 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.49 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:37 | 0 | 85.02 | 85 | 0 | 39.98 | 40 | 104.6016 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.2442 | 0 | 31.59 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9887 | 19.0119
2008-08-05 10:38 | 0 | 84.95 | 85 | 0 | 40 | 40 | 104.6016 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.2442 | 0 | 31.66 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 10:39 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 104.5511 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.63 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:40 | 0 | 85 | 85 | 0 | 39.92 | 40 | 104.6016 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.2442 | 0 | 31.61 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:41 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 104.5006 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.57 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 10:42 | 0 | 85 | 85 | 0 | 39.96 | 40 | 104.6016 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.52 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 10:43 | 0 | 85.03 | 85 | 0 | 39.93 | 40 | 104.5511 | 0 | 30.2357 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.5 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:44 | 0 | 84.95 | 85 | 0 | 39.91 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.5 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:45 | 0 | 84.92 | 85 | 0 | 39.93 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.44 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 10:46 | 0 | 84.97 | 85 | 0 | 40.01 | 40 | 104.5511 | 0 | 30.2515 | 29.8298 | 0 |  | 0.96148 | 31.01 | 30.2442 | 0 | 31.4 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9902 | 19.0119
2008-08-05 10:47 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.39 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9902 | 19.0119
2008-08-05 10:48 | 0 | 85.03 | 85 | 0 | 40 | 40 | 104.6016 | 0 | 30.2515 | 29.8298 | 0 |  | 0.96194 | 31.01 | 30.232 | 0 | 31.4 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9887 | 19.0119
2008-08-05 10:49 | 0 | 84.92 | 85 | 0 | 40.03 | 40 | 104.6016 | 0 | 30.2662 | 29.8154 | 0 |  | 0.96194 | 31.01 | 30.2442 | 0 | 31.42 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 10:50 | 0 | 85 | 85 | 0 | 40.05 | 40 | 104.5511 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.47 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0011 | 19.0119
2008-08-05 10:51 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 104.5511 | 0 | 30.2652 | 29.8298 | 0 |  | 0.96148 | 31.01 | 30.232 | 0 | 31.52 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 10:52 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 104.5511 | 0 | 30.2357 | 29.8 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.51 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9541 | 19.0119
2008-08-05 10:53 | 0 | 85.05 | 85 | 0 | 40.03 | 40 | 104.5006 | 0 | 30.2357 | 29.8154 | 0 |  | 0.96101 | 31.01 | 30.232 | 0 | 31.49 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9541 | 19.0119
2008-08-05 10:54 | 0 | 84.91 | 85 | 0 | 39.98 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.48 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9541 | 19.0119
2008-08-05 10:55 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 104.5511 | 0 | 30.2515 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.1954 | 0 | 31.46 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 10:56 | 0 | 84.97 | 85 | 0 | 40.05 | 40 | 104.5006 | 0 | 30.2367 | 29.801 | 0 |  | 0.96101 | 31.01 | 30.1954 | 0 | 31.36 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9939 | 19.0119
2008-08-05 10:57 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 104.5511 | 0 | 30.2515 | 29.8 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.35 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0052 | 19.0119
2008-08-05 10:58 | 0 | 85 | 85 | 0 | 39.96 | 40 | 104.5006 | 0 | 30.2367 | 29.8154 | 0 |  | 0.96101 | 31.01 | 30.232 | 0 | 31.33 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0052 | 19.0119
2008-08-05 10:59 | 0 | 85 | 85 | 0 | 39.9 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.32 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 11:00 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 104.5511 | 0 | 30.2367 | 29.801 | 0 |  | 0.96148 | 31.01 | 30.2076 | 0 | 31.31 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9928 | 19.0119
2008-08-05 11:01 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.29 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9753 | 19.0119
2008-08-05 11:02 | 0 | 85.05 | 85 | 0 | 39.93 | 40 | 104.5511 | 0 | 30.2662 | 29.8154 | 0 |  | 0.96148 | 31.01 | 30.2076 | 0 | 31.36 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9877 | 19.0119
2008-08-05 11:03 | 0 | 85 | 85 | 0 | 39.94 | 40 | 104.6016 | 0 | 30.2357 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.2076 | 0 | 31.47 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9753 | 19.0119
2008-08-05 11:04 | 0 | 85 | 85 | 0 | 39.92 | 40 | 104.5511 | 0 | 30.2515 | 29.8288 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.52 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9902 | 19.0119
2008-08-05 11:05 | 0 | 84.92 | 85 | 0 | 39.84 | 40 | 104.5511 | 0 | 30.2515 | 29.8298 | 0 |  | 0.96148 | 31.01 | 30.232 | 0 | 31.52 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9902 | 19.0119
2008-08-05 11:06 | 0 | 85 | 85 | 0 | 40.08 | 40 | 104.6016 | 0 | 30.2367 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.47 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0052 | 19.0119
2008-08-05 11:07 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 104.6016 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.2442 | 0 | 31.45 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9887 | 19.0119
2008-08-05 11:08 | 0 | 85 | 85 | 0 | 39.95 | 40 | 104.5511 | 0 | 30.2504 | 29.8288 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.39 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 11:09 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.6016 | 0 | 30.2504 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.35 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:10 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 104.6016 | 0 | 30.2515 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.36 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9902 | 19.0119
2008-08-05 11:11 | 0 | 84.97 | 85 | 0 | 40.01 | 40 | 104.5511 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.32 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9753 | 19.0119
2008-08-05 11:12 | 0 | 85 | 85 | 0 | 40.01 | 40 | 104.6016 | 0 | 30.2504 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.33 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:13 | 0 | 85 | 85 | 0 | 39.94 | 40 | 104.5511 | 0 | 30.2662 | 29.8154 | 0 |  | 0.96148 | 31.01 | 30.232 | 0 | 31.4 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9753 | 19.0119
2008-08-05 11:14 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 104.5511 | 0 | 30.2515 | 29.8288 | 0 |  | 0.96148 | 31.01 | 30.232 | 0 | 31.47 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9902 | 19.0119
2008-08-05 11:15 | 0 | 84.97 | 85 | 0 | 40 | 40 | 104.5511 | 0 | 30.2515 | 29.8288 | 0 |  | 0.96148 | 31.01 | 30.232 | 0 | 31.53 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:16 | 0 | 85 | 85 | 0 | 39.93 | 40 | 104.5006 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.51 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0052 | 19.0119
2008-08-05 11:17 | 0 | 84.96 | 85 | 0 | 40.05 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.49 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:18 | 0 | 84.92 | 85 | 0 | 39.96 | 40 | 104.5006 | 0 | 30.2504 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.2076 | 0 | 31.44 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:19 | 0 | 84.94 | 85 | 0 | 39.97 | 40 | 104.6016 | 0 | 30.2504 | 29.8 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.42 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:20 | 0 | 85 | 85 | 0 | 40 | 40 | 104.5511 | 0 | 30.2357 | 29.8288 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.39 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:21 | 0 | 84.94 | 85 | 0 | 40.01 | 40 | 104.5511 | 0 | 30.2504 | 29.8 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.36 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9516 | 19.0119
2008-08-05 11:22 | 0 | 85.02 | 85 | 0 | 39.92 | 40 | 104.5006 | 0 | 30.2357 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.2076 | 0 | 31.32 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9691 | 19.0119
2008-08-05 11:23 | 0 | 85.03 | 85 | 0 | 40.03 | 40 | 104.5006 | 0 | 30.2504 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.31 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9789 | 19.0119
2008-08-05 11:24 | 0 | 84.96 | 85 | 0 | 40 | 40 | 104.5006 | 0 | 30.2357 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.36 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9541 | 19.0119
2008-08-05 11:25 | 0 | 85.02 | 85 | 0 | 39.95 | 40 | 104.5511 | 0 | 30.2357 | 29.8 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.43 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.968 | 19.0119
2008-08-05 11:26 | 0 | 84.96 | 85 | 0 | 40.07 | 40 | 104.5006 | 0 | 30.2367 | 29.8154 | 0 |  | 0.96101 | 31.01 | 30.2076 | 0 | 31.48 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9804 | 19.0119
2008-08-05 11:27 | 0 | 84.95 | 85 | 0 | 39.89 | 40 | 104.5511 | 0 | 30.2504 | 29.8 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.47 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.0119
2008-08-05 11:28 | 0 | 85 | 85 | 0 | 40.04 | 40 | 104.5511 | 0 | 30.2357 | 29.8 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.43 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9866 | 19.0119
2008-08-05 11:29 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 104.5511 | 0 | 30.2357 | 29.8 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.42 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9789 | 19.0119
2008-08-05 11:30 | 0 | 85 | 85 | 0 | 40 | 40 | 104.3995 | 0 | 30.2504 | 29.7712 | 0 |  | 0.96005 | 31.01 | 30.2076 | 0 | 31.38 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.012
2008-08-05 11:31 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 104.5006 | 0 | 30.2357 | 29.7712 | 0 |  | 0.96098 | 31.01 | 30.2076 | 0 | 31.39 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9691 | 19.012
2008-08-05 11:32 | 0 | 84.97 | 85 | 0 | 40.08 | 40 | 104.5006 | 0 | 30.2357 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.1954 | 0 | 31.35 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0115 | 19.0119
2008-08-05 11:33 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 104.3995 | 0 | 30.2357 | 29.8 | 0 |  | 0.96005 | 31.01 | 30.1954 | 0 | 31.35 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 11:34 | 0 | 85.02 | 85 | 0 | 40.07 | 40 | 104.5006 | 0 | 30.2357 | 29.7712 | 0 |  | 0.96098 | 31.01 | 30.2076 | 0 | 31.35 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.012
2008-08-05 11:35 | 0 | 85 | 85 | 0 | 39.97 | 40 | 104.5006 | 0 | 30.2357 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.1954 | 0 | 31.46 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9541 | 19.0119
2008-08-05 11:36 | 0 | 84.95 | 85 | 0 | 39.84 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.52 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9541 | 19.0119
2008-08-05 11:37 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 104.5006 | 0 | 30.2357 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.55 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.0119
2008-08-05 11:38 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 104.5006 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.53 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:39 | 0 | 85 | 85 | 0 | 39.98 | 40 | 104.5006 | 0 | 30.2504 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.49 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.0119
2008-08-05 11:40 | 0 | 84.99 | 85 | 0 | 40.07 | 40 | 104.6016 | 0 | 30.2357 | 29.8 | 0 |  | 0.96191 | 31.01 | 30.2198 | 0 | 31.45 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:41 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 104.5511 | 0 | 30.2357 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.4 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:42 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 104.5511 | 0 | 30.2357 | 29.8288 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.43 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9815 | 19.0119
2008-08-05 11:43 | 0 | 84.94 | 85 | 0 | 39.96 | 40 | 104.5006 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.5 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 11:44 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.61 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:45 | 0 | 85.03 | 85 | 0 | 39.95 | 40 | 104.5511 | 0 | 30.2357 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.69 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 11:46 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 104.6016 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.74 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:47 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.5006 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.7 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 11:48 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 104.6016 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.2442 | 0 | 31.69 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 11:49 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 104.6016 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.2198 | 0 | 31.63 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 11:50 | 0 | 84.96 | 85 | 0 | 40 | 40 | 104.5006 | 0 | 30.2357 | 29.8288 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.62 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 11:51 | 0 | 84.95 | 85 | 0 | 40 | 40 | 104.5511 | 0 | 30.2357 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.58 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:52 | 0 | 85 | 85 | 0 | 40 | 40 | 104.5006 | 0 | 30.2357 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.1954 | 0 | 31.65 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9789 | 19.0119
2008-08-05 11:53 | 0 | 84.89 | 85 | 0 | 39.97 | 40 | 104.5006 | 0 | 30.2357 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.73 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 11:54 | 0 | 84.95 | 85 | 0 | 40 | 40 | 104.5006 | 0 | 30.2504 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.8 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.0119
2008-08-05 11:55 | 0 | 85 | 85 | 0 | 40.11 | 40 | 104.5006 | 0 | 30.2357 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.2076 | 0 | 31.8 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9939 | 19.0119
2008-08-05 11:56 | 0 | 84.92 | 85 | 0 | 40.15 | 40 | 104.5006 | 0 | 30.2504 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.78 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9815 | 19.0119
2008-08-05 11:57 | 0 | 84.95 | 85 | 0 | 40.07 | 40 | 104.5006 | 0 | 30.2357 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.2076 | 0 | 31.76 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 11:58 | 0 | 84.92 | 85 | 0 | 40 | 40 | 104.5006 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.73 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.0119
2008-08-05 11:59 | 0 | 84.91 | 85 | 0 | 40 | 40 | 104.3995 | 0 | 30.2357 | 29.8 | 0 |  | 0.96005 | 31.01 | 30.2076 | 0 | 31.69 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0363 | 19.0119
2008-08-05 12:00 | 0 | 85 | 85 | 0 | 39.88 | 40 | 104.3995 | 0 | 30.2357 | 29.8 | 0 |  | 0.96005 | 31.01 | 30.1954 | 0 | 31.65 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.0119
2008-08-05 12:01 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 104.5006 | 0 | 30.2357 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.2076 | 0 | 31.62 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:02 | 0 | 85.03 | 85 | 0 | 39.91 | 40 | 104.5006 | 0 | 30.2504 | 29.8 | 0 |  | 0.96098 | 31.01 | 30.1832 | 0 | 31.58 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:03 | 0 | 85 | 85 | 0 | 39.95 | 40 | 104.5006 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.2076 | 0 | 31.54 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.0119
2008-08-05 12:04 | 0 | 84.94 | 85 | 0 | 39.96 | 40 | 104.5006 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.53 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:05 | 0 | 85 | 85 | 0 | 40.03 | 40 | 104.5006 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.2076 | 0 | 31.61 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9815 | 19.0119
2008-08-05 12:06 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 104.5006 | 0 | 30.2357 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.72 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 12:07 | 0 | 84.92 | 85 | 0 | 39.95 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.81 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 12:08 | 0 | 85 | 85 | 0 | 40 | 40 | 104.5511 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.82 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 12:09 | 0 | 84.94 | 85 | 0 | 40 | 40 | 104.6016 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.2076 | 0 | 31.78 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:10 | 0 | 84.95 | 85 | 0 | 40 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.74 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:11 | 0 | 85.05 | 85 | 0 | 39.89 | 40 | 104.5006 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.2076 | 0 | 31.7 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 12:12 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 104.5511 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.69 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 12:13 | 0 | 84.92 | 85 | 0 | 40.03 | 40 | 104.6016 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.2076 | 0 | 31.63 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:14 | 0 | 85.03 | 85 | 0 | 39.98 | 40 | 104.5511 | 0 | 30.2504 | 29.8288 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.72 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:15 | 0 | 85.08 | 85 | 0 | 40.03 | 40 | 104.6016 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.84 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:16 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 104.6016 | 0 | 30.2504 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.89 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 12:17 | 0 | 84.97 | 85 | 0 | 40 | 40 | 104.5511 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.89 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 12:18 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 104.6521 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96237 | 31.01 | 30.2442 | 0 | 31.87 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9887 | 19.0119
2008-08-05 12:19 | 0 | 84.92 | 85 | 0 | 40.03 | 40 | 104.6016 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.82 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 12:20 | 0 | 84.94 | 85 | 0 | 40.01 | 40 | 104.5511 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.8 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:21 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.75 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 12:22 | 0 | 84.96 | 85 | 0 | 40 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2076 | 0 | 31.74 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 12:23 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 104.5511 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.74 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 12:24 | 0 | 84.99 | 85 | 0 | 40 | 40 | 104.6016 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.2442 | 0 | 31.86 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 12:25 | 0 | 85 | 85 | 0 | 39.91 | 40 | 104.5511 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.96 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:26 | 0 | 85.03 | 85 | 0 | 40 | 40 | 104.5511 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.98 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 12:27 | 0 | 84.94 | 85 | 0 | 39.89 | 40 | 104.6016 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.2442 | 0 | 31.94 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 12:28 | 0 | 85 | 85 | 0 | 39.87 | 40 | 104.5511 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.9 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9887 | 19.0119
2008-08-05 12:29 | 0 | 85 | 85 | 0 | 39.88 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.87 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:30 | 0 | 85 | 85 | 0 | 39.96 | 40 | 104.6016 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.84 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 12:31 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 104.6016 | 0 | 30.2504 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.2442 | 0 | 31.82 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:32 | 0 | 84.94 | 85 | 0 | 39.94 | 40 | 104.5511 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.78 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:33 | 0 | 84.96 | 85 | 0 | 40 | 40 | 104.6016 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.77 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.0119
2008-08-05 12:34 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.78 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 12:35 | 0 | 84.96 | 85 | 0 | 40.01 | 40 | 104.6521 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96237 | 31.01 | 30.2442 | 0 | 31.82 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0147 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 12:36 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.82 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 12:37 | 0 | 84.94 | 85 | 0 | 39.91 | 40 | 104.6016 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96191 | 31.01 | 30.2442 | 0 | 31.79 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9887 | 19.0119
2008-08-05 12:38 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 104.6521 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96237 | 31.01 | 30.2442 | 0 | 31.75 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 12:39 | 0 | 84.94 | 85 | 0 | 39.97 | 40 | 104.5511 | 0 | 30.2504 | 29.8 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.72 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 12:40 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 104.6016 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.7 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9707 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 12:41 | 0 | 85 | 85 | 0 | 40.04 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.63 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9877 | 19.0119
2008-08-05 12:42 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.61 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 12:43 | 0 | 85 | 85 | 0 | 39.9 | 40 | 104.6016 | 0 | 30.2515 | 29.8154 | 0 |  | 0.96194 | 31.01 | 30.232 | 0 | 31.59 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9887 | 19.0119
2008-08-05 12:44 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 104.6016 | 0 | 30.2504 | 29.8154 | 0 |  | 0.96194 | 31.01 | 30.2442 | 0 | 31.59 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.978 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 12:45 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 104.6016 | 0 | 30.2662 | 29.8298 | 0 |  | 0.96194 | 31.01 | 30.2442 | 0 | 31.66 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9753 | 19.0119
2008-08-05 12:46 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 104.6016 | 0 | 30.2515 | 29.8298 | 0 |  | 0.96194 | 31.01 | 30.232 | 0 | 31.75 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9902 | 19.0119
2008-08-05 12:47 | 0 | 84.99 | 85 | 0 | 40 | 40 | 104.5511 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.83 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9877 | 19.0119
2008-08-05 12:48 | 0 | 84.92 | 85 | 0 | 39.98 | 40 | 104.5511 | 0 | 30.2662 | 29.8154 | 0 |  | 0.96148 | 31.01 | 30.232 | 0 | 31.82 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9763 | 19.0119
2008-08-05 12:49 | 0 | 84.91 | 85 | 0 | 39.97 | 40 | 104.5511 | 0 | 30.2504 | 29.8288 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.81 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:50 | 0 | 85.02 | 85 | 0 | 40.03 | 40 | 104.5511 | 0 | 30.2504 | 29.8154 | 0 |  | 0.96148 | 31.01 | 30.232 | 0 | 31.73 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.0119
2008-08-05 12:51 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 104.5006 | 0 | 30.2515 | 29.8298 | 0 |  | 0.96101 | 31.01 | 30.1954 | 0 | 31.72 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 12:52 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 104.5511 | 0 | 30.2662 | 29.8154 | 0 |  | 0.96148 | 31.01 | 30.232 | 0 | 31.68 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:53 | 0 | 85 | 85 | 0 | 39.9 | 40 | 104.5006 | 0 | 30.2367 | 29.8154 | 0 |  | 0.96101 | 31.01 | 30.232 | 0 | 31.65 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20.0073 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 12:54 | 0 | 85 | 85 | 0 | 39.93 | 40 | 104.6016 | 0 | 30.2515 | 29.8154 | 0 |  | 0.96194 | 31.01 | 30.232 | 0 | 31.61 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9902 | 19.0119
2008-08-05 12:55 | 0 | 85 | 85 | 0 | 40.01 | 40 | 104.5511 | 0 | 30.2515 | 29.8154 | 0 |  | 0.96148 | 31.01 | 30.232 | 0 | 31.63 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9655 | 19.0119
2008-08-05 12:56 | 0 | 84.96 | 85 | 0 | 40.05 | 40 | 104.6016 | 0 | 30.2652 | 29.8 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.75 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9902 | 19.0119
2008-08-05 12:57 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 104.6016 | 0 | 30.2662 | 29.8154 | 0 |  | 0.96148 | 31.01 | 30.232 | 0 | 31.86 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9665 | 19.0119
2008-08-05 12:58 | 0 | 84.95 | 85 | 0 | 40.07 | 40 | 104.5511 | 0 | 30.2515 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.2442 | 0 | 31.87 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0026 | 19.0119
2008-08-05 12:59 | 0 | 85 | 85 | 0 | 40.1 | 40 | 104.5006 | 0 | 30.2504 | 29.8288 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.84 | 20.399 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0063 | 19.0119
2008-08-05 13:00 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 104.6016 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.2076 | 0 | 31.82 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9887 | 19.0119
2008-08-05 13:01 | 0 | 84.96 | 85 | 0 | 39.91 | 40 | 104.5006 | 0 | 30.2357 | 29.8144 | 0 |  | 0.96098 | 31.01 | 30.232 | 0 | 31.77 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 31.0037 | 19.0119
2008-08-05 13:02 | 0 | 84.95 | 85 | 0 | 40.04 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.75 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 20 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 13:03 | 0 | 85 | 85 | 0 | 40.05 | 40 | 104.5511 | 0 | 30.2357 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.66 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 13:04 | 0 | 85.05 | 85 | 0 | 40.03 | 40 | 104.5511 | 0 | 30.2504 | 29.8144 | 0 |  | 0.96144 | 31.01 | 30.232 | 0 | 31.69 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9913 | 19.0119
2008-08-05 13:05 | 0 | 85 | 85 | 0 | 40 | 40 | 104.6016 | 0 | 30.2652 | 29.8144 | 0 |  | 0.96191 | 31.01 | 30.232 | 0 | 31.77 | 20.3991 | 0 | 0.85172 | 20 | 20.021 |  | 19.9927 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 20.0512 | -1 | 20.031 | 30.9753 | 19.0119
2008-08-05 13:06 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 104.6016 | 0 | 30.2652 | 29.8288 | 0 |  | 0.96191 | 30.89 | 30.232 | 0 | 31.94 | 40.8029 | 0 | 1.7036 | 15 | 20.021 |  | 19.9853 | 0 | 0 | 0 | 25 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 20.0512 | -1 | 15.0309 | 30.9887 | 38.0283
2008-08-05 13:07 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 99.8536 | 0 | 28.4257 | 27.9951 | 0 |  | 0.90622 | 30.89 | 28.3883 | 0 | 31.55 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 23.89 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0284 | -1 | 15.0309 | 30.8703 | 0
2008-08-05 13:08 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 99.9546 | 0 | 28.3626 | 27.9632 | 0 |  | 0.90525 | 30.89 | 28.3394 | 0 | 30.87 | 0 | 0 | 0 | 15 | 15 |  | 14.967 | 0 | 0 | 0 | 25 | 24.51 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0284 | -1 | 15.0309 | 30.8674 | 0
2008-08-05 13:09 | 0 | 85.03 | 85 | 0 | 40.05 | 40 | 99.9546 | 0 | 28.3636 | 27.9359 | 0 |  | 0.90437 | 30.89 | 28.3394 | 0 | 30.37 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.8 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0284 | -1 | 15.0309 | 30.8613 | 0
2008-08-05 13:10 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 99.8536 | 0 | 28.3542 | 27.8993 | 0 |  | 0.90318 | 30.89 | 28.3028 | 0 | 30.12 | 0 | 0 | 0 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 24.89 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0284 | -1 | 15.0309 | 30.8333 | 0
2008-08-05 13:11 | 0 | 85 | 85 | 0 | 39.92 | 40 | 99.8536 | 0 | 28.3212 | 27.8946 | 0 |  | 0.90303 | 30.89 | 28.3028 | 0 | 29.98 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.94 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8686 | 0
2008-08-05 13:12 | 0 | 85 | 85 | 0 | 39.97 | 40 | 99.8031 | 0 | 28.3043 | 27.8918 | 0 |  | 0.90248 | 30.89 | 28.2784 | 0 | 29.84 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.97 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8582 | 0
2008-08-05 13:13 | 0 | 85.05 | 85 | 0 | 40.01 | 40 | 99.8536 | 0 | 28.3175 | 27.8909 | 0 |  | 0.90291 | 30.89 | 28.2906 | 0 | 29.79 | 0 | 0 | 0 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 24.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8603 | 0
2008-08-05 13:14 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 99.8536 | 0 | 28.3166 | 27.8899 | 0 |  | 0.90288 | 30.89 | 28.2906 | 0 | 29.78 | 0 | 0 | 0 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8478 | 0
2008-08-05 13:15 | 0 | 84.94 | 85 | 0 | 39.92 | 40 | 99.8536 | 0 | 28.3156 | 27.8749 | 0 |  | 0.90285 | 30.89 | 28.2906 | 0 | 29.73 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8488 | 0
2008-08-05 13:16 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 99.8031 | 0 | 28.3015 | 27.8598 | 0 |  | 0.90236 | 30.89 | 28.2906 | 0 | 29.71 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8645 | 0
2008-08-05 13:17 | 0 | 84.94 | 85 | 0 | 39.97 | 40 | 99.8536 | 0 | 28.3015 | 27.889 | 0 |  | 0.90285 | 30.89 | 28.2906 | 0 | 29.67 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.879 | 0
2008-08-05 13:18 | 0 | 84.99 | 85 | 0 | 40.01 | 40 | 99.7526 | 0 | 28.3005 | 27.8881 | 0 |  | 0.9019 | 30.89 | 28.2906 | 0 | 29.75 | 0 | 0 | 0 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.879 | 0
2008-08-05 13:19 | 0 | 85.05 | 85 | 0 | 40.03 | 40 | 99.8536 | 0 | 28.3005 | 27.8881 | 0 |  | 0.90282 | 30.89 | 28.2906 | 0 | 29.88 | 0 | 0 | 0 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8634 | 0
2008-08-05 13:20 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 99.8536 | 0 | 28.2864 | 27.8881 | 0 |  | 0.90282 | 30.89 | 28.2784 | 0 | 29.99 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8499 | 0
2008-08-05 13:21 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 99.8031 | 0 | 28.3147 | 27.8881 | 0 |  | 0.90236 | 30.89 | 28.2784 | 0 | 30.03 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8811 | 0
2008-08-05 13:22 | 0 | 84.95 | 85 | 0 | 39.9 | 40 | 99.8031 | 0 | 28.3005 | 27.8881 | 0 |  | 0.90236 | 30.89 | 28.2906 | 0 | 30.04 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.879 | 0
2008-08-05 13:23 | 0 | 84.99 | 85 | 0 | 40 | 40 | 99.8536 | 0 | 28.3147 | 27.8881 | 0 |  | 0.90282 | 30.89 | 28.2906 | 0 | 30 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8655 | 0
2008-08-05 13:24 | 0 | 85 | 85 | 0 | 39.98 | 40 | 99.8536 | 0 | 28.3147 | 27.8881 | 0 |  | 0.90282 | 30.89 | 28.2906 | 0 | 29.95 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8946 | 0
2008-08-05 13:25 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 99.8536 | 0 | 28.3147 | 27.8881 | 0 |  | 0.90282 | 30.89 | 28.2662 | 0 | 29.89 | 0 | 0 | 0 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8634 | 0
2008-08-05 13:26 | 0 | 84.99 | 85 | 0 | 39.91 | 40 | 99.8536 | 0 | 28.3005 | 27.8881 | 0 |  | 0.90282 | 30.89 | 28.2906 | 0 | 29.88 | 0 | 0 | 0 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8655 | 0
2008-08-05 13:27 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 99.8536 | 0 | 28.3005 | 27.889 | 0 |  | 0.90285 | 30.89 | 28.2906 | 0 | 29.91 | 0 | 0 | 0 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8499 | 0
2008-08-05 13:28 | 0 | 85.02 | 85 | 0 | 40 | 40 | 99.8536 | 0 | 28.3005 | 27.8881 | 0 |  | 0.90282 | 30.89 | 28.2906 | 0 | 30.01 | 0 | 0 | 0 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8645 | 0
2008-08-05 13:29 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 99.8536 | 0 | 28.2874 | 27.889 | 0 |  | 0.90285 | 30.89 | 28.2784 | 0 | 30.08 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.879 | 0
2008-08-05 13:30 | 0 | 85 | 85 | 0 | 39.93 | 40 | 99.8536 | 0 | 28.3005 | 27.8881 | 0 |  | 0.90282 | 30.89 | 28.2906 | 0 | 30.09 | 0 | 0 | 0 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8655 | 0
2008-08-05 13:31 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 99.8031 | 0 | 28.3156 | 27.8608 | 0 |  | 0.90239 | 30.89 | 28.2784 | 0 | 30.04 | 0 | 0 | 0 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8499 | 0
2008-08-05 13:32 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 99.8536 | 0 | 28.3005 | 27.8739 | 0 |  | 0.90282 | 30.89 | 28.2906 | 0 | 30.02 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.878 | 0
2008-08-05 13:33 | 0 | 85 | 85 | 0 | 40.03 | 40 | 99.8536 | 0 | 28.3156 | 27.8749 | 0 |  | 0.90285 | 30.89 | 28.2784 | 0 | 29.96 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8655 | 0
2008-08-05 13:34 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 99.8031 | 0 | 28.3005 | 27.8881 | 0 |  | 0.90236 | 30.89 | 28.2906 | 0 | 29.94 | 0 | 0 | 0 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8655 | 0
2008-08-05 13:35 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 99.8031 | 0 | 28.3015 | 27.889 | 0 |  | 0.90239 | 30.89 | 28.2784 | 0 | 29.89 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8499 | 0
2008-08-05 13:36 | 0 | 84.92 | 85 | 0 | 39.88 | 40 | 99.8536 | 0 | 28.3156 | 27.889 | 0 |  | 0.90285 | 30.89 | 28.2784 | 0 | 29.89 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8623 | 0
2008-08-05 13:37 | 0 | 85.03 | 85 | 0 | 40 | 40 | 99.8536 | 0 | 28.2874 | 27.889 | 0 |  | 0.90285 | 30.89 | 28.2906 | 0 | 29.93 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8488 | 0
2008-08-05 13:38 | 0 | 85 | 85 | 0 | 40.03 | 40 | 99.7526 | 0 | 28.3156 | 27.8749 | 0 |  | 0.90194 | 30.89 | 28.2906 | 0 | 30 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8623 | 0
2008-08-05 13:39 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 99.8536 | 0 | 28.3156 | 27.889 | 0 |  | 0.90285 | 30.89 | 28.2784 | 0 | 30.05 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8645 | 0
2008-08-05 13:40 | 0 | 85 | 85 | 0 | 39.9 | 40 | 99.8536 | 0 | 28.3156 | 27.889 | 0 |  | 0.90285 | 30.89 | 28.2784 | 0 | 30.02 | 0 | 0 | 0 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8218 | 0
2008-08-05 13:41 | 0 | 85 | 85 | 0 | 39.89 | 40 | 99.8536 | 0 | 28.3156 | 27.8749 | 0 |  | 0.90285 | 30.89 | 28.2906 | 0 | 29.98 | 0 | 0 | 0 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.878 | 0
2008-08-05 13:42 | 0 | 85.05 | 85 | 0 | 39.91 | 40 | 99.8031 | 0 | 28.3156 | 27.889 | 0 |  | 0.90239 | 30.89 | 28.2906 | 0 | 29.94 | 0 | 0 | 0 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8623 | 0
2008-08-05 13:43 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 99.8536 | 0 | 28.3015 | 27.8608 | 0 |  | 0.90285 | 30.89 | 28.2906 | 0 | 29.88 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8488 | 0
2008-08-05 13:44 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 99.8031 | 0 | 28.3015 | 27.889 | 0 |  | 0.90239 | 30.89 | 28.2784 | 0 | 29.83 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.878 | 0
2008-08-05 13:45 | 0 | 84.96 | 85 | 0 | 40 | 40 | 99.8536 | 0 | 28.3156 | 27.889 | 0 |  | 0.90285 | 30.89 | 28.2784 | 0 | 29.8 | 0 | 0 | 0 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8488 | 0
2008-08-05 13:46 | 0 | 84.92 | 85 | 0 | 39.98 | 40 | 99.8536 | 0 | 28.3015 | 27.8749 | 0 |  | 0.90285 | 30.89 | 28.2784 | 0 | 29.82 | 0 | 0 | 0 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8374 | 0
2008-08-05 13:47 | 0 | 84.92 | 85 | 0 | 39.96 | 40 | 99.8536 | 0 | 28.3166 | 27.8758 | 0 |  | 0.90288 | 30.89 | 28.2784 | 0 | 29.83 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8623 | 0
2008-08-05 13:48 | 0 | 85.03 | 85 | 0 | 39.96 | 40 | 99.8536 | 0 | 28.3156 | 27.889 | 0 |  | 0.90285 | 30.89 | 28.2784 | 0 | 29.88 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8634 | 0
2008-08-05 13:49 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 99.8536 | 0 | 28.3166 | 27.8899 | 0 |  | 0.90288 | 30.89 | 28.2784 | 0 | 29.91 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8623 | 0
2008-08-05 13:50 | 0 | 84.96 | 85 | 0 | 40.11 | 40 | 99.8536 | 0 | 28.3024 | 27.8899 | 0 |  | 0.90288 | 30.89 | 28.2784 | 0 | 29.91 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8904 | 0
2008-08-05 13:51 | 0 | 84.92 | 85 | 0 | 39.96 | 40 | 99.8536 | 0 | 28.3156 | 27.889 | 0 |  | 0.90285 | 30.89 | 28.2784 | 0 | 29.88 | 0 | 0 | 0 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0283 | -1 | 15.0309 | 30.8478 | 0
2008-08-05 13:52 | 0 | 85.03 | 85 | 0 | 40.01 | 40 | 102.8337 | 0 | 29.1496 | 28.7213 | 0 |  | 0.92979 | 30.89 | 29.1087 | 0 | 29.84 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.845 | 38.0311
2008-08-05 13:53 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 102.7832 | 0 | 29.1496 | 28.7072 | 0 |  | 0.92934 | 30.89 | 29.1087 | 0 | 29.79 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8602 | 38.0311
2008-08-05 13:54 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 102.8337 | 0 | 29.1355 | 28.7072 | 0 |  | 0.92979 | 30.89 | 29.0965 | 0 | 29.76 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8612 | 38.0311
2008-08-05 13:55 | 0 | 85 | 85 | 0 | 39.97 | 40 | 102.7327 | 0 | 29.1355 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 29.75 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8764 | 38.0312
2008-08-05 13:56 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 102.7832 | 0 | 29.1355 | 28.7213 | 0 |  | 0.92934 | 30.89 | 29.1209 | 0 | 29.75 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.846 | 38.0311
2008-08-05 13:57 | 0 | 84.99 | 85 | 0 | 39.84 | 40 | 102.7327 | 0 | 29.1355 | 28.7072 | 0 |  | 0.92888 | 30.89 | 29.1087 | 0 | 29.71 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.846 | 38.0311
2008-08-05 13:58 | 0 | 85.07 | 85 | 0 | 39.97 | 40 | 102.8337 | 0 | 29.1355 | 28.6931 | 0 |  | 0.92979 | 30.89 | 29.1087 | 0 | 29.78 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8481 | 38.0312
2008-08-05 13:59 | 0 | 85.02 | 85 | 0 | 39.98 | 40 | 102.7832 | 0 | 29.1355 | 28.7213 | 0 |  | 0.92934 | 30.89 | 29.1087 | 0 | 29.89 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8612 | 38.0311
2008-08-05 14:00 | 0 | 84.97 | 85 | 0 | 40 | 40 | 102.8337 | 0 | 29.1496 | 28.7072 | 0 |  | 0.92979 | 30.89 | 29.1209 | 0 | 29.99 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8591 | 38.0311
2008-08-05 14:01 | 0 | 84.95 | 85 | 0 | 39.89 | 40 | 102.7327 | 0 | 29.1355 | 28.7213 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.01 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0311
2008-08-05 14:02 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 102.7832 | 0 | 29.1345 | 28.7063 | 0 |  | 0.92931 | 30.89 | 29.1087 | 0 | 30.02 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8491 | 38.0311
2008-08-05 14:03 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 102.7832 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92934 | 30.89 | 29.1087 | 0 | 30.02 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8764 | 38.0312
2008-08-05 14:04 | 0 | 85 | 85 | 0 | 40 | 40 | 102.7832 | 0 | 29.1345 | 28.6639 | 0 |  | 0.92931 | 30.89 | 29.1087 | 0 | 30.01 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8622 | 38.0312
2008-08-05 14:05 | 0 | 84.94 | 85 | 0 | 40.01 | 40 | 102.7832 | 0 | 29.1355 | 28.7072 | 0 |  | 0.92931 | 30.89 | 29.0965 | 0 | 29.98 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8481 | 38.0311
2008-08-05 14:06 | 0 | 84.97 | 85 | 0 | 39.81 | 40 | 102.7832 | 0 | 29.1204 | 28.6649 | 0 |  | 0.92934 | 30.89 | 29.1087 | 0 | 29.96 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8774 | 38.0312
2008-08-05 14:07 | 0 | 84.94 | 85 | 0 | 39.95 | 40 | 102.7832 | 0 | 29.1355 | 28.7063 | 0 |  | 0.92931 | 30.89 | 29.0965 | 0 | 29.95 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.967 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8491 | 38.0311
2008-08-05 14:08 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 102.7832 | 0 | 29.1345 | 28.7072 | 0 |  | 0.92934 | 30.89 | 29.0965 | 0 | 29.96 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8491 | 38.0311
2008-08-05 14:09 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 102.7832 | 0 | 29.1214 | 28.7063 | 0 |  | 0.92931 | 30.89 | 29.1087 | 0 | 30.07 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8612 | 38.0311
2008-08-05 14:10 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 102.7327 | 0 | 29.1355 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.1087 | 0 | 30.18 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8612 | 38.0312
2008-08-05 14:11 | 0 | 85 | 85 | 0 | 40.03 | 40 | 102.7832 | 0 | 29.1345 | 28.6922 | 0 |  | 0.92931 | 30.89 | 29.1087 | 0 | 30.28 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8764 | 38.0312
2008-08-05 14:12 | 0 | 84.99 | 85 | 0 | 40.04 | 40 | 102.7832 | 0 | 29.1204 | 28.7063 | 0 |  | 0.92931 | 30.89 | 29.0965 | 0 | 30.25 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8622 | 38.0311
2008-08-05 14:13 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 102.7832 | 0 | 29.1345 | 28.6922 | 0 |  | 0.92931 | 30.89 | 29.1087 | 0 | 30.27 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8622 | 38.0312
2008-08-05 14:14 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 102.7832 | 0 | 29.1345 | 28.6922 | 0 |  | 0.92931 | 30.89 | 29.0965 | 0 | 30.24 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8491 | 38.0312
2008-08-05 14:15 | 0 | 84.96 | 85 | 0 | 40.07 | 40 | 102.7327 | 0 | 29.1345 | 28.7072 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.21 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8622 | 38.0311
2008-08-05 14:16 | 0 | 84.96 | 85 | 0 | 40.07 | 40 | 102.7832 | 0 | 29.1345 | 28.6922 | 0 |  | 0.92931 | 30.89 | 29.072 | 0 | 30.21 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8774 | 38.0312
2008-08-05 14:17 | 0 | 85 | 85 | 0 | 40.11 | 40 | 102.8337 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92979 | 30.89 | 29.0965 | 0 | 30.17 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8491 | 38.0312
2008-08-05 14:18 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 102.7327 | 0 | 29.1345 | 28.7204 | 0 |  | 0.92885 | 30.89 | 29.1087 | 0 | 30.21 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8622 | 38.0311
2008-08-05 14:19 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 102.7832 | 0 | 29.1496 | 28.7072 | 0 |  | 0.92934 | 30.89 | 29.0842 | 0 | 30.21 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8471 | 38.0311
2008-08-05 14:20 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 102.7832 | 0 | 29.1345 | 28.7063 | 0 |  | 0.92934 | 30.89 | 29.1087 | 0 | 30.32 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8774 | 38.0311
2008-08-05 14:21 | 0 | 84.99 | 85 | 0 | 40.07 | 40 | 102.7832 | 0 | 29.1345 | 28.7063 | 0 |  | 0.92931 | 30.89 | 29.0965 | 0 | 30.39 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8936 | 38.0311
2008-08-05 14:22 | 0 | 85 | 85 | 0 | 39.97 | 40 | 102.7832 | 0 | 29.1355 | 28.6931 | 0 |  | 0.92934 | 30.89 | 29.1087 | 0 | 30.39 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8612 | 38.0312
2008-08-05 14:23 | 0 | 85.02 | 85 | 0 | 40.04 | 40 | 102.7327 | 0 | 29.1345 | 28.7063 | 0 |  | 0.92885 | 30.89 | 29.0965 | 0 | 30.4 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8612 | 38.0311
2008-08-05 14:24 | 0 | 85.03 | 85 | 0 | 40 | 40 | 102.7327 | 0 | 29.1355 | 28.7072 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.35 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8491 | 38.0311
2008-08-05 14:25 | 0 | 84.99 | 85 | 0 | 40.04 | 40 | 102.7327 | 0 | 29.1345 | 28.7063 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.32 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8643 | 38.0311
2008-08-05 14:26 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 102.7327 | 0 | 29.1204 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.1087 | 0 | 30.29 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8774 | 38.0312
2008-08-05 14:27 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 102.7327 | 0 | 29.1345 | 28.6922 | 0 |  | 0.92885 | 30.89 | 29.1087 | 0 | 30.25 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8612 | 38.0312
2008-08-05 14:28 | 0 | 84.92 | 85 | 0 | 39.96 | 40 | 102.7832 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92934 | 30.89 | 29.072 | 0 | 30.23 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8622 | 38.0312
2008-08-05 14:29 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 102.7327 | 0 | 29.1355 | 28.7072 | 0 |  | 0.92888 | 30.89 | 29.1087 | 0 | 30.2 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.9068 | 38.0311
2008-08-05 14:30 | 0 | 84.97 | 85 | 0 | 40.07 | 40 | 102.7327 | 0 | 29.1355 | 28.7072 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.24 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8612 | 38.0311
2008-08-05 14:31 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 102.7327 | 0 | 29.0931 | 28.7072 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.21 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0311
2008-08-05 14:32 | 0 | 85.02 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.27 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 14:33 | 0 | 85.02 | 85 | 0 | 39.94 | 40 | 102.7327 | 0 | 29.1355 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.33 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8481 | 38.0312
2008-08-05 14:34 | 0 | 84.92 | 85 | 0 | 40 | 40 | 102.7327 | 0 | 29.0931 | 28.7072 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.37 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8805 | 38.0311
2008-08-05 14:35 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 102.6317 | 0 | 29.1355 | 28.7072 | 0 |  | 0.92797 | 30.89 | 29.0965 | 0 | 30.35 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0311
2008-08-05 14:36 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 102.7832 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92934 | 30.89 | 29.072 | 0 | 30.33 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.9068 | 38.0312
2008-08-05 14:37 | 0 | 84.94 | 85 | 0 | 40.03 | 40 | 102.7832 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92934 | 30.89 | 29.0965 | 0 | 30.31 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8936 | 38.0312
2008-08-05 14:38 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.27 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0312
2008-08-05 14:39 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 102.7327 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.23 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8936 | 38.0312
2008-08-05 14:40 | 0 | 85 | 85 | 0 | 39.97 | 40 | 102.7327 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.17 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8805 | 38.0312
2008-08-05 14:41 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.16 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 14:42 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.13 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0312
2008-08-05 14:43 | 0 | 84.99 | 85 | 0 | 39.84 | 40 | 102.7327 | 0 | 29.1214 | 28.7072 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.1 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0311
2008-08-05 14:44 | 0 | 85 | 85 | 0 | 40.04 | 40 | 102.7327 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0598 | 0 | 30.06 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 14:45 | 0 | 85 | 85 | 0 | 39.95 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.05 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0312
2008-08-05 14:46 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 102.7327 | 0 | 29.1223 | 28.6941 | 0 |  | 0.92891 | 30.89 | 29.0842 | 0 | 30.1 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.967 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8491 | 38.0312
2008-08-05 14:47 | 0 | 85 | 85 | 0 | 39.96 | 40 | 102.7327 | 0 | 29.1355 | 28.7072 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.21 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8926 | 38.0311
2008-08-05 14:48 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 102.7327 | 0 | 29.1223 | 28.6941 | 0 |  | 0.92891 | 30.89 | 29.0842 | 0 | 30.29 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8805 | 38.0312
2008-08-05 14:49 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 102.6317 | 0 | 29.0931 | 28.7072 | 0 |  | 0.92797 | 30.89 | 29.0842 | 0 | 30.29 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0311
2008-08-05 14:50 | 0 | 84.94 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.0965 | 0 | 30.27 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8936 | 38.0312
2008-08-05 14:51 | 0 | 84.99 | 85 | 0 | 40 | 40 | 102.7327 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.25 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8764 | 38.0312
2008-08-05 14:52 | 0 | 85 | 85 | 0 | 40 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.18 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0312
2008-08-05 14:53 | 0 | 84.97 | 85 | 0 | 39.87 | 40 | 102.7327 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.18 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8805 | 38.0312
2008-08-05 14:54 | 0 | 85 | 85 | 0 | 40.1 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.13 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0312
2008-08-05 14:55 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 102.7327 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.12 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8481 | 38.0312
2008-08-05 14:56 | 0 | 84.95 | 85 | 0 | 39.92 | 40 | 102.7327 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.1087 | 0 | 30.07 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0312
2008-08-05 14:57 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 102.6317 | 0 | 29.0941 | 28.6941 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.04 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.967 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8795 | 38.0312
2008-08-05 14:58 | 0 | 85 | 85 | 0 | 39.95 | 40 | 102.7832 | 0 | 29.1365 | 28.7082 | 0 |  | 0.92937 | 30.89 | 29.0965 | 0 | 30.13 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8339 | 38.0311
2008-08-05 14:59 | 0 | 85.02 | 85 | 0 | 40.03 | 40 | 102.7327 | 0 | 29.1223 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.24 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0312
2008-08-05 15:00 | 0 | 84.92 | 85 | 0 | 39.86 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.32 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8491 | 38.0312
2008-08-05 15:01 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 102.6317 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.0842 | 0 | 30.33 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 15:02 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 102.7327 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.29 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8805 | 38.0312
2008-08-05 15:03 | 0 | 84.96 | 85 | 0 | 40.05 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.28 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8805 | 38.0312
2008-08-05 15:04 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 102.6317 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.0842 | 0 | 30.29 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0312
2008-08-05 15:05 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 102.7327 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.21 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0312
2008-08-05 15:06 | 0 | 84.94 | 85 | 0 | 40.01 | 40 | 102.7327 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0598 | 0 | 30.2 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 15:07 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.25 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9963 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 15:08 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 102.7327 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0598 | 0 | 30.36 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 15:09 | 0 | 84.94 | 85 | 0 | 39.94 | 40 | 102.6317 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.0842 | 0 | 30.42 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.967 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0312
2008-08-05 15:10 | 0 | 85 | 85 | 0 | 39.87 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.44 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0312
2008-08-05 15:11 | 0 | 85.05 | 85 | 0 | 39.98 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.39 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8239 | 38.0312
2008-08-05 15:12 | 0 | 85.02 | 85 | 0 | 40.01 | 40 | 102.7327 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.41 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 15:13 | 0 | 85 | 85 | 0 | 39.93 | 40 | 102.6317 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.0965 | 0 | 30.37 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 15:14 | 0 | 84.92 | 85 | 0 | 39.93 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.33 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 15:15 | 0 | 84.96 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.0842 | 0 | 30.25 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8239 | 38.0312
2008-08-05 15:16 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 102.7327 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.25 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8936 | 38.0312
2008-08-05 15:17 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.24 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 15:18 | 0 | 85.03 | 85 | 0 | 39.94 | 40 | 102.7327 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.1087 | 0 | 30.25 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8936 | 38.0312
2008-08-05 15:19 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.29 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0312
2008-08-05 15:20 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.34 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8936 | 38.0312
2008-08-05 15:21 | 0 | 85.02 | 85 | 0 | 40.04 | 40 | 102.7832 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92934 | 30.89 | 29.1087 | 0 | 30.34 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 15:22 | 0 | 84.92 | 85 | 0 | 39.95 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.34 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 15:23 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.29 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0312
2008-08-05 15:24 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 102.7832 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92934 | 30.89 | 29.0965 | 0 | 30.27 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0312
2008-08-05 15:25 | 0 | 84.96 | 85 | 0 | 40.01 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.21 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0312
2008-08-05 15:26 | 0 | 84.99 | 85 | 0 | 39.9 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.21 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 15:27 | 0 | 85 | 85 | 0 | 40.04 | 40 | 102.7832 | 0 | 29.1214 | 28.7072 | 0 |  | 0.92934 | 30.89 | 29.0598 | 0 | 30.16 | 40.8059 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0311
2008-08-05 15:28 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 102.7327 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.13 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8936 | 38.0312
2008-08-05 15:29 | 0 | 84.94 | 85 | 0 | 40.03 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.13 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8805 | 38.0312
2008-08-05 15:30 | 0 | 84.97 | 85 | 0 | 40.07 | 40 | 102.7327 | 0 | 29.1214 | 28.6941 | 0 |  | 0.92891 | 30.89 | 29.0842 | 0 | 30.21 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0312
2008-08-05 15:31 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 102.6317 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.0965 | 0 | 30.34 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8622 | 38.0312
2008-08-05 15:32 | 0 | 84.96 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.43 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 15:33 | 0 | 84.97 | 85 | 0 | 40.01 | 40 | 102.7327 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.46 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0312
2008-08-05 15:34 | 0 | 84.95 | 85 | 0 | 40 | 40 | 102.7327 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0598 | 0 | 30.41 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 15:35 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 102.6317 | 0 | 29.079 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.37 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8239 | 38.0312
2008-08-05 15:36 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 102.7327 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.31 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0312
2008-08-05 15:37 | 0 | 85 | 85 | 0 | 39.96 | 40 | 102.7327 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.29 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8694 | 38.0312
2008-08-05 15:38 | 0 | 84.97 | 85 | 0 | 39.91 | 40 | 102.5812 | 0 | 29.079 | 28.6649 | 0 |  | 0.92751 | 30.89 | 29.0598 | 0 | 30.27 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9963 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8826 | 38.0312
2008-08-05 15:39 | 0 | 85 | 85 | 0 | 39.9 | 40 | 102.7327 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.29 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8239 | 38.0312
2008-08-05 15:40 | 0 | 84.94 | 85 | 0 | 39.96 | 40 | 102.6317 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.36 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 15:41 | 0 | 85 | 85 | 0 | 39.97 | 40 | 102.6317 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.51 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8239 | 38.0312
2008-08-05 15:42 | 0 | 84.91 | 85 | 0 | 39.83 | 40 | 102.7327 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.58 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 15:43 | 0 | 84.94 | 85 | 0 | 39.9 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.58 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0312
2008-08-05 15:44 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 102.5812 | 0 | 29.079 | 28.6508 | 0 |  | 0.92751 | 30.89 | 29.0598 | 0 | 30.55 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8805 | 38.0313
2008-08-05 15:45 | 0 | 84.99 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.51 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 15:46 | 0 | 84.88 | 85 | 0 | 40.04 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.48 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8936 | 38.0312
2008-08-05 15:47 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 102.7327 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.44 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0312
2008-08-05 15:48 | 0 | 84.96 | 85 | 0 | 40 | 40 | 102.7327 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.42 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0312
2008-08-05 15:49 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 102.7327 | 0 | 29.0931 | 28.6508 | 0 |  | 0.92888 | 30.89 | 29.0598 | 0 | 30.39 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0313
2008-08-05 15:50 | 0 | 84.95 | 85 | 0 | 40 | 40 | 102.7327 | 0 | 29.0931 | 28.6508 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.37 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0313
2008-08-05 15:51 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 102.7327 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.41 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8239 | 38.0312
2008-08-05 15:52 | 0 | 85.07 | 85 | 0 | 40.04 | 40 | 102.7327 | 0 | 29.079 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.46 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8936 | 38.0312
2008-08-05 15:53 | 0 | 84.94 | 85 | 0 | 39.95 | 40 | 102.6317 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.46 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8936 | 38.0312
2008-08-05 15:54 | 0 | 84.94 | 85 | 0 | 39.92 | 40 | 102.6317 | 0 | 29.0931 | 28.6508 | 0 |  | 0.92797 | 30.89 | 29.0965 | 0 | 30.45 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0313
2008-08-05 15:55 | 0 | 85 | 85 | 0 | 39.97 | 40 | 102.7327 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.43 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8805 | 38.0312
2008-08-05 15:56 | 0 | 84.96 | 85 | 0 | 40.01 | 40 | 102.7327 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.38 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.967 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8239 | 38.0312
2008-08-05 15:57 | 0 | 84.99 | 85 | 0 | 39.9 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.36 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0312
2008-08-05 15:58 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 102.6317 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.0842 | 0 | 30.31 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8239 | 38.0312
2008-08-05 15:59 | 0 | 84.94 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.0842 | 0 | 30.25 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 16:00 | 0 | 85 | 85 | 0 | 39.94 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.0842 | 0 | 30.25 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 16:01 | 0 | 85 | 85 | 0 | 39.92 | 40 | 102.7327 | 0 | 29.1214 | 28.6941 | 0 |  | 0.92891 | 30.89 | 29.0598 | 0 | 30.31 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8826 | 38.0312
2008-08-05 16:02 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.38 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 16:03 | 0 | 85 | 85 | 0 | 39.97 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.41 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8663 | 38.0312
2008-08-05 16:04 | 0 | 85 | 85 | 0 | 39.93 | 40 | 102.5812 | 0 | 29.079 | 28.6508 | 0 |  | 0.92751 | 30.89 | 29.072 | 0 | 30.38 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8532 | 38.0313
2008-08-05 16:05 | 0 | 84.96 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.37 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8694 | 38.0312
2008-08-05 16:06 | 0 | 85 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.35 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 16:07 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.31 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8694 | 38.0312
2008-08-05 16:08 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 102.5812 | 0 | 29.0799 | 28.6659 | 0 |  | 0.92754 | 30.89 | 29.072 | 0 | 30.24 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 16:09 | 0 | 84.95 | 85 | 0 | 39.88 | 40 | 102.6317 | 0 | 29.0941 | 28.6649 | 0 |  | 0.928 | 30.89 | 29.072 | 0 | 30.23 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8532 | 38.0312
2008-08-05 16:10 | 0 | 85 | 85 | 0 | 39.97 | 40 | 102.5812 | 0 | 29.0941 | 28.6518 | 0 |  | 0.92754 | 30.89 | 29.0354 | 0 | 30.18 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8532 | 38.0313
2008-08-05 16:11 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.21 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8795 | 38.0312
2008-08-05 16:12 | 0 | 85 | 85 | 0 | 40.01 | 40 | 102.6317 | 0 | 29.079 | 28.6659 | 0 |  | 0.928 | 30.89 | 29.0598 | 0 | 30.29 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0312
2008-08-05 16:13 | 0 | 85 | 85 | 0 | 39.95 | 40 | 102.5812 | 0 | 29.0941 | 28.6649 | 0 |  | 0.92751 | 30.89 | 29.072 | 0 | 30.38 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8229 | 38.0312
2008-08-05 16:14 | 0 | 85.03 | 85 | 0 | 39.94 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.39 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 16:15 | 0 | 84.92 | 85 | 0 | 39.88 | 40 | 102.6317 | 0 | 29.0799 | 28.6659 | 0 |  | 0.928 | 30.89 | 29.072 | 0 | 30.38 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 16:16 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.33 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8826 | 38.0312
2008-08-05 16:17 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 102.7327 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.29 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8805 | 38.0312
2008-08-05 16:18 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 102.6317 | 0 | 29.0941 | 28.6941 | 0 |  | 0.928 | 30.89 | 29.0842 | 0 | 30.25 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8663 | 38.0312
2008-08-05 16:19 | 0 | 85.02 | 85 | 0 | 39.92 | 40 | 102.6317 | 0 | 29.079 | 28.6518 | 0 |  | 0.928 | 30.89 | 29.072 | 0 | 30.27 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8826 | 38.0313
2008-08-05 16:20 | 0 | 84.94 | 85 | 0 | 39.94 | 40 | 102.6317 | 0 | 29.079 | 28.6649 | 0 |  | 0.928 | 30.89 | 29.072 | 0 | 30.35 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8826 | 38.0312
2008-08-05 16:21 | 0 | 84.97 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.0931 | 28.6659 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.45 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.836 | 38.0312
2008-08-05 16:22 | 0 | 84.94 | 85 | 0 | 40.03 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.5 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 16:23 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.5 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 16:24 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.47 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 16:25 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 102.6317 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.46 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 16:26 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 102.7327 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.41 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 16:27 | 0 | 85 | 85 | 0 | 40.01 | 40 | 102.5812 | 0 | 29.1214 | 28.6508 | 0 |  | 0.92751 | 30.89 | 29.072 | 0 | 30.36 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0313
2008-08-05 16:28 | 0 | 85 | 85 | 0 | 39.98 | 40 | 102.7327 | 0 | 29.079 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.31 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 16:29 | 0 | 85 | 85 | 0 | 39.93 | 40 | 102.6317 | 0 | 29.0931 | 28.6508 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.28 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0313
2008-08-05 16:30 | 0 | 84.97 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.079 | 28.6508 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.29 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8694 | 38.0313
2008-08-05 16:31 | 0 | 84.97 | 85 | 0 | 40.05 | 40 | 102.6317 | 0 | 29.0931 | 28.6508 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.31 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8239 | 38.0313
2008-08-05 16:32 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.35 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 16:33 | 0 | 85.02 | 85 | 0 | 39.94 | 40 | 102.6317 | 0 | 29.0931 | 28.6508 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.41 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8532 | 38.0313
2008-08-05 16:34 | 0 | 85 | 85 | 0 | 39.93 | 40 | 102.6317 | 0 | 29.1214 | 28.6508 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.42 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.837 | 38.0313
2008-08-05 16:35 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 102.6317 | 0 | 29.1214 | 28.6508 | 0 |  | 0.92797 | 30.89 | 29.0842 | 0 | 30.41 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8229 | 38.0313
2008-08-05 16:36 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 102.6317 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.37 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 16:37 | 0 | 85.02 | 85 | 0 | 40.05 | 40 | 102.6317 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.35 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0312
2008-08-05 16:38 | 0 | 85 | 85 | 0 | 39.92 | 40 | 102.6317 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.32 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8632 | 38.0312
2008-08-05 16:39 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 102.7327 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.28 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8826 | 38.0312
2008-08-05 16:40 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.24 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 16:41 | 0 | 84.96 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.17 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8795 | 38.0312
2008-08-05 16:42 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 102.6317 | 0 | 29.0799 | 28.6941 | 0 |  | 0.928 | 30.89 | 29.0842 | 0 | 30.14 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8947 | 38.0312
2008-08-05 16:43 | 0 | 84.97 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.0799 | 28.6941 | 0 |  | 0.928 | 30.89 | 29.0842 | 0 | 30.21 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 16:44 | 0 | 84.97 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.32 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8663 | 38.0312
2008-08-05 16:45 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 102.6317 | 0 | 29.0941 | 28.6941 | 0 |  | 0.928 | 30.89 | 29.072 | 0 | 30.41 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.836 | 38.0312
2008-08-05 16:46 | 0 | 85 | 85 | 0 | 39.92 | 40 | 102.7327 | 0 | 29.0941 | 28.6941 | 0 |  | 0.92891 | 30.89 | 29.0598 | 0 | 30.43 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8795 | 38.0312
2008-08-05 16:47 | 0 | 84.91 | 85 | 0 | 40.07 | 40 | 102.7327 | 0 | 29.0941 | 28.6659 | 0 |  | 0.92891 | 30.89 | 29.072 | 0 | 30.4 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8501 | 38.0312
2008-08-05 16:48 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 102.7327 | 0 | 29.1214 | 28.6931 | 0 |  | 0.92888 | 30.89 | 29.0965 | 0 | 30.36 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8805 | 38.0312
2008-08-05 16:49 | 0 | 85 | 85 | 0 | 39.92 | 40 | 102.5812 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92751 | 30.89 | 29.072 | 0 | 30.35 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8239 | 38.0312
2008-08-05 16:50 | 0 | 84.96 | 85 | 0 | 39.84 | 40 | 102.7327 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0598 | 0 | 30.31 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 16:51 | 0 | 85 | 85 | 0 | 39.84 | 40 | 102.7327 | 0 | 29.079 | 28.6508 | 0 |  | 0.92888 | 30.89 | 29.0598 | 0 | 30.28 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8694 | 38.0313
2008-08-05 16:52 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 102.5812 | 0 | 29.079 | 28.6649 | 0 |  | 0.92751 | 30.89 | 29.0598 | 0 | 30.27 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8239 | 38.0312
2008-08-05 16:53 | 0 | 84.95 | 85 | 0 | 39.9 | 40 | 102.5812 | 0 | 29.079 | 28.6649 | 0 |  | 0.92751 | 30.89 | 29.0598 | 0 | 30.28 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 16:54 | 0 | 85.03 | 85 | 0 | 39.97 | 40 | 102.5812 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92751 | 30.89 | 29.0598 | 0 | 30.4 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8936 | 38.0312
2008-08-05 16:55 | 0 | 85 | 85 | 0 | 39.91 | 40 | 102.7327 | 0 | 29.1214 | 28.6508 | 0 |  | 0.92888 | 30.89 | 29.0598 | 0 | 30.51 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8694 | 38.0313
2008-08-05 16:56 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 102.7327 | 0 | 29.1214 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.0842 | 0 | 30.53 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8239 | 38.0312
2008-08-05 16:57 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.0842 | 0 | 30.5 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 16:58 | 0 | 85.02 | 85 | 0 | 39.93 | 40 | 102.5812 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92751 | 30.89 | 29.072 | 0 | 30.46 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8805 | 38.0312
2008-08-05 16:59 | 0 | 85 | 85 | 0 | 40 | 40 | 102.7327 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92888 | 30.89 | 29.072 | 0 | 30.39 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8694 | 38.0312
2008-08-05 17:00 | 0 | 85 | 85 | 0 | 40 | 40 | 102.6317 | 0 | 29.079 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.37 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0312
2008-08-05 17:01 | 0 | 85 | 85 | 0 | 39.91 | 40 | 102.6317 | 0 | 29.079 | 28.6508 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.34 | 40.8061 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8542 | 38.0313
2008-08-05 17:02 | 0 | 84.94 | 85 | 0 | 39.97 | 40 | 102.6317 | 0 | 29.0931 | 28.6931 | 0 |  | 0.92797 | 30.89 | 29.0598 | 0 | 30.33 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 17:03 | 0 | 85 | 85 | 0 | 39.97 | 40 | 102.5812 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92751 | 30.89 | 29.0598 | 0 | 30.31 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9817 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8694 | 38.0312
2008-08-05 17:04 | 0 | 84.99 | 85 | 0 | 39.91 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.29 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.9744 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 17:05 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 102.6317 | 0 | 29.0931 | 28.6649 | 0 |  | 0.92797 | 30.89 | 29.072 | 0 | 30.27 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8674 | 38.0312
2008-08-05 17:06 | 0 | 84.97 | 85 | 0 | 40 | 40 | 102.7327 | 0 | 29.0941 | 28.6931 | 0 |  | 0.92891 | 30.89 | 29.0842 | 0 | 30.29 | 40.806 | 0 | 1.7038 | 15 | 15 |  | 14.989 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 15.0291 | -1 | 15.0309 | 30.8532 | 38.0312
2008-08-05 17:07 | 0 | 84.63 | 85 | 0 | 40 | 40 | 99.096 | 0 | 3.9721 | 3.7869 | 0 |  | 0.90388 | 4.19 | 3.9805 | 0 | 31.14 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 22.4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2565 | -1 | 25.0042 | 4.2188 | 0
2008-08-05 17:08 | 0 | 85.05 | 85 | 0 | 40 | 40 | 99.0454 | 0 | 3.9898 | 3.807 | 0 |  | 0.90821 | 4.19 | 4.0171 | 0 | 31.83 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 20.84 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2565 | -1 | 25.0042 | 4.2238 | 0
2008-08-05 17:09 | 0 | 84.92 | 85 | 0 | 40.03 | 40 | 99.0454 | 0 | 4.0067 | 3.8223 | 0 |  | 0.91231 | 4.19 | 4.0293 | 0 | 32.48 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.1795 | 0 | 0 | 0 | 4 | 19.52 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2565 | -1 | 25.0042 | 4.2168 | 0
2008-08-05 17:10 | 0 | 84.96 | 85 | 0 | 39.84 | 40 | 99.0454 | 0 | 4.0215 | 3.8378 | 0 |  | 0.91594 | 4.19 | 4.0415 | 0 | 32.98 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 18.36 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2565 | -1 | 25.0042 | 4.2126 | 0
2008-08-05 17:11 | 0 | 84.91 | 85 | 0 | 39.93 | 40 | 99.0454 | 0 | 4.0352 | 3.8517 | 0 |  | 0.91931 | 4.19 | 4.0537 | 0 | 33.41 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 17.29 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2566 | -1 | 25.0042 | 4.216 | 0
2008-08-05 17:12 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 98.9949 | 0 | 4.0456 | 3.8626 | 0 |  | 0.92192 | 4.19 | 4.0659 | 0 | 33.71 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 16.32 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2566 | -1 | 25.0042 | 4.2121 | 0
2008-08-05 17:13 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.92706 | 0 | 0.51282 | 0 | 33.97 | 0 | 0 | 0 | 0 | 0 |  | 0.043956 | 0 | 0 | 0 | 4 | 15.45 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.56242 | 0
2008-08-05 17:14 | 0 | 84.94 | 85 | 0 | 40.01 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.92977 | 0 | 0.4884 | 0 | 34.2 | 0 | 0 | 0 | 0 | 0 |  | 0.029304 | 0 | 0 | 0 | 4 | 14.61 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.48754 | 0
2008-08-05 17:15 | 0 | 84.96 | 85 | 0 | 40 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.93226 | 0 | 0.46398 | 0 | 34.43 | 0 | 0 | 0 | 0 | 0 |  | 0.021978 | 0 | 0 | 0 | 4 | 13.84 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.46215 | 0
2008-08-05 17:16 | 0 | 85.03 | 85 | 0 | 39.93 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.93461 | 0 | 0.47619 | 0 | 34.71 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 13.1 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.47306 | 0
2008-08-05 17:17 | 0 | 85.05 | 85 | 0 | 39.9 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.93683 | 0 | 0.46398 | 0 | 34.9 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 12.43 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45973 | 0
2008-08-05 17:18 | 0 | 85 | 85 | 0 | 40.03 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.93894 | 0 | 0.46398 | 0 | 35 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 11.79 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45864 | 0
2008-08-05 17:19 | 0 | 85.02 | 85 | 0 | 39.93 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.94092 | 0 | 0.46398 | 0 | 35 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 11.19 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45785 | 0
2008-08-05 17:20 | 0 | 85.08 | 85 | 0 | 40.04 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.94281 | 0 | 0.46398 | 0 | 35.03 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 10.62 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44485 | 0
2008-08-05 17:21 | 0 | 85.17 | 85 | 0 | 40.03 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.9446 | 0 | 0.46398 | 0 | 35.01 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 10.08 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44415 | 0
2008-08-05 17:22 | 0 | 85 | 85 | 0 | 40.01 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.94634 | 0 | 0.45177 | 0 | 35.03 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 9.56 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44305 | 0
2008-08-05 17:23 | 0 | 85.09 | 85 | 0 | 40 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.94798 | 0 | 0.46398 | 0 | 35 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 9.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44266 | 0
2008-08-05 17:24 | 0 | 85.03 | 85 | 0 | 40.05 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.94957 | 0 | 0.46398 | 0 | 34.98 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 8.6 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44186 | 0
2008-08-05 17:25 | 0 | 85.05 | 85 | 0 | 39.96 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.95105 | 0 | 0.45177 | 0 | 34.92 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 8.16 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44088 | 0
2008-08-05 17:26 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.95254 | 0 | 0.45177 | 0 | 34.91 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 7.73 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44018 | 0
2008-08-05 17:27 | 0 | 85.03 | 85 | 0 | 40.04 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.95393 | 0 | 0.46398 | 0 | 34.89 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 7.32 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45137 | 0
2008-08-05 17:28 | 0 | 84.99 | 85 | 0 | 40 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.9553 | 0 | 0.46398 | 0 | 34.83 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 6.92 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.439 | 0
2008-08-05 17:29 | 0 | 85.05 | 85 | 0 | 40.05 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.95577 | 0 | 0.46398 | 0 | 34.85 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 6.77 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45078 | 0
2008-08-05 17:30 | 0 | 85.05 | 85 | 0 | 40.05 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.95577 | 0 | 0.46398 | 0 | 34.85 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 6.77 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45078 | 0
2008-08-05 17:31 | 0 | 85.05 | 85 | 0 | 40.05 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.95577 | 0 | 0.46398 | 0 | 34.85 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 6.77 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45078 | 0
2008-08-05 17:33 | 0 | 85.02 | 85 | 0 | 40.01 | 40 | 99.197 | 0 | 0 | 0 | 0 |  | 0.96026 | 0 | 0.46398 | 0 | 34.85 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 5.33 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44849 | 0
2008-08-05 17:34 | 0 | 85.03 | 85 | 0 | 40.03 | 40 | 99.197 | 0 | 0 | 0 | 0 |  | 0.96133 | 0 | 0.46398 | 0 | 34.78 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 5.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44773 | 0
2008-08-05 17:35 | 0 | 85.03 | 85 | 0 | 40.07 | 40 | 99.197 | 0 | 0 | 0 | 0 |  | 0.96282 | 0 | 0.45177 | 0 | 34.75 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4.72 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44746 | 0
2008-08-05 17:36 | 0 | 85 | 85 | 0 | 40.03 | 40 | 99.197 | 0 | 0 | 0 | 0 |  | 0.96331 | 0 | 0.46398 | 0 | 34.69 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4.44 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44698 | 0
2008-08-05 17:37 | 0 | 85.03 | 85 | 0 | 40.04 | 40 | 99.197 | 0 | 0 | 0 | 0 |  | 0.96424 | 0 | 0.46398 | 0 | 34.59 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4.17 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45837 | 0
2008-08-05 17:38 | 0 | 85 | 85 | 0 | 40.03 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.96536 | 0 | 0.46398 | 0 | 34.57 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44627 | 0
2008-08-05 17:39 | 0 | 84.89 | 85 | 0 | 39.81 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.96501 | 0 | 0.46398 | 0 | 34.46 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.4462 | 0
2008-08-05 17:40 | 0 | 84.96 | 85 | 0 | 40.07 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.96553 | 0 | 0.46398 | 0 | 34.31 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44617 | 0
2008-08-05 17:41 | 0 | 85 | 85 | 0 | 40.07 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.96504 | 0 | 0.46398 | 0 | 34.27 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44593 | 0
2008-08-05 17:42 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.96553 | 0 | 0.46398 | 0 | 34.27 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44593 | 0
2008-08-05 17:43 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.96655 | 0 | 0.47619 | 0 | 34.27 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.94 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44593 | 0
2008-08-05 17:44 | 0 | 85 | 85 | 0 | 39.96 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.96553 | 0 | 0.47619 | 0 | 34.25 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45766 | 0
2008-08-05 17:45 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 99.3485 | 0 | 0 | 0 | 0 |  | 0.96606 | 0 | 0.46398 | 0 | 34.14 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.94 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45766 | 0
2008-08-05 17:46 | 0 | 84.95 | 85 | 0 | 40 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.96557 | 0 | 0.47619 | 0 | 34.04 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.94 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45791 | 0
2008-08-05 17:47 | 0 | 84.95 | 85 | 0 | 40 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.96603 | 0 | 0.46398 | 0 | 33.93 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44617 | 0
2008-08-05 17:48 | 0 | 85 | 85 | 0 | 39.95 | 40 | 99.298 | 0 | 0 | 0 | 0 |  | 0.96504 | 0 | 0.47619 | 0 | 33.88 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45745 | 0
2008-08-05 17:49 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.96553 | 0 | 0.46398 | 0 | 33.78 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44572 | 0
2008-08-05 17:50 | 0 | 84.96 | 85 | 0 | 39.91 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.96603 | 0 | 0.47619 | 0 | 33.66 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44594 | 0
2008-08-05 17:51 | 0 | 85.03 | 85 | 0 | 39.9 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.96603 | 0 | 0.47619 | 0 | 33.59 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.4577 | 0
2008-08-05 17:52 | 0 | 84.95 | 85 | 0 | 40.05 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.9655 | 0 | 0.46398 | 0 | 33.59 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45746 | 0
2008-08-05 17:53 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.96599 | 0 | 0.46398 | 0 | 33.59 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45723 | 0
2008-08-05 17:54 | 0 | 84.94 | 85 | 0 | 39.9 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.9655 | 0 | 0.47619 | 0 | 33.59 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44596 | 0
2008-08-05 17:55 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 99.2475 | 0 | 0 | 0 | 0 |  | 0.9655 | 0 | 0.47619 | 0 | 33.57 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.44596 | 0
2008-08-05 17:56 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 99.298 | 0 | 0 | 0 | 0 |  | 0.96546 | 0 | 0.47619 | 0 | 33.55 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 3.97 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.45771 | 0
2008-08-05 17:57 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 98.9949 | 0 | 4.2154 | 4.035 | 0 |  | 0.9635 | 4.19 | 4.2369 | 0 | 33.52 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2234 | 0 | 0 | 0 | 4 | 3.97 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.1998 | 0
2008-08-05 17:58 | 0 | 85 | 85 | 0 | 40.01 | 40 | 98.9949 | 0 | 4.2154 | 4.0371 | 0 |  | 0.96301 | 4.19 | 4.2369 | 0 | 33.52 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.97 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.1999 | 0
2008-08-05 17:59 | 0 | 84.87 | 85 | 0 | 39.86 | 40 | 99.0454 | 0 | 4.2154 | 4.0371 | 0 |  | 0.96301 | 4.19 | 4.2369 | 0 | 33.52 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.97 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.213 | 0
2008-08-05 18:00 | 0 | 84.92 | 85 | 0 | 39.95 | 40 | 98.9949 | 0 | 4.2154 | 4.0371 | 0 |  | 0.9635 | 4.19 | 4.2247 | 0 | 33.46 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2308 | 0 | 0 | 0 | 4 | 3.97 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.213 | 0
2008-08-05 18:01 | 0 | 84.92 | 85 | 0 | 39.94 | 40 | 99.096 | 0 | 4.2132 | 4.0369 | 0 |  | 0.96346 | 4.19 | 4.2247 | 0 | 33.41 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2151 | 0
2008-08-05 18:02 | 0 | 84.92 | 85 | 0 | 39.84 | 40 | 99.0454 | 0 | 4.2132 | 4.0371 | 0 |  | 0.96301 | 4.19 | 4.2369 | 0 | 33.33 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.97 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2131 | 0
2008-08-05 18:03 | 0 | 84.88 | 85 | 0 | 39.92 | 40 | 99.0454 | 0 | 4.2132 | 4.0349 | 0 |  | 0.96346 | 4.19 | 4.2369 | 0 | 33.31 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2153 | 0
2008-08-05 18:04 | 0 | 84.94 | 85 | 0 | 39.97 | 40 | 98.9949 | 0 | 4.2152 | 4.0349 | 0 |  | 0.96297 | 4.19 | 4.2247 | 0 | 33.31 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2153 | 0
2008-08-05 18:05 | 0 | 84.94 | 85 | 0 | 39.97 | 40 | 98.9949 | 0 | 4.2172 | 4.0369 | 0 |  | 0.96346 | 4.19 | 4.2369 | 0 | 33.29 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2131 | 0
2008-08-05 18:06 | 0 | 84.92 | 85 | 0 | 39.91 | 40 | 98.9949 | 0 | 4.2132 | 4.0369 | 0 |  | 0.96297 | 4.19 | 4.2369 | 0 | 33.26 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2022 | 0
2008-08-05 18:07 | 0 | 84.87 | 85 | 0 | 39.94 | 40 | 98.9949 | 0 | 4.2152 | 4.0349 | 0 |  | 0.96346 | 4.19 | 4.2369 | 0 | 33.24 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2022 | 0
2008-08-05 18:08 | 0 | 85 | 85 | 0 | 39.86 | 40 | 98.9949 | 0 | 4.2152 | 4.0369 | 0 |  | 0.96297 | 4.19 | 4.2369 | 0 | 33.23 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2131 | 0
2008-08-05 18:09 | 0 | 84.91 | 85 | 0 | 39.87 | 40 | 99.0454 | 0 | 4.2132 | 4.039 | 0 |  | 0.96346 | 4.19 | 4.2369 | 0 | 33.22 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2015 | 0 | 0 | 0 | 4 | 3.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2153 | 0
2008-08-05 18:10 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 99.0454 | 0 | 4.2151 | 4.0388 | 0 |  | 0.96343 | 4.19 | 4.2369 | 0 | 33.22 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2161 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.1979 | 0
2008-08-05 18:11 | 0 | 84.91 | 85 | 0 | 39.8 | 40 | 99.0454 | 0 | 4.2132 | 4.0349 | 0 |  | 0.96297 | 4.19 | 4.2369 | 0 | 33.19 | 0 | 0 | 0 | 25 | 25.2525 |  | 25.2088 | 0 | 0 | 0 | 4 | 3.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 25.2567 | -1 | 25.0042 | 4.2022 | 0

## Ref_Raw
DATETIME | CALIB_AUXOUTPUTSELECT | CALIB_CALIBAIRPREHEATER_PV | CALIB_CALIBAIRPREHEATER_SP | CALIB_CALIBFAULTLATCH | CALIB_CALIBHEATEDPLATE_PV | CALIB_CALIBHEATEDPLATE_SP | CALIB_CALPRESSURE | CALIB_CALPRESSURESENSORFAIL | CALIB_CALSOURCEACTUAL | CALIB_CALSOURCEADJUST | CALIB_CALSOURCECOMPARE | CALIB_CALSOURCECORRECTION | CALIB_CALSOURCEFACTOR | CALIB_CALSOURCEFLOW | CALIB_CALSOURCEREADBACKRAW | CALIB_CALSOURCESTABLE | CALIB_CASEPV | CALIB_CONCENTRATION | CALIB_DILUTECOMPARE | CALIB_DILUTECONCENTRATION | CALIB_DILUTEFLOW | CALIB_DILUTEFLOWACTUAL | CALIB_DILUTEFLOWCORRECTION | CALIB_DILUTEFLOWREADBACK | CALIB_DILUTIONGAS | CALIB_INJECTSELECT | CALIB_LED | CALIB_MERCURYSOURCESP | CALIB_MERCURYSOURCETEMP | CALIB_NISTHGTABLE | CALIB_PLCSTATUS | CALIB_REMOTECONTROL | CALIB_SOURCEACTIVATE | CALIB_SOURCEHEATERENABLE | CALIB_STATTEMP | CALIB_STATTEMPLOGIC | CALIB_STATUS | CALIB_STATUSER1 | CALIB_STATUSER2 | CALIB_STATUSER3 | CALIB_TOTALFLOW | CALIB_TOTALFLOWACTUAL | CALIB_TOTALFLOWALTERNATE | CALIB_TOTALFLOWX | CALIB_CALSOURCEREADBACK | CALIB_CONCEPA
2008-08-04 14:39 | 0 | 85.03 | 85 | 0 | 40.02 | 40 | 100.7321 | 0 | 0 | 0 | 0 |  | 0.9798 | 0 | 0.18315 | 0 | 20.25 | 0 | 0 | 0 | 0 | 0 |  | 0.24176 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17712 | 0
2008-08-04 14:40 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 100.7321 | 0 | 0 | 0 | 0 |  | 0.9798 | 0 | 0.18315 | 0 | 20.12 | 0 | 0 | 0 | 0 | 0 |  | 0.23443 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17721 | 0
2008-08-04 14:41 | 0 | 84.9 | 85 | 0 | 39.96 | 40 | 100.7321 | 0 | 0 | 0 | 0 |  | 0.9798 | 0 | 0.18315 | 0 | 19.93 | 0 | 0 | 0 | 0 | 0 |  | 0.24176 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17712 | 0
2008-08-04 14:42 | 0 | 84.89 | 85 | 0 | 39.99 | 40 | 100.7826 | 0 | 0 | 0 | 0 |  | 0.98029 | 0 | 0.18315 | 0 | 19.92 | 0 | 0 | 0 | 0 | 0 |  | 0.24176 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17721 | 0
2008-08-04 14:43 | 0 | 84.87 | 85 | 0 | 39.94 | 40 | 107.3995 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 19.7 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5195 | 1.5004
2008-08-04 14:44 | 0 | 84.44 | 85 | 0 | 39.98 | 40 | 107.551 | 0 | 13.4137 | 13.1394 | 0 |  | 1.0461 | 12.56 | 13.3822 | 0 | 19.42 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5304 | 1.5004
2008-08-04 14:45 | 0 | 84.41 | 85 | 0 | 40.02 | 40 | 107.551 | 0 | 13.4141 | 13.1394 | 0 |  | 1.0461 | 12.56 | 13.3944 | 0 | 19.45 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5304 | 1.5004
2008-08-04 14:46 | 0 | 84.47 | 85 | 0 | 39.94 | 40 | 107.45 | 0 | 13.4011 | 13.127 | 0 |  | 1.0451 | 12.56 | 13.37 | 0 | 19.48 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5307 | 1.5004
2008-08-04 14:47 | 0 | 84.58 | 85 | 0 | 39.93 | 40 | 107.45 | 0 | 13.4011 | 13.127 | 0 |  | 1.0451 | 12.56 | 13.3822 | 0 | 19.65 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5422 | 1.5004
2008-08-04 14:48 | 0 | 84.66 | 85 | 0 | 39.92 | 40 | 107.551 | 0 | 13.4011 | 13.127 | 0 |  | 1.0451 | 12.56 | 13.3822 | 0 | 20.35 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5422 | 1.5004
2008-08-04 14:49 | 0 | 84.76 | 85 | 0 | 39.95 | 40 | 107.45 | 0 | 13.4011 | 13.127 | 0 |  | 1.0451 | 12.56 | 13.3822 | 0 | 21.43 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5422 | 1.5004
2008-08-04 14:50 | 0 | 84.8 | 85 | 0 | 39.98 | 40 | 107.551 | 0 | 13.4006 | 13.1265 | 0 |  | 1.0451 | 12.56 | 13.3822 | 0 | 21.69 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5422 | 1.5004
2008-08-04 14:51 | 0 | 84.84 | 85 | 0 | 39.86 | 40 | 107.45 | 0 | 13.4006 | 13.1265 | 0 |  | 1.0451 | 12.56 | 13.3822 | 0 | 21.27 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5426 | 1.5004
2008-08-04 14:52 | 0 | 84.84 | 85 | 0 | 39.94 | 40 | 107.45 | 0 | 13.4006 | 13.1265 | 0 |  | 1.0451 | 12.56 | 13.3944 | 0 | 20.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5426 | 1.5004
2008-08-04 14:53 | 0 | 84.91 | 85 | 0 | 39.97 | 40 | 107.45 | 0 | 13.4006 | 13.1265 | 0 |  | 1.0451 | 12.56 | 13.3822 | 0 | 20.41 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5431 | 1.5004
2008-08-04 14:54 | 0 | 84.91 | 85 | 0 | 40.04 | 40 | 107.45 | 0 | 13.4001 | 13.1261 | 0 |  | 1.0451 | 12.56 | 13.37 | 0 | 19.89 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5431 | 1.5004
2008-08-04 14:55 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 107.45 | 0 | 13.4001 | 13.1261 | 0 |  | 1.0451 | 12.56 | 13.37 | 0 | 19.8 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5431 | 1.5004
2008-08-04 14:56 | 0 | 84.96 | 85 | 0 | 39.9 | 40 | 107.45 | 0 | 13.4001 | 13.1261 | 0 |  | 1.0451 | 12.56 | 13.37 | 0 | 19.58 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5311 | 1.5004
2008-08-04 14:57 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 107.45 | 0 | 13.4006 | 13.1261 | 0 |  | 1.0451 | 12.56 | 13.3822 | 0 | 19.53 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5311 | 1.5004
2008-08-04 14:58 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 107.45 | 0 | 13.4006 | 13.1265 | 0 |  | 1.0451 | 12.56 | 13.37 | 0 | 19.5 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5316 | 1.5004
2008-08-04 14:59 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 107.45 | 0 | 13.4006 | 13.1265 | 0 |  | 1.0451 | 12.56 | 13.3822 | 0 | 20.22 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5311 | 1.5004
2008-08-04 15:00 | 0 | 85.03 | 85 | 0 | 39.86 | 40 | 107.45 | 0 | 13.4001 | 13.1261 | 0 |  | 1.0451 | 12.56 | 13.37 | 0 | 20.37 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5311 | 1.5004
2008-08-04 15:01 | 0 | 84.98 | 85 | 0 | 40 | 40 | 107.45 | 0 | 13.4001 | 13.1261 | 0 |  | 1.0451 | 12.56 | 13.37 | 0 | 20.35 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5426 | 1.5004
2008-08-04 15:02 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 107.45 | 0 | 13.4006 | 13.1265 | 0 |  | 1.0451 | 12.56 | 13.3822 | 0 | 20.21 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5458 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:03 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 107.45 | 0 | 13.4001 | 13.1261 | 0 |  | 1.0451 | 12.56 | 13.3822 | 0 | 20.29 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5431 | 1.5004
2008-08-04 15:04 | 0 | 84.98 | 85 | 0 | 39.9 | 40 | 107.3995 | 0 | 13.4001 | 13.1261 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 20.39 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5316 | 1.5004
2008-08-04 15:05 | 0 | 85.01 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.4001 | 13.1199 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 20.44 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:06 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 107.45 | 0 | 13.3939 | 13.1199 | 0 |  | 1.0446 | 12.56 | 13.3822 | 0 | 20.37 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5311 | 1.5004
2008-08-04 15:07 | 0 | 84.95 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.4001 | 13.1265 | 0 |  | 1.0446 | 12.56 | 13.3822 | 0 | 20.37 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5196 | 1.5004
2008-08-04 15:08 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3944 | 13.1265 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 20.37 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5311 | 1.5004
2008-08-04 15:09 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 107.3995 | 0 | 13.3944 | 13.1265 | 0 |  | 1.0451 | 12.56 | 13.37 | 0 | 20.39 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5375 | 1.5004
2008-08-04 15:10 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3944 | 13.1204 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 20.66 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5375 | 1.5004
2008-08-04 15:11 | 0 | 84.97 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.3944 | 13.1204 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 21.34 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:12 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.4001 | 13.1204 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 21.43 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:13 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3939 | 13.1199 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 21.14 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5311 | 1.5004
2008-08-04 15:14 | 0 | 85.03 | 85 | 0 | 39.92 | 40 | 107.3995 | 0 | 13.3944 | 13.1199 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 20.6 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:15 | 0 | 85.02 | 85 | 0 | 39.95 | 40 | 107.3995 | 0 | 13.3944 | 13.1204 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 20.29 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:16 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3939 | 13.1199 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 19.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5375 | 1.5004
2008-08-04 15:17 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 107.45 | 0 | 13.3939 | 13.1199 | 0 |  | 1.0451 | 12.56 | 13.37 | 0 | 19.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5429 | 1.5004
2008-08-04 15:18 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 107.3995 | 0 | 13.3944 | 13.1204 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 19.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:19 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 107.3995 | 0 | 13.3944 | 13.1204 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 20.14 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:20 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3944 | 13.1204 | 0 |  | 1.0446 | 12.56 | 13.3578 | 0 | 20.66 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5255 | 1.5004
2008-08-04 15:21 | 0 | 84.98 | 85 | 0 | 39.92 | 40 | 107.3995 | 0 | 13.3944 | 13.1204 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 21.22 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:22 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 107.3995 | 0 | 13.3944 | 13.1204 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 21.24 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5255 | 1.5004
2008-08-04 15:23 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3944 | 13.1204 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 21.16 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:24 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3944 | 13.1204 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 20.54 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5458 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:25 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 13.3881 | 13.1142 | 0 |  | 1.0441 | 12.56 | 13.3578 | 0 | 20.62 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5314 | 1.5004
2008-08-04 15:26 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3944 | 13.1208 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 20.22 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5458 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:27 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 107.3995 | 0 | 13.3944 | 13.1204 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 20.12 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:28 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3944 | 13.1204 | 0 |  | 1.0446 | 12.56 | 13.37 | 0 | 20.01 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.537 | 1.5004
2008-08-04 15:29 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3944 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3822 | 0 | 19.82 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 15:30 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3944 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 20.01 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 15:31 | 0 | 85.01 | 85 | 0 | 39.93 | 40 | 107.3995 | 0 | 13.3944 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 20.04 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 15:32 | 0 | 85.01 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 20.22 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 15:33 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 20.85 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 15:34 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 21.03 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 15:35 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 20.64 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 15:36 | 0 | 84.95 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 20.58 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 15:37 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 20.28 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 15:38 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 19.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5458 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5429 | 1.5004
2008-08-04 15:39 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 19.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 15:40 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 107.3489 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 15:41 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 19.82 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 15:42 | 0 | 84.96 | 85 | 0 | 39.92 | 40 | 107.3995 | 0 | 13.3886 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 19.73 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 15:43 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 19.76 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 15:44 | 0 | 85.02 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3886 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 20.02 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 15:45 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 20.67 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 15:46 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 107.3995 | 0 | 13.3953 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 21.05 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 15:47 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 21.23 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5165 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 15:48 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 20.77 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 15:49 | 0 | 85.46 | 85 | 0 | 39.94 | 40 | 100.5805 | 0 | 0 | 0 | 0 |  | 0.97833 | 0 | 0.45177 | 0 | 20.65 | 0 | 0 | 0 | 0 | 0 |  | 0.28571 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.42593 | 0
2008-08-04 15:50 | 0 | 85.45 | 85 | 0 | 40.02 | 40 | 100.5805 | 0 | 0 | 0 | 0 |  | 0.97833 | 0 | 0.26862 | 0 | 20.3 | 0 | 0 | 0 | 0 | 0 |  | 0.30037 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.28396 | 0
2008-08-04 15:51 | 0 | 85.46 | 85 | 0 | 39.99 | 40 | 100.53 | 0 | 4.2651 | 4.0971 | 1 |  | 0.97784 | 4.19 | 0.29304 | 0 | 20.05 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.4579 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 0.23663 | 0
2008-08-04 15:52 | 0 | 84.87 | 85 | 0 | 40 | 40 | 100.53 | 0 | 4.2672 | 4.0992 | 0 |  | 0.97833 | 4.19 | 4.2613 | 0 | 20.06 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.184 | 0
2008-08-04 15:53 | 0 | 84.8 | 85 | 0 | 39.98 | 40 | 100.53 | 0 | 4.2674 | 4.0973 | 0 |  | 0.97787 | 4.19 | 4.2613 | 0 | 19.87 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1715 | 0
2008-08-04 15:54 | 0 | 84.74 | 85 | 0 | 39.99 | 40 | 100.5805 | 0 | 4.2653 | 4.0993 | 0 |  | 0.97787 | 4.19 | 4.2613 | 0 | 19.94 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1736 | 0
2008-08-04 15:55 | 0 | 84.79 | 85 | 0 | 39.95 | 40 | 100.5805 | 0 | 4.2674 | 4.0993 | 0 |  | 0.97833 | 4.19 | 4.2613 | 0 | 20.04 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1715 | 0
2008-08-04 15:56 | 0 | 84.83 | 85 | 0 | 39.95 | 40 | 100.53 | 0 | 4.2674 | 4.0993 | 0 |  | 0.97787 | 4.19 | 4.2613 | 0 | 20.85 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1736 | 0
2008-08-04 15:57 | 0 | 84.89 | 85 | 0 | 39.95 | 40 | 100.53 | 0 | 4.2674 | 4.0993 | 0 |  | 0.97787 | 4.19 | 4.2613 | 0 | 21.82 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1838 | 0
2008-08-04 15:58 | 0 | 84.89 | 85 | 0 | 39.95 | 40 | 100.5805 | 0 | 4.2653 | 4.0993 | 0 |  | 0.97787 | 4.19 | 4.2613 | 0 | 21.75 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1838 | 0
2008-08-04 15:59 | 0 | 84.9 | 85 | 0 | 39.98 | 40 | 100.53 | 0 | 4.2653 | 4.0993 | 0 |  | 0.97787 | 4.19 | 4.2613 | 0 | 21.11 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1838 | 0
2008-08-04 16:00 | 0 | 84.91 | 85 | 0 | 39.99 | 40 | 100.5805 | 0 | 4.2674 | 4.0993 | 0 |  | 0.97787 | 4.19 | 4.2613 | 0 | 20.73 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1737 | 0
2008-08-04 16:01 | 0 | 84.9 | 85 | 0 | 39.96 | 40 | 100.53 | 0 | 4.2651 | 4.0973 | 0 |  | 0.97787 | 4.19 | 4.2491 | 0 | 20.45 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1737 | 0
2008-08-04 16:02 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 100.53 | 0 | 4.2651 | 4.0971 | 0 |  | 0.97787 | 4.19 | 4.2613 | 0 | 20.24 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1737 | 0
2008-08-04 16:03 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 100.53 | 0 | 4.2653 | 4.0993 | 0 |  | 0.97787 | 4.19 | 4.2613 | 0 | 20.06 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1736 | 0
2008-08-04 16:04 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 100.53 | 0 | 4.2653 | 4.0973 | 0 |  | 0.97787 | 4.19 | 4.2613 | 0 | 19.96 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1736 | 0
2008-08-04 16:05 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 100.5805 | 0 | 4.2653 | 4.0973 | 0 |  | 0.97787 | 4.19 | 4.2613 | 0 | 20 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1736 | 0
2008-08-04 16:06 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.3879 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.2491 | 0 | 20.26 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5842 | -1 | 25.0042 | 4.1736 | 0.50071
2008-08-04 16:07 | 0 | 84.9 | 85 | 0 | 39.95 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0442 | 4.19 | 4.5299 | 0 | 20.65 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 16:08 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.95 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 16:09 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.84 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 16:10 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 107.2984 | 0 | 4.5459 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.55 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 16:11 | 0 | 84.93 | 85 | 0 | 39.94 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.28 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 16:12 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 107.2984 | 0 | 4.5459 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.13 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 16:13 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.02 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 16:14 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.96 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 16:15 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.92 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 16:16 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.78 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5165 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 16:17 | 0 | 84.93 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.83 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 16:18 | 0 | 85.02 | 85 | 0 | 39.94 | 40 | 107.2984 | 0 | 4.5438 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.84 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 16:19 | 0 | 85.01 | 85 | 0 | 39.94 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.26 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 16:20 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5459 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.92 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 16:21 | 0 | 85.03 | 85 | 0 | 39.92 | 40 | 107.3489 | 0 | 4.5436 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.82 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 16:22 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 4.5457 | 4.375 | 0 |  | 1.0442 | 4.19 | 4.5299 | 0 | 20.6 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5165 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.175 | 0.50071
2008-08-04 16:23 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5436 | 4.375 | 0 |  | 1.0442 | 4.19 | 4.5299 | 0 | 20.21 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1866 | 0.50071
2008-08-04 16:24 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.3489 | 0 | 4.5457 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.175 | 0.50071
2008-08-04 16:25 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5457 | 4.375 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.81 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1866 | 0.50071
2008-08-04 16:26 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 107.2984 | 0 | 4.5436 | 4.375 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.71 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 16:27 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 4.5436 | 4.375 | 0 |  | 1.0442 | 4.19 | 4.5299 | 0 | 19.68 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 16:28 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.53 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 16:29 | 0 | 84.93 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5436 | 4.3752 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.57 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 16:30 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 107.3489 | 0 | 4.5436 | 4.3731 | 0 |  | 1.0442 | 4.19 | 4.5421 | 0 | 19.5 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 16:31 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5457 | 4.375 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.41 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1749 | 0.50071
2008-08-04 16:32 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5436 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.5 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 16:33 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 4.5436 | 4.373 | 0 |  | 1.0442 | 4.19 | 4.5299 | 0 | 20.53 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 16:34 | 0 | 84.98 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5459 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 21.36 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 16:35 | 0 | 85.01 | 85 | 0 | 40.02 | 40 | 107.3489 | 0 | 4.5438 | 4.3752 | 0 |  | 1.0442 | 4.19 | 4.5299 | 0 | 21.47 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1866 | 0.50071
2008-08-04 16:36 | 0 | 85.02 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 21.07 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 16:37 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 107.1974 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.61 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 16:38 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5395 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.2 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 16:39 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5436 | 4.3689 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.97 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 16:40 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.85 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1809 | 0.50071
2008-08-04 16:41 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 107.2984 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.75 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1925 | 0.50071
2008-08-04 16:42 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 107.1974 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.61 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1925 | 0.50071
2008-08-04 16:43 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 107.1974 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 19.57 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1925 | 0.50071
2008-08-04 16:44 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 107.1974 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 19.5 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 16:45 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 107.1974 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.48 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1809 | 0.50071
2008-08-04 16:46 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 107.1974 | 0 | 4.5396 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.38 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 16:47 | 0 | 84.93 | 85 | 0 | 39.98 | 40 | 107.1974 | 0 | 4.5438 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 19.37 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5165 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 16:48 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 107.2984 | 0 | 4.5396 | 4.369 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.33 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 16:49 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 107.1974 | 0 | 4.5396 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 19.32 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1924 | 0.50071
2008-08-04 16:50 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 107.1974 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.36 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5165 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1924 | 0.50071
2008-08-04 16:51 | 0 | 84.91 | 85 | 0 | 39.94 | 40 | 107.1974 | 0 | 4.5396 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 20.45 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1924 | 0.50071
2008-08-04 16:52 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 21.61 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 16:53 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0442 | 4.19 | 4.5299 | 0 | 22.13 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 16:54 | 0 | 84.95 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 21.53 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 16:55 | 0 | 85.01 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.99 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 16:56 | 0 | 85.01 | 85 | 0 | 39.92 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.45 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 16:57 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 107.2984 | 0 | 4.5457 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.14 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.175 | 0.50071
2008-08-04 16:58 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.98 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 16:59 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.84 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 17:00 | 0 | 84.93 | 85 | 0 | 39.94 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.74 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:01 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.81 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 17:02 | 0 | 84.96 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 21.46 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 17:03 | 0 | 84.95 | 85 | 0 | 40.04 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 21.45 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 17:04 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.97 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:05 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.53 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 17:06 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.39 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 17:07 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 107.1974 | 0 | 4.5436 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 20.16 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1809 | 0.50071
2008-08-04 17:08 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 107.1974 | 0 | 4.5395 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.96 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1809 | 0.50071
2008-08-04 17:09 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 107.1974 | 0 | 4.5395 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.8 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1809 | 0.50071
2008-08-04 17:10 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 107.1974 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.81 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:11 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 107.1974 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 19.64 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1925 | 0.50071
2008-08-04 17:12 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 107.1974 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 19.6 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1809 | 0.50071
2008-08-04 17:13 | 0 | 84.91 | 85 | 0 | 39.98 | 40 | 107.1974 | 0 | 4.5436 | 4.373 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.88 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:14 | 0 | 84.95 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5436 | 4.3689 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.65 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1808 | 0.50071
2008-08-04 17:15 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5396 | 4.369 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 21.59 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 17:16 | 0 | 84.96 | 85 | 0 | 39.92 | 40 | 107.2984 | 0 | 4.5396 | 4.369 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 21.96 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1925 | 0.50071
2008-08-04 17:17 | 0 | 85.03 | 85 | 0 | 39.94 | 40 | 107.1974 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 21.47 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1925 | 0.50071
2008-08-04 17:18 | 0 | 85.01 | 85 | 0 | 39.94 | 40 | 107.1974 | 0 | 4.5395 | 4.373 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 20.91 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1809 | 0.50071
2008-08-04 17:19 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 107.2984 | 0 | 4.5436 | 4.3689 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.47 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:20 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 107.1974 | 0 | 4.5436 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 20.2 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1809 | 0.50071
2008-08-04 17:21 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 107.1974 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 20 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1809 | 0.50071
2008-08-04 17:22 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.9 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:23 | 0 | 84.98 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.62 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:24 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.66 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:25 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.57 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:26 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.64 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:27 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.5 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:28 | 0 | 84.98 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.49 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 17:29 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.12 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 17:30 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 21.16 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 17:31 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 107.2984 | 0 | 4.5438 | 4.3752 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 21.48 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 17:32 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0442 | 4.19 | 4.5421 | 0 | 21.2 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 17:33 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.74 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:34 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 107.2984 | 0 | 4.5436 | 4.375 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.37 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.175 | 0.50071
2008-08-04 17:35 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 107.2984 | 0 | 4.5436 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.1 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 17:36 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 107.2984 | 0 | 4.5436 | 4.375 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.86 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:37 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 107.1974 | 0 | 4.5438 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.77 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 17:38 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.66 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1808 | 0.50071
2008-08-04 17:39 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 107.1974 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.63 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1924 | 0.50071
2008-08-04 17:40 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 107.1974 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 19.74 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 17:41 | 0 | 84.93 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 20.75 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 17:42 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 107.1974 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 21.74 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1924 | 0.50071
2008-08-04 17:43 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5436 | 4.3731 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 21.89 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 17:44 | 0 | 85.02 | 85 | 0 | 39.93 | 40 | 107.1974 | 0 | 4.5395 | 4.373 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 21.35 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1925 | 0.50071
2008-08-04 17:45 | 0 | 84.96 | 85 | 0 | 39.9 | 40 | 107.2984 | 0 | 4.5395 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.83 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5458 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1809 | 0.50071
2008-08-04 17:46 | 0 | 84.98 | 85 | 0 | 40.03 | 40 | 107.1974 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 20.37 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 17:47 | 0 | 84.96 | 85 | 0 | 39.9 | 40 | 107.1974 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.16 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1925 | 0.50071
2008-08-04 17:48 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 107.1974 | 0 | 4.5395 | 4.373 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 20.08 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1925 | 0.50071
2008-08-04 17:49 | 0 | 84.98 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.9 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1809 | 0.50071
2008-08-04 17:50 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5395 | 4.3689 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.21 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1809 | 0.50071
2008-08-04 17:51 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 107.1974 | 0 | 4.5436 | 4.3689 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 21.22 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:52 | 0 | 85.01 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 22 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 17:53 | 0 | 84.95 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 21.97 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 17:54 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 107.2984 | 0 | 4.5457 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 21.32 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1866 | 0.50071
2008-08-04 17:55 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5436 | 4.375 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 20.83 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:56 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.56 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:57 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.24 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:58 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 20.04 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 17:59 | 0 | 85.02 | 85 | 0 | 39.94 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.97 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5458 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 18:00 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.79 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 18:01 | 0 | 85.01 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.82 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 18:02 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.75 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 18:03 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.67 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 18:04 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.65 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:05 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5459 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.73 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1749 | 0.50071
2008-08-04 18:06 | 0 | 85.01 | 85 | 0 | 39.95 | 40 | 107.3489 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0442 | 4.19 | 4.5299 | 0 | 19.69 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:07 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 107.2984 | 0 | 4.5396 | 4.369 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.69 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1808 | 0.50071
2008-08-04 18:08 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5438 | 4.369 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.6 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:09 | 0 | 84.9 | 85 | 0 | 39.99 | 40 | 107.1974 | 0 | 4.5396 | 4.3731 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 19.61 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1808 | 0.50071
2008-08-04 18:10 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 107.1974 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.5 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1924 | 0.50071
2008-08-04 18:11 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 107.1974 | 0 | 4.5396 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 19.58 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1808 | 0.50071
2008-08-04 18:12 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 107.2984 | 0 | 4.5396 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.5 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1808 | 0.50071
2008-08-04 18:13 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.48 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:14 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 107.1974 | 0 | 4.5438 | 4.369 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.42 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:15 | 0 | 85.03 | 85 | 0 | 40 | 40 | 107.1974 | 0 | 4.5438 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 19.44 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:16 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 107.1974 | 0 | 4.5396 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.38 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1924 | 0.50071
2008-08-04 18:17 | 0 | 84.93 | 85 | 0 | 40 | 40 | 107.1974 | 0 | 4.5438 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.33 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:18 | 0 | 85.01 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5396 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.5 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1808 | 0.50071
2008-08-04 18:19 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5438 | 4.369 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.34 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:20 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 107.1974 | 0 | 4.5438 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.38 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:21 | 0 | 84.95 | 85 | 0 | 39.92 | 40 | 107.2984 | 0 | 4.5396 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.42 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1808 | 0.50071
2008-08-04 18:22 | 0 | 84.95 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.4 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:23 | 0 | 85.02 | 85 | 0 | 40.07 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.29 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:24 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.36 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:25 | 0 | 84.97 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.32 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:26 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.28 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:27 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.32 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:28 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 107.3489 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.37 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:29 | 0 | 84.95 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.37 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:30 | 0 | 85.01 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.42 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:31 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.34 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:32 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.42 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:33 | 0 | 84.93 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.38 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:34 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.32 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:35 | 0 | 85.01 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.34 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:36 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.21 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:37 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 107.1974 | 0 | 4.5396 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.26 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5165 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1924 | 0.50071
2008-08-04 18:38 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 107.1974 | 0 | 4.5396 | 4.3731 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.25 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:39 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5438 | 4.369 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.21 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:40 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.22 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:41 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.21 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:42 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.1974 | 0 | 4.5438 | 4.369 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.28 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:43 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.3 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:44 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.24 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:45 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5438 | 4.369 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.25 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:46 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 107.1974 | 0 | 4.5438 | 4.369 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.09 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:47 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.08 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:48 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 107.1974 | 0 | 4.5438 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 19.08 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:49 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5396 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 19.13 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1924 | 0.50071
2008-08-04 18:50 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.24 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:51 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5438 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.14 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1808 | 0.50071
2008-08-04 18:52 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.21 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 18:53 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 107.2984 | 0 | 4.5459 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.2 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1864 | 0.50071
2008-08-04 18:54 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0442 | 4.19 | 4.5421 | 0 | 19.13 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1864 | 0.50071
2008-08-04 18:55 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.18 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:56 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.17 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1749 | 0.50071
2008-08-04 18:57 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.16 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:58 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 107.2984 | 0 | 4.5436 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.21 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 18:59 | 0 | 84.95 | 85 | 0 | 40.04 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.13 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 19:00 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5436 | 4.375 | 0 |  | 1.0442 | 4.19 | 4.5421 | 0 | 19.06 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1866 | 0.50071
2008-08-04 19:01 | 0 | 85.02 | 85 | 0 | 39.92 | 40 | 107.2984 | 0 | 4.5436 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.1 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 19:02 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 107.2984 | 0 | 4.5459 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.05 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 19:03 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 107.3489 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0442 | 4.19 | 4.5421 | 0 | 19.16 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 19:04 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.04 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1864 | 0.50071
2008-08-04 19:05 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.01 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 19:06 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.04 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 19:07 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 107.1974 | 0 | 4.5438 | 4.369 | 0 |  | 1.0427 | 4.19 | 4.5421 | 0 | 19.14 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 19:08 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 107.1974 | 0 | 4.5438 | 4.373 | 0 |  | 1.0427 | 4.19 | 4.5299 | 0 | 19.08 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 19:09 | 0 | 84.93 | 85 | 0 | 39.98 | 40 | 107.1974 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.04 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 19:10 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.17 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1884 | 0.50071
2008-08-04 19:11 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.05 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1924 | 0.50071
2008-08-04 19:12 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 107.2984 | 0 | 4.5436 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.06 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 19:13 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.16 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 19:14 | 0 | 84.96 | 85 | 0 | 40 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.05 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 19:15 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.05 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.177 | 0.50071
2008-08-04 19:16 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.05 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5165 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 19:17 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.08 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 19:18 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.2984 | 0 | 4.5436 | 4.373 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.1 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 19:19 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5421 | 0 | 19.13 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1768 | 0.50071
2008-08-04 19:20 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 107.2984 | 0 | 4.5438 | 4.3731 | 0 |  | 1.0437 | 4.19 | 4.5299 | 0 | 19.09 | 0.53725 | 0 | 0.022432 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.5843 | -1 | 25.0042 | 4.1886 | 0.50071
2008-08-04 19:21 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 107.3489 | 0 | 13.3828 | 13.109 | 1 |  | 1.0442 | 12.56 | 9.8657 | 0 | 19.02 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5931 | -1 | 25.0126 | 4.1884 | 1.5004
2008-08-04 19:22 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.13 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 19:23 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.3886 | 13.1208 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.01 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5305 | 1.5004
2008-08-04 19:24 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 19:25 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.98 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 19:26 | 0 | 84.97 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.389 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.02 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 19:27 | 0 | 84.98 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.09 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 19:28 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 107.3489 | 0 | 13.389 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.01 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5305 | 1.5004
2008-08-04 19:29 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.97 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 19:30 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.94 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 19:31 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.98 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5305 | 1.5004
2008-08-04 19:32 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 107.3489 | 0 | 13.3886 | 13.1213 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.97 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 19:33 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.01 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 19:34 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.97 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 19:35 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 107.3489 | 0 | 13.389 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.93 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 19:36 | 0 | 84.97 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5305 | 1.5004
2008-08-04 19:37 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 19.01 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 19:38 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 19 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 19:39 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.98 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 19:40 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 107.3995 | 0 | 13.3953 | 13.1208 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 19:41 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 19:42 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 19:43 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 19.01 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 19:44 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 19.02 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5165 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 19:45 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.98 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5165 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 19:46 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3953 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 19:47 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.94 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 19:48 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 19:49 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 107.45 | 0 | 13.3948 | 13.1213 | 0 |  | 1.0452 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 19:50 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 19:51 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3886 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 19.04 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 19:52 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.97 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 19:53 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.93 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 19:54 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 19:55 | 0 | 85.03 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 19.01 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 19:56 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.01 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 19:57 | 0 | 85.01 | 85 | 0 | 39.95 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.13 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 19:58 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.94 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 19:59 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 13.389 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5305 | 1.5004
2008-08-04 20:00 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.93 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 20:01 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 20:02 | 0 | 84.95 | 85 | 0 | 39.92 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:03 | 0 | 84.93 | 85 | 0 | 39.95 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.85 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 20:04 | 0 | 84.95 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 20:05 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 20:06 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.9 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:07 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.85 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 20:08 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 19 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 20:09 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 107.3995 | 0 | 13.3948 | 13.1151 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 20:10 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 107.3995 | 0 | 13.3953 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.98 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 20:11 | 0 | 84.93 | 85 | 0 | 39.95 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 20:12 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 20:13 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 107.3995 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.94 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 20:14 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.97 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:15 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 107.3489 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.94 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:16 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.93 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 20:17 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:18 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 107.3489 | 0 | 13.389 | 13.1208 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.94 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 20:19 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.8 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:20 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 107.3995 | 0 | 13.389 | 13.1151 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 20:21 | 0 | 84.95 | 85 | 0 | 39.92 | 40 | 107.3995 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.98 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 20:22 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.05 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 20:23 | 0 | 84.98 | 85 | 0 | 40.04 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 20:24 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:25 | 0 | 85.01 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.04 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 20:26 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.01 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 20:27 | 0 | 85.01 | 85 | 0 | 40.02 | 40 | 107.3489 | 0 | 13.389 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 19 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 20:28 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 13.3886 | 13.1085 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:29 | 0 | 84.97 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.06 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 20:30 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 107.2984 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0437 | 12.56 | 13.3578 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 20:31 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.3886 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.85 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 20:32 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 19.01 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 20:33 | 0 | 84.98 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5305 | 1.5004
2008-08-04 20:34 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 13.3823 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5931 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 20:35 | 0 | 85.01 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 13.3886 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:36 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 107.3489 | 0 | 13.3886 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.81 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 20:37 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 20:38 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.98 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5305 | 1.5004
2008-08-04 20:39 | 0 | 84.96 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.3886 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.85 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5246 | 1.5004
2008-08-04 20:40 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 20:41 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.3489 | 0 | 13.3886 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.94 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 20:42 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 107.3995 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.94 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 20:43 | 0 | 84.99 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.389 | 13.1213 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 20:44 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 13.3886 | 13.1208 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 20:45 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 20:46 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 20:47 | 0 | 84.93 | 85 | 0 | 39.92 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 20:48 | 0 | 84.98 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.3953 | 13.1147 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.94 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 20:49 | 0 | 84.97 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.3953 | 13.1151 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.86 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 20:50 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.82 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 20:51 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.85 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 20:52 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.3953 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:53 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 107.3489 | 0 | 13.3953 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.9 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:54 | 0 | 84.97 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:55 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.93 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 20:56 | 0 | 84.96 | 85 | 0 | 39.9 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.89 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 20:57 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.389 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.85 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:58 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.89 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 20:59 | 0 | 84.96 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 21:00 | 0 | 84.96 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 21:01 | 0 | 84.95 | 85 | 0 | 39.92 | 40 | 107.3489 | 0 | 13.3886 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.82 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 21:02 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.86 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 21:03 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.86 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 21:04 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 21:05 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.93 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5305 | 1.5004
2008-08-04 21:06 | 0 | 84.96 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.3886 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.93 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 21:07 | 0 | 85.01 | 85 | 0 | 39.98 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 21:08 | 0 | 84.97 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.3886 | 13.1208 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.89 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 21:09 | 0 | 84.89 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 19 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 21:10 | 0 | 84.99 | 85 | 0 | 40.04 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.94 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:11 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 107.3995 | 0 | 13.3953 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 21:12 | 0 | 84.99 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0452 | 12.56 | 13.37 | 0 | 18.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 21:13 | 0 | 85.01 | 85 | 0 | 39.98 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 19.01 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 21:14 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 107.3995 | 0 | 13.3948 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 21:15 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 21:16 | 0 | 84.95 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.86 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 21:17 | 0 | 84.97 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.3948 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.86 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 21:18 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.8 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:19 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.8 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 21:20 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3886 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.9 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:21 | 0 | 84.93 | 85 | 0 | 39.95 | 40 | 107.3995 | 0 | 13.3948 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:22 | 0 | 84.99 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.89 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 21:23 | 0 | 84.98 | 85 | 0 | 40.04 | 40 | 107.3489 | 0 | 13.3886 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 21:24 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5305 | 1.5004
2008-08-04 21:25 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.86 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 21:26 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 107.3489 | 0 | 13.389 | 13.1213 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.86 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 21:27 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 21:28 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 107.3489 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 21:29 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 21:30 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.9 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 21:31 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 13.389 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.78 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:32 | 0 | 85.02 | 85 | 0 | 40.02 | 40 | 107.3489 | 0 | 13.389 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.85 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:33 | 0 | 84.93 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:34 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 107.3995 | 0 | 13.3886 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.85 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:35 | 0 | 84.93 | 85 | 0 | 39.92 | 40 | 107.3995 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 21:36 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 107.3489 | 0 | 13.3886 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.88 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 21:37 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.93 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 21:38 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3822 | 0 | 18.96 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5476 | 1.5004
2008-08-04 21:39 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:40 | 0 | 84.97 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3822 | 0 | 18.89 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:41 | 0 | 84.98 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.3953 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.82 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 21:42 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3953 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.9 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:43 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 107.3995 | 0 | 13.4016 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3822 | 0 | 18.98 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:44 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.93 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:45 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3948 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.89 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:46 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.85 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:47 | 0 | 84.93 | 85 | 0 | 39.94 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.92 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5246 | 1.5004
2008-08-04 21:48 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0451 | 12.56 | 13.37 | 0 | 18.85 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 21:49 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:50 | 0 | 84.99 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 21:51 | 0 | 84.98 | 85 | 0 | 39.92 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5307 | 1.5004
2008-08-04 21:52 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 107.3489 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.82 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 21:53 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.81 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 21:54 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 13.3948 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.8 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:55 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 107.3489 | 0 | 13.3886 | 13.1208 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.78 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 21:56 | 0 | 84.93 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.76 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 21:57 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3953 | 13.1147 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.78 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 21:58 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.77 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5165 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 21:59 | 0 | 84.98 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.82 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 22:00 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 107.3995 | 0 | 13.3953 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.78 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 22:01 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 107.3489 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:02 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 107.3489 | 0 | 13.3948 | 13.1213 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:03 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:04 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3822 | 0 | 18.81 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:05 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 107.3995 | 0 | 13.389 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.7 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 22:06 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 107.3995 | 0 | 13.3948 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.82 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 22:07 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.76 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 22:08 | 0 | 84.93 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.4016 | 13.127 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5302 | 1.5004
2008-08-04 22:09 | 0 | 84.98 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.3948 | 13.127 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.86 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:10 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.4011 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.77 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5458 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5307 | 1.5004
2008-08-04 22:11 | 0 | 84.97 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.82 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5302 | 1.5004
2008-08-04 22:12 | 0 | 85.06 | 85 | 0 | 39.99 | 40 | 107.45 | 0 | 13.4011 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.82 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:13 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5361 | 1.5004
2008-08-04 22:14 | 0 | 84.93 | 85 | 0 | 39.9 | 40 | 107.3995 | 0 | 13.3948 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.85 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5246 | 1.5004
2008-08-04 22:15 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3948 | 13.127 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.8 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:16 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3953 | 13.1213 | 0 |  | 1.0447 | 12.56 | 13.3822 | 0 | 18.76 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:17 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 107.3995 | 0 | 13.3948 | 13.127 | 0 |  | 1.0451 | 12.56 | 13.37 | 0 | 18.81 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:18 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.8 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:19 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.4011 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.8 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 22:20 | 0 | 84.96 | 85 | 0 | 40 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.8 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5187 | 1.5004
2008-08-04 22:21 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 107.45 | 0 | 13.4011 | 13.1213 | 0 |  | 1.0452 | 12.56 | 13.37 | 0 | 18.73 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:22 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.81 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:23 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.76 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 22:24 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.8 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:25 | 0 | 84.95 | 85 | 0 | 39.92 | 40 | 107.3489 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:26 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 107.3995 | 0 | 13.3886 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.76 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 22:27 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3953 | 13.1151 | 0 |  | 1.0442 | 12.56 | 13.3578 | 0 | 18.68 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 22:28 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.84 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 22:29 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 107.3995 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.82 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5251 | 1.5004
2008-08-04 22:30 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 107.3995 | 0 | 13.3948 | 13.1208 | 0 |  | 1.0447 | 12.56 | 13.3578 | 0 | 18.78 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5366 | 1.5004
2008-08-04 22:31 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 107.3489 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.85 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 22:32 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 107.3995 | 0 | 13.3948 | 13.1147 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.73 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5246 | 1.5004
2008-08-04 22:33 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.86 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.542 | 1.5004
2008-08-04 22:34 | 0 | 84.95 | 85 | 0 | 40 | 40 | 107.3489 | 0 | 13.3886 | 13.1147 | 0 |  | 1.0447 | 12.56 | 13.37 | 0 | 18.82 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.531 | 1.5004
2008-08-04 22:35 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 107.3489 | 0 | 13.3953 | 13.1147 | 0 |  | 1.0442 | 12.56 | 13.37 | 0 | 18.76 | 1.6099 | 0 | 0.067218 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0126 | 23.5932 | -1 | 25.0126 | 12.5425 | 1.5004
2008-08-04 22:36 | 0 | 84.61 | 85 | 0 | 40.06 | 40 | 100.7321 | 0 | 7.2254 | 6.7825 | 1 |  | 0.98008 | 6.73 | 9.4994 | 0 | 18.35 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.8571 | 0 | 0 | 0 | 15 | 3.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7759 | -1 | 20.0067 | 12.5425 | 0
2008-08-04 22:37 | 0 | 85.11 | 85 | 0 | 39.96 | 40 | 100.6311 | 0 | 6.6368 | 6.4451 | 0 |  | 0.95723 | 6.73 | 6.63 | 0 | 18.29 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 10.25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7753 | -1 | 20.0067 | 6.7182 | 0
2008-08-04 22:38 | 0 | 85.03 | 85 | 0 | 39.97 | 40 | 100.6311 | 0 | 6.5648 | 6.3736 | 0 |  | 0.94668 | 6.73 | 6.5568 | 0 | 18.35 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 13.41 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7228 | 0
2008-08-04 22:39 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 100.6311 | 0 | 6.5395 | 6.3567 | 0 |  | 0.9445 | 6.73 | 6.5324 | 0 | 18.06 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.37 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7199 | 0
2008-08-04 22:40 | 0 | 85.04 | 85 | 0 | 39.99 | 40 | 100.6311 | 0 | 6.5304 | 6.3419 | 0 |  | 0.9423 | 6.73 | 6.5201 | 0 | 18.2 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.74 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7183 | 0
2008-08-04 22:41 | 0 | 85.02 | 85 | 0 | 39.95 | 40 | 100.6311 | 0 | 6.5271 | 6.3386 | 0 |  | 0.94185 | 6.73 | 6.5201 | 0 | 18.4 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 14.88 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7225 | 0
2008-08-04 22:42 | 0 | 84.99 | 85 | 0 | 40 | 40 | 100.6311 | 0 | 6.5255 | 6.3371 | 0 |  | 0.94162 | 6.73 | 6.5079 | 0 | 18.4 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7115 | 0
2008-08-04 22:43 | 0 | 85.04 | 85 | 0 | 39.99 | 40 | 100.6311 | 0 | 6.5244 | 6.336 | 0 |  | 0.94145 | 6.73 | 6.5079 | 0 | 18.47 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7124 | 0
2008-08-04 22:44 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 100.6311 | 0 | 6.5237 | 6.3353 | 0 |  | 0.94136 | 6.73 | 6.5079 | 0 | 18.35 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15.03 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7131 | 0
2008-08-04 22:45 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 100.6311 | 0 | 6.5233 | 6.3349 | 0 |  | 0.94129 | 6.73 | 6.5079 | 0 | 18.36 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15.05 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7007 | 0
2008-08-04 22:46 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 100.6311 | 0 | 6.5231 | 6.3347 | 0 |  | 0.94126 | 6.73 | 6.5079 | 0 | 18.42 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.06 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7141 | 0
2008-08-04 22:47 | 0 | 84.95 | 85 | 0 | 40 | 40 | 100.6311 | 0 | 6.5228 | 6.3344 | 0 |  | 0.94122 | 6.73 | 6.5079 | 0 | 18.48 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7141 | 0
2008-08-04 22:48 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 100.6311 | 0 | 6.5228 | 6.3344 | 0 |  | 0.94122 | 6.73 | 6.5079 | 0 | 18.51 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7141 | 0
2008-08-04 22:49 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 100.6311 | 0 | 6.5228 | 6.3344 | 0 |  | 0.94122 | 6.73 | 6.4957 | 0 | 18.54 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7143 | 0
2008-08-04 22:50 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 100.6311 | 0 | 6.5228 | 6.3344 | 0 |  | 0.94122 | 6.73 | 6.5079 | 0 | 18.42 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7143 | 0
2008-08-04 22:51 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 100.6311 | 0 | 6.5228 | 6.3344 | 0 |  | 0.94122 | 6.73 | 6.5079 | 0 | 18.37 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7143 | 0
2008-08-04 22:52 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 100.6311 | 0 | 6.5228 | 6.3344 | 0 |  | 0.94122 | 6.73 | 6.5079 | 0 | 18.46 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7141 | 0
2008-08-04 22:53 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 100.6311 | 0 | 6.5231 | 6.3344 | 0 |  | 0.94122 | 6.73 | 6.5079 | 0 | 18.46 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.07 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7141 | 0
2008-08-04 22:54 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 100.6311 | 0 | 6.5231 | 6.3347 | 0 |  | 0.94126 | 6.73 | 6.4957 | 0 | 18.4 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15.06 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7141 | 0
2008-08-04 22:55 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 100.6311 | 0 | 6.5231 | 6.3347 | 0 |  | 0.94126 | 6.73 | 6.5079 | 0 | 18.47 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15.06 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7141 | 0
2008-08-04 22:56 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 100.6311 | 0 | 6.5231 | 6.3349 | 0 |  | 0.94078 | 6.73 | 6.5079 | 0 | 18.55 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.06 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7138 | 0
2008-08-04 22:57 | 0 | 84.9 | 85 | 0 | 39.96 | 40 | 100.6311 | 0 | 6.5233 | 6.3349 | 0 |  | 0.94129 | 6.73 | 6.5079 | 0 | 18.46 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.05 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7138 | 0
2008-08-04 22:58 | 0 | 85.04 | 85 | 0 | 39.97 | 40 | 100.6311 | 0 | 6.5233 | 6.3349 | 0 |  | 0.94129 | 6.73 | 6.5079 | 0 | 18.49 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.05 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7138 | 0
2008-08-04 22:59 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 100.6311 | 0 | 6.5233 | 6.3349 | 0 |  | 0.94129 | 6.73 | 6.5079 | 0 | 18.51 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.04 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7138 | 0
2008-08-04 23:00 | 0 | 84.96 | 85 | 0 | 40 | 40 | 100.6311 | 0 | 6.5235 | 6.3351 | 0 |  | 0.94132 | 6.73 | 6.5079 | 0 | 18.43 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.04 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7136 | 0
2008-08-04 23:01 | 0 | 84.93 | 85 | 0 | 39.95 | 40 | 100.6311 | 0 | 6.5235 | 6.3351 | 0 |  | 0.94132 | 6.73 | 6.5079 | 0 | 18.39 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.04 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7136 | 0
2008-08-04 23:02 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 100.6311 | 0 | 6.5237 | 6.3353 | 0 |  | 0.94136 | 6.73 | 6.5079 | 0 | 18.42 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.03 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7134 | 0
2008-08-04 23:03 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 100.6311 | 0 | 6.5237 | 6.3353 | 0 |  | 0.94136 | 6.73 | 6.5079 | 0 | 18.42 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.03 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7134 | 0
2008-08-04 23:04 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 100.6311 | 0 | 6.5237 | 6.3353 | 0 |  | 0.94136 | 6.73 | 6.5079 | 0 | 18.45 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15.03 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7134 | 0
2008-08-04 23:05 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 100.6311 | 0 | 6.5237 | 6.3353 | 0 |  | 0.94136 | 6.73 | 6.5079 | 0 | 18.51 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.03 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7134 | 0
2008-08-04 23:06 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 100.6311 | 0 | 6.5237 | 6.3353 | 0 |  | 0.94139 | 6.73 | 6.5201 | 0 | 18.42 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.03 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7134 | 0
2008-08-04 23:07 | 0 | 85.01 | 85 | 0 | 39.98 | 40 | 100.6311 | 0 | 6.524 | 6.3355 | 0 |  | 0.94139 | 6.73 | 6.5079 | 0 | 18.47 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7134 | 0
2008-08-04 23:08 | 0 | 85.01 | 85 | 0 | 39.9 | 40 | 100.6311 | 0 | 6.524 | 6.3355 | 0 |  | 0.94139 | 6.73 | 6.5079 | 0 | 18.46 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7131 | 0
2008-08-04 23:09 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 100.6311 | 0 | 6.524 | 6.3355 | 0 |  | 0.94139 | 6.73 | 6.5079 | 0 | 18.44 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7131 | 0
2008-08-04 23:10 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 100.6311 | 0 | 6.5242 | 6.3355 | 0 |  | 0.94139 | 6.73 | 6.5079 | 0 | 18.45 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7129 | 0
2008-08-04 23:11 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 100.6311 | 0 | 6.524 | 6.3358 | 0 |  | 0.94142 | 6.73 | 6.5079 | 0 | 18.43 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7129 | 0
2008-08-04 23:12 | 0 | 84.96 | 85 | 0 | 40 | 40 | 100.6311 | 0 | 6.5242 | 6.3355 | 0 |  | 0.94142 | 6.73 | 6.5079 | 0 | 18.42 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7129 | 0
2008-08-04 23:13 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 100.6311 | 0 | 6.5242 | 6.3358 | 0 |  | 0.94142 | 6.73 | 6.5201 | 0 | 18.5 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7129 | 0
2008-08-04 23:14 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 100.6311 | 0 | 6.5242 | 6.3358 | 0 |  | 0.94142 | 6.73 | 6.5079 | 0 | 18.37 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7129 | 0
2008-08-04 23:15 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 100.6311 | 0 | 6.5242 | 6.3358 | 0 |  | 0.94142 | 6.73 | 6.5079 | 0 | 18.42 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7129 | 0
2008-08-04 23:16 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 100.6311 | 0 | 6.5242 | 6.3358 | 0 |  | 0.94142 | 6.73 | 6.5079 | 0 | 18.5 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7129 | 0
2008-08-04 23:17 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 100.6311 | 0 | 6.5242 | 6.3358 | 0 |  | 0.94142 | 6.73 | 6.5079 | 0 | 18.43 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7129 | 0
2008-08-04 23:18 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 100.6311 | 0 | 6.5242 | 6.3358 | 0 |  | 0.94142 | 6.73 | 6.5079 | 0 | 18.49 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7129 | 0
2008-08-04 23:19 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 100.6311 | 0 | 6.5242 | 6.3358 | 0 |  | 0.94145 | 6.73 | 6.5079 | 0 | 18.45 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7129 | 0
2008-08-04 23:20 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 100.6311 | 0 | 6.5242 | 6.3358 | 0 |  | 0.94142 | 6.73 | 6.5079 | 0 | 18.44 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7127 | 0
2008-08-04 23:21 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.5244 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.5079 | 0 | 18.39 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7752 | -1 | 20.0067 | 6.7127 | 2.7021
2008-08-04 23:22 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.7888 | 0 | 18.39 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:23 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98351 | 6.73 | 6.801 | 0 | 18.57 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:24 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.49 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:25 | 0 | 85.01 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:26 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:27 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.49 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:28 | 0 | 85.04 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.43 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.718 | 2.7021
2008-08-04 23:29 | 0 | 85.03 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98395 | 6.73 | 6.801 | 0 | 18.53 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.718 | 2.7021
2008-08-04 23:30 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.42 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:31 | 0 | 84.93 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.47 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:32 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8129 | 6.6283 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.51 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:33 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8129 | 6.622 | 0 |  | 0.98395 | 6.73 | 6.7888 | 0 | 18.59 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:34 | 0 | 84.97 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98395 | 6.73 | 6.801 | 0 | 18.54 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-04 23:35 | 0 | 84.99 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8193 | 6.6283 | 0 |  | 0.98395 | 6.73 | 6.801 | 0 | 18.41 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7236 | 2.7021
2008-08-04 23:36 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 6.8129 | 6.622 | 0 |  | 0.98395 | 6.73 | 6.801 | 0 | 18.42 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:37 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.51 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-04 23:38 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 105.177 | 0 | 6.8129 | 6.622 | 0 |  | 0.98395 | 6.73 | 6.801 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.718 | 2.7021
2008-08-04 23:39 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-04 23:40 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 6.8129 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.8132 | 0 | 18.46 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7054 | 2.7021
2008-08-04 23:41 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 6.8129 | 6.622 | 0 |  | 0.98395 | 6.73 | 6.801 | 0 | 18.49 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:42 | 0 | 84.93 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.7888 | 0 | 18.41 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7054 | 2.7021
2008-08-04 23:43 | 0 | 85.01 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8129 | 6.622 | 0 |  | 0.98395 | 6.73 | 6.801 | 0 | 18.53 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.718 | 2.7021
2008-08-04 23:44 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.718 | 2.7021
2008-08-04 23:45 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 6.8129 | 6.622 | 0 |  | 0.98395 | 6.73 | 6.8132 | 0 | 18.44 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:46 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:47 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.5 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.718 | 2.7021
2008-08-04 23:48 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.59 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:49 | 0 | 84.98 | 85 | 0 | 39.92 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.52 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:50 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.5 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:51 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.56 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:52 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.49 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:53 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.51 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:54 | 0 | 85.03 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.7888 | 0 | 18.51 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:55 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.44 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:56 | 0 | 84.99 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.6 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:57 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.59 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:58 | 0 | 84.97 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.55 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-04 23:59 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.5 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.619 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:01 | 0 | 84.9 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.53 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.721 | 2.7021
2008-08-05 00:02 | 0 | 84.93 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6286 | 0 |  | 0.98493 | 6.73 | 6.8132 | 0 | 18.65 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:03 | 0 | 85.02 | 85 | 0 | 40.03 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.8132 | 0 | 18.56 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:04 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6286 | 0 |  | 0.98398 | 6.73 | 6.8132 | 0 | 18.4 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:05 | 0 | 84.98 | 85 | 0 | 39.9 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.41 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 00:06 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.42 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7236 | 2.7021
2008-08-05 00:07 | 0 | 84.95 | 85 | 0 | 40.07 | 40 | 105.278 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:08 | 0 | 84.96 | 85 | 0 | 39.92 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.4 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:09 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.39 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:10 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8132 | 6.6286 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.47 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:11 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7236 | 2.7021
2008-08-05 00:12 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.47 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 00:13 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.46 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 00:14 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.57 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:15 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.52 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:16 | 0 | 84.98 | 85 | 0 | 40.03 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.8132 | 0 | 18.56 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7106 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:17 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.47 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:18 | 0 | 84.96 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:19 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.44 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:20 | 0 | 84.96 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:21 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 6.8134 | 6.6224 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.55 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:22 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.55 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:23 | 0 | 84.99 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:24 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.41 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:25 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.4 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:26 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:27 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.54 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 00:28 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.55 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 00:29 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.58 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:30 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.51 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:31 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.61 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:32 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 6.8196 | 6.6286 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.55 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7236 | 2.7021
2008-08-05 00:33 | 0 | 84.98 | 85 | 0 | 40.04 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.58 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:34 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.8132 | 0 | 18.49 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 00:35 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.8132 | 0 | 18.52 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 00:36 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8132 | 6.6286 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.44 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:37 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:38 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8196 | 6.6286 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 00:39 | 0 | 84.99 | 85 | 0 | 40 | 40 | 105.278 | 0 | 6.8132 | 6.6286 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.53 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:40 | 0 | 85.02 | 85 | 0 | 39.9 | 40 | 105.278 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.52 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 00:41 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.5 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 00:42 | 0 | 84.93 | 85 | 0 | 40.03 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.49 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:43 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8196 | 6.6286 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.47 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 00:44 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 00:45 | 0 | 84.96 | 85 | 0 | 40 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.56 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:46 | 0 | 84.97 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6286 | 0 |  | 0.98493 | 6.73 | 6.8132 | 0 | 18.49 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:47 | 0 | 85.01 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:48 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.56 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:49 | 0 | 85.03 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.55 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:50 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:51 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:52 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:53 | 0 | 85.01 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.56 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:54 | 0 | 84.96 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.49 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:55 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.47 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:56 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.51 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:57 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.47 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:58 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.41 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 00:59 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.46 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:00 | 0 | 84.98 | 85 | 0 | 40.04 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.68 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:01 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.5 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7175 | 2.7021
2008-08-05 01:02 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 6.8132 | 6.6286 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.44 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 01:03 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 6.8132 | 6.6286 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.56 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 01:04 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 6.8132 | 6.6286 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.6 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 01:05 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.58 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 01:06 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.61 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 01:07 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.278 | 0 | 6.8196 | 6.6286 | 0 |  | 0.98493 | 6.73 | 6.8132 | 0 | 18.46 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7236 | 2.7021
2008-08-05 01:08 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.49 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:09 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6286 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.51 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7236 | 2.7021
2008-08-05 01:10 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.47 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7423 | 2.7021
2008-08-05 01:11 | 0 | 85.01 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.8132 | 0 | 18.51 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7236 | 2.7021
2008-08-05 01:12 | 0 | 85.01 | 85 | 0 | 39.93 | 40 | 105.278 | 0 | 6.8196 | 6.6286 | 0 |  | 0.98493 | 6.73 | 6.7888 | 0 | 18.41 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7236 | 2.7021
2008-08-05 01:13 | 0 | 84.9 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.52 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7175 | 2.7021
2008-08-05 01:14 | 0 | 85.02 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.5 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 01:15 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 6.8196 | 6.6286 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.53 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 01:16 | 0 | 84.96 | 85 | 0 | 39.92 | 40 | 105.278 | 0 | 6.8132 | 6.6286 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.52 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 01:17 | 0 | 84.96 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:18 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.52 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:19 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.44 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:20 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.56 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:21 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.58 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7106 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:22 | 0 | 84.93 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.58 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:23 | 0 | 84.93 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.53 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:24 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:25 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.57 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:26 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.64 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:27 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.59 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:28 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.5 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:29 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 105.1265 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98351 | 6.73 | 6.801 | 0 | 18.57 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:30 | 0 | 84.96 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.46 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:31 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.47 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:32 | 0 | 85.01 | 85 | 0 | 39.92 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.8132 | 0 | 18.46 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 01:33 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.54 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:34 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.52 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:35 | 0 | 85.01 | 85 | 0 | 40.06 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.52 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 01:36 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.51 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 01:37 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.8132 | 0 | 18.46 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 01:38 | 0 | 84.91 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.8132 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:39 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.51 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7423 | 2.7021
2008-08-05 01:40 | 0 | 84.95 | 85 | 0 | 40 | 40 | 105.278 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.55 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:41 | 0 | 84.93 | 85 | 0 | 39.93 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.56 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:42 | 0 | 84.93 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6286 | 0 |  | 0.98493 | 6.73 | 6.8132 | 0 | 18.47 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 01:43 | 0 | 85.01 | 85 | 0 | 39.92 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.43 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 01:44 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.8132 | 0 | 18.4 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 01:45 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98493 | 6.73 | 6.8132 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:46 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8196 | 6.6286 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.53 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7236 | 2.7021
2008-08-05 01:47 | 0 | 85.01 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.53 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:48 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.8132 | 0 | 18.42 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:49 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.46 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:50 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.51 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:51 | 0 | 84.99 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.44 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:52 | 0 | 84.97 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.54 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 01:53 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.47 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7106 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:54 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:55 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.42 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:56 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.6 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:57 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.57 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:58 | 0 | 84.98 | 85 | 0 | 40.04 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 01:59 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.46 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:00 | 0 | 85.01 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.44 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 02:01 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.61 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:02 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.57 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:03 | 0 | 85.04 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.52 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:04 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.49 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:05 | 0 | 85.03 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.7888 | 0 | 18.65 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:06 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 105.278 | 0 | 6.8132 | 6.6286 | 0 |  | 0.98493 | 6.73 | 6.801 | 0 | 18.53 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:07 | 0 | 85.01 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.56 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:08 | 0 | 84.93 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.59 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 02:09 | 0 | 85.03 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.46 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:10 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 105.177 | 0 | 6.8196 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.53 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7113 | 2.7021
2008-08-05 02:11 | 0 | 84.99 | 85 | 0 | 39.9 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.5 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:12 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.53 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:13 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.45 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:14 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.46 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:15 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.47 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:16 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.43 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:17 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.52 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:18 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.7888 | 0 | 18.64 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:19 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.51 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:20 | 0 | 84.93 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:21 | 0 | 85.01 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.47 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:22 | 0 | 84.96 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.46 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:23 | 0 | 84.93 | 85 | 0 | 40.03 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.5 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:24 | 0 | 84.93 | 85 | 0 | 40.06 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.52 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.73 | 2.7021
2008-08-05 02:25 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.48 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:26 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.44 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:27 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.55 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:28 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.7888 | 0 | 18.54 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:29 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8132 | 6.6224 | 0 |  | 0.98402 | 6.73 | 6.801 | 0 | 18.61 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:30 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.57 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:31 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.46 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:32 | 0 | 84.9 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.59 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:33 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.52 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:34 | 0 | 85.01 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.5 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:35 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 6.8132 | 6.6222 | 0 |  | 0.98398 | 6.73 | 6.801 | 0 | 18.49 | 2.8992 | 0 | 0.12105 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0067 | 18.7755 | -1 | 20.0067 | 6.7177 | 2.7021
2008-08-05 02:36 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 1 |  | 0.98398 | 14.21 | 7.558 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 6.7177 | 5.7031
2008-08-05 02:37 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.177 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 02:38 | 0 | 85.01 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.49 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 02:39 | 0 | 84.93 | 85 | 0 | 40 | 40 | 105.278 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 02:40 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.48 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1871 | 5.7031
2008-08-05 02:41 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 02:42 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.52 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 02:43 | 0 | 84.96 | 85 | 0 | 40 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.56 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 02:44 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.69 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 02:45 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 105.278 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.6 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 02:46 | 0 | 84.91 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.66 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 02:47 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.53 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 02:48 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 105.278 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 02:49 | 0 | 85.04 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.45 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 02:50 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 02:51 | 0 | 84.98 | 85 | 0 | 40.11 | 40 | 105.278 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.47 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 02:52 | 0 | 84.95 | 85 | 0 | 40 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1871 | 5.7031
2008-08-05 02:53 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.47 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 02:54 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.52 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 02:55 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.52 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 02:56 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 02:57 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.52 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1681 | 5.7031
2008-08-05 02:58 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.52 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 02:59 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 14.2911 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.52 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:00 | 0 | 84.97 | 85 | 0 | 40.09 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:01 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.3285 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:02 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:03 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:04 | 0 | 84.91 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1744 | 5.7031
2008-08-05 03:05 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2369 | 0 | 18.56 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 03:06 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.53 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:07 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.46 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:08 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 03:09 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.48 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 03:10 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:11 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 03:12 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.52 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 03:13 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.59 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 03:14 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.52 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 03:15 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.63 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 03:16 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.48 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:17 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 105.278 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 03:18 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.49 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 03:19 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.46 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 03:20 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:21 | 0 | 84.93 | 85 | 0 | 39.93 | 40 | 105.177 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.6 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7106 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:22 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.61 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:23 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.59 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:24 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.49 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 03:25 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1871 | 5.7031
2008-08-05 03:26 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.47 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:27 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.6 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.2007 | 5.7031
2008-08-05 03:28 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.61 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:29 | 0 | 84.93 | 85 | 0 | 39.98 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.45 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:30 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.59 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:31 | 0 | 84.95 | 85 | 0 | 40 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:32 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 105.278 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:33 | 0 | 84.95 | 85 | 0 | 40 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2369 | 0 | 18.64 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:34 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1871 | 5.7031
2008-08-05 03:35 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 105.278 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.61 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:36 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:37 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.56 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:38 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:39 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.52 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:40 | 0 | 85.01 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.53 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7473 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:41 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.58 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 03:42 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:43 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.48 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:44 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.48 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 03:45 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.61 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 03:46 | 0 | 84.91 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.65 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:47 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:48 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.59 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:49 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.65 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:50 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:51 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 03:52 | 0 | 85.01 | 85 | 0 | 39.89 | 40 | 105.278 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:53 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:54 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2369 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1871 | 5.7031
2008-08-05 03:55 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.53 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:56 | 0 | 85.02 | 85 | 0 | 39.92 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.61 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:57 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 105.177 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 03:58 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 105.278 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.53 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 03:59 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:00 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 04:01 | 0 | 84.93 | 85 | 0 | 39.93 | 40 | 105.278 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:02 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:03 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 04:04 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 105.278 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.46 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:05 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:06 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2369 | 0 | 18.63 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 04:07 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.58 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:08 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.49 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 04:09 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.56 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:10 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.47 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:11 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 105.177 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.2007 | 5.7031
2008-08-05 04:12 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.58 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:13 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:14 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.47 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:15 | 0 | 84.95 | 85 | 0 | 39.92 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.48 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:16 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 04:17 | 0 | 84.96 | 85 | 0 | 39.9 | 40 | 105.177 | 0 | 14.2847 | 13.9829 | 0 |  | 0.98496 | 14.21 | 14.2369 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:18 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.49 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:19 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.56 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:20 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 04:21 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.48 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:22 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.4 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:23 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:24 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:25 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:26 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.45 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:27 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 14.2842 | 13.9963 | 0 |  | 0.98496 | 14.21 | 14.2613 | 0 | 18.52 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:28 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2847 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.43 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:29 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2847 | 13.9829 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.52 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1744 | 5.7031
2008-08-05 04:30 | 0 | 84.95 | 85 | 0 | 40 | 40 | 105.278 | 0 | 14.2842 | 13.9963 | 0 |  | 0.98496 | 14.21 | 14.2491 | 0 | 18.44 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:31 | 0 | 84.96 | 85 | 0 | 40.06 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:32 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.56 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1744 | 5.7031
2008-08-05 04:33 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.43 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:34 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2842 | 13.9963 | 0 |  | 0.98496 | 14.21 | 14.2491 | 0 | 18.53 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:35 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.62 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1744 | 5.7031
2008-08-05 04:36 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:37 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 105.278 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.53 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 04:38 | 0 | 84.95 | 85 | 0 | 40 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 04:39 | 0 | 84.99 | 85 | 0 | 40 | 40 | 105.177 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.49 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 04:40 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 105.278 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:41 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1744 | 5.7031
2008-08-05 04:42 | 0 | 84.93 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.63 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 04:43 | 0 | 84.95 | 85 | 0 | 39.9 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.49 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 04:44 | 0 | 84.97 | 85 | 0 | 40 | 40 | 105.177 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.47 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 04:45 | 0 | 84.97 | 85 | 0 | 40 | 40 | 105.177 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.59 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:46 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.63 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 04:47 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.6 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:48 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.47 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 04:49 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 04:50 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2369 | 0 | 18.46 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.2007 | 5.7031
2008-08-05 04:51 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 14.2706 | 13.9963 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1622 | 5.7031
2008-08-05 04:52 | 0 | 85.01 | 85 | 0 | 39.92 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.56 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:53 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:54 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:55 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.56 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:56 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.47 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 04:57 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 105.278 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.58 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:58 | 0 | 84.97 | 85 | 0 | 40 | 40 | 105.278 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.53 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 04:59 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.42 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:00 | 0 | 85.02 | 85 | 0 | 40.03 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2369 | 0 | 18.56 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:01 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.47 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:02 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.47 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:03 | 0 | 84.9 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2369 | 0 | 18.62 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 05:04 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.56 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:05 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.6 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:06 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:07 | 0 | 84.99 | 85 | 0 | 40 | 40 | 105.278 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2369 | 0 | 18.59 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 05:08 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.52 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 05:09 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 105.177 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.49 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 05:10 | 0 | 84.99 | 85 | 0 | 40 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 05:11 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.53 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 05:12 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:13 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:14 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:15 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2369 | 0 | 18.47 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 05:16 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 05:17 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.48 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.188 | 5.7031
2008-08-05 05:18 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.177 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.68 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:19 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.65 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:20 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.6 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 05:21 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.278 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.58 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 05:22 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:23 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:24 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.53 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:25 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.56 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:26 | 0 | 85.03 | 85 | 0 | 39.99 | 40 | 105.278 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.5 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:27 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.58 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:28 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.49 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:29 | 0 | 85.01 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1871 | 5.7031
2008-08-05 05:30 | 0 | 85.01 | 85 | 0 | 39.98 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.49 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:31 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.6 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:32 | 0 | 84.96 | 85 | 0 | 39.92 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2613 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:33 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.59 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7106 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:34 | 0 | 84.93 | 85 | 0 | 39.95 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1871 | 5.7031
2008-08-05 05:35 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.6 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:36 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2369 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:37 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.53 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 05:38 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.66 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 05:39 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.69 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 05:40 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 105.177 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.51 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 05:41 | 0 | 84.93 | 85 | 0 | 40.03 | 40 | 105.177 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.6 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1627 | 5.7031
2008-08-05 05:42 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.278 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98493 | 14.21 | 14.2369 | 0 | 18.55 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 05:43 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.59 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 05:44 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.177 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 05:45 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 105.177 | 0 | 14.2706 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2491 | 0 | 18.54 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:46 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.177 | 0 | 14.2842 | 13.9824 | 0 |  | 0.98398 | 14.21 | 14.2369 | 0 | 18.61 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1763 | 5.7031
2008-08-05 05:47 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 105.278 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.59 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:48 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 14.2706 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.58 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:49 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2491 | 0 | 18.57 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1749 | 5.7031
2008-08-05 05:50 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.278 | 0 | 14.2842 | 13.9958 | 0 |  | 0.98493 | 14.21 | 14.2369 | 0 | 18.63 | 6.1192 | 0 | 0.2555 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0142 | 18.7829 | -1 | 20.0142 | 14.1885 | 5.7031
2008-08-05 05:51 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.3285 | 0 | 23.6329 | 23.3106 | 0 |  | 0.98493 | 23.69 | 22.0513 | 0 | 18.52 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 19.2205 | 9.5035
2008-08-05 05:52 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 105.379 | 0 | 23.6441 | 23.3553 | 0 |  | 0.98587 | 23.69 | 23.5897 | 0 | 18.6 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6348 | 9.5034
2008-08-05 05:53 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.379 | 0 | 23.6441 | 23.3553 | 0 |  | 0.9854 | 23.69 | 23.5775 | 0 | 18.59 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6348 | 9.5034
2008-08-05 05:54 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.379 | 0 | 23.6553 | 23.3441 | 0 |  | 0.98587 | 23.69 | 23.5897 | 0 | 18.56 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6348 | 9.5034
2008-08-05 05:55 | 0 | 84.95 | 85 | 0 | 40 | 40 | 105.3285 | 0 | 23.6329 | 23.3441 | 0 |  | 0.9854 | 23.69 | 23.5775 | 0 | 18.56 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6224 | 9.5034
2008-08-05 05:56 | 0 | 85.02 | 85 | 0 | 39.93 | 40 | 105.379 | 0 | 23.6553 | 23.3441 | 0 |  | 0.9854 | 23.69 | 23.5897 | 0 | 18.56 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6348 | 9.5034
2008-08-05 05:57 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 105.3285 | 0 | 23.6441 | 23.3553 | 0 |  | 0.9854 | 23.69 | 23.5897 | 0 | 18.66 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6472 | 9.5034
2008-08-05 05:58 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.3285 | 0 | 23.6553 | 23.3553 | 0 |  | 0.98587 | 23.69 | 23.5897 | 0 | 18.6 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6235 | 9.5034
2008-08-05 05:59 | 0 | 84.98 | 85 | 0 | 39.92 | 40 | 105.379 | 0 | 23.6441 | 23.3441 | 0 |  | 0.98587 | 23.69 | 23.5897 | 0 | 18.6 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6224 | 9.5034
2008-08-05 06:00 | 0 | 85.02 | 85 | 0 | 39.98 | 40 | 105.379 | 0 | 23.6441 | 23.3441 | 0 |  | 0.98587 | 23.69 | 23.5897 | 0 | 18.63 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6235 | 9.5034
2008-08-05 06:01 | 0 | 84.93 | 85 | 0 | 39.9 | 40 | 105.379 | 0 | 23.6553 | 23.3441 | 0 |  | 0.98587 | 23.69 | 23.5897 | 0 | 18.6 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6235 | 9.5034
2008-08-05 06:02 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.379 | 0 | 23.6441 | 23.3441 | 0 |  | 0.9854 | 23.69 | 23.5897 | 0 | 18.58 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6348 | 9.5034
2008-08-05 06:03 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.379 | 0 | 23.6441 | 23.3665 | 0 |  | 0.98587 | 23.69 | 23.5775 | 0 | 18.61 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6224 | 9.5034
2008-08-05 06:04 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 105.379 | 0 | 23.6553 | 23.3553 | 0 |  | 0.98587 | 23.69 | 23.5775 | 0 | 19.02 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6235 | 9.5034
2008-08-05 06:05 | 0 | 84.93 | 85 | 0 | 39.94 | 40 | 105.379 | 0 | 23.6561 | 23.3561 | 0 |  | 0.98591 | 23.69 | 23.5897 | 0 | 20.47 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6235 | 9.5034
2008-08-05 06:06 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6553 | 23.3553 | 0 |  | 0.98587 | 23.69 | 23.5897 | 0 | 21.67 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 06:07 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 22.29 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 06:08 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 22.27 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:09 | 0 | 84.96 | 85 | 0 | 40.06 | 40 | 105.4295 | 0 | 23.6657 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.602 | 0 | 21.82 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6377 | 9.5034
2008-08-05 06:10 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6657 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.6142 | 0 | 21.11 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6254 | 9.5034
2008-08-05 06:11 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6657 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.602 | 0 | 20.48 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6377 | 9.5034
2008-08-05 06:12 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6657 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.6142 | 0 | 19.97 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6377 | 9.5034
2008-08-05 06:13 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6657 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.602 | 0 | 19.81 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6377 | 9.5034
2008-08-05 06:14 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6657 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.602 | 0 | 19.5 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6254 | 9.5034
2008-08-05 06:15 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6657 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.602 | 0 | 19.46 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7106 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6377 | 9.5034
2008-08-05 06:16 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6657 | 23.3545 | 0 |  | 0.98631 | 23.69 | 23.602 | 0 | 19.42 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6254 | 9.5034
2008-08-05 06:17 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 105.4295 | 0 | 23.6545 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.602 | 0 | 20.28 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6254 | 9.5034
2008-08-05 06:18 | 0 | 84.93 | 85 | 0 | 40.03 | 40 | 105.4295 | 0 | 23.6657 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.6142 | 0 | 21.08 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:19 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6545 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.602 | 0 | 21.59 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6491 | 9.5034
2008-08-05 06:20 | 0 | 84.96 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6657 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.602 | 0 | 21.39 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6377 | 9.5034
2008-08-05 06:21 | 0 | 85.03 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6657 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.6142 | 0 | 20.84 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6254 | 9.5034
2008-08-05 06:22 | 0 | 85.01 | 85 | 0 | 40.02 | 40 | 105.379 | 0 | 23.6545 | 23.3545 | 0 |  | 0.98584 | 23.69 | 23.602 | 0 | 20.33 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6491 | 9.5034
2008-08-05 06:23 | 0 | 85.01 | 85 | 0 | 40.02 | 40 | 105.379 | 0 | 23.6545 | 23.3545 | 0 |  | 0.98584 | 23.69 | 23.602 | 0 | 20.04 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6367 | 9.5034
2008-08-05 06:24 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.379 | 0 | 23.6545 | 23.3545 | 0 |  | 0.98584 | 23.69 | 23.602 | 0 | 19.71 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6232 | 9.5034
2008-08-05 06:25 | 0 | 85.01 | 85 | 0 | 39.99 | 40 | 105.379 | 0 | 23.6657 | 23.3545 | 0 |  | 0.98631 | 23.69 | 23.602 | 0 | 19.58 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6243 | 9.5034
2008-08-05 06:26 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6545 | 23.3545 | 0 |  | 0.98584 | 23.69 | 23.602 | 0 | 19.41 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6254 | 9.5034
2008-08-05 06:27 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.379 | 0 | 23.6441 | 23.3665 | 0 |  | 0.98587 | 23.69 | 23.5897 | 0 | 19.42 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6472 | 9.5034
2008-08-05 06:28 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 105.379 | 0 | 23.6545 | 23.3657 | 0 |  | 0.98587 | 23.69 | 23.602 | 0 | 19.58 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6359 | 9.5034
2008-08-05 06:29 | 0 | 85.02 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6553 | 23.3553 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20.45 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6359 | 9.5034
2008-08-05 06:30 | 0 | 84.95 | 85 | 0 | 39.92 | 40 | 105.379 | 0 | 23.6553 | 23.3553 | 0 |  | 0.98587 | 23.69 | 23.602 | 0 | 21.41 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 06:31 | 0 | 85.03 | 85 | 0 | 39.97 | 40 | 105.379 | 0 | 23.6665 | 23.3553 | 0 |  | 0.98587 | 23.69 | 23.602 | 0 | 21.62 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 06:32 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 105.379 | 0 | 23.6665 | 23.3553 | 0 |  | 0.98587 | 23.69 | 23.602 | 0 | 21.22 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6122 | 9.5034
2008-08-05 06:33 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.379 | 0 | 23.6657 | 23.3545 | 0 |  | 0.98584 | 23.69 | 23.6142 | 0 | 20.6 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 06:34 | 0 | 85.01 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6553 | 23.3545 | 0 |  | 0.98631 | 23.69 | 23.602 | 0 | 20.17 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:35 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 105.379 | 0 | 23.6553 | 23.3657 | 0 |  | 0.98584 | 23.69 | 23.5897 | 0 | 19.83 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6235 | 9.5034
2008-08-05 06:36 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 105.379 | 0 | 23.6545 | 23.3553 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.66 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:37 | 0 | 85.02 | 85 | 0 | 39.9 | 40 | 105.4295 | 0 | 23.6657 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.6142 | 0 | 19.54 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:38 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.4 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:39 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.56 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 06:40 | 0 | 85.03 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.26 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:41 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 21.4 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:42 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 22.08 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:43 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 21.9 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:44 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 21.24 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:45 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6657 | 23.3657 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20.65 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:46 | 0 | 85.01 | 85 | 0 | 39.92 | 40 | 105.4295 | 0 | 23.6657 | 23.3657 | 0 |  | 0.98631 | 23.69 | 23.602 | 0 | 20.16 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:47 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.04 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:48 | 0 | 85.01 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.87 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:49 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.65 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:50 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.54 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:51 | 0 | 84.93 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.49 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:52 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.48 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 06:53 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.67 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 06:54 | 0 | 84.98 | 85 | 0 | 40.03 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20.25 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 06:55 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 20.84 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 06:56 | 0 | 85.02 | 85 | 0 | 39.94 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 21.05 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 06:57 | 0 | 85.01 | 85 | 0 | 39.93 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.74 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 06:58 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.3 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 06:59 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.01 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:00 | 0 | 84.98 | 85 | 0 | 40.03 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.87 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:01 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.67 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:02 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.54 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 07:03 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 105.4295 | 0 | 23.6665 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.602 | 0 | 19.5 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:04 | 0 | 85.01 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.602 | 0 | 20.02 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:05 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 20.56 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6237 | 9.5034
2008-08-05 07:06 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 21.14 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 07:07 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.79 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:08 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.51 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:09 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.14 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:10 | 0 | 84.95 | 85 | 0 | 40.04 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.87 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:11 | 0 | 84.98 | 85 | 0 | 39.9 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.73 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:12 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.57 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:13 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.42 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:14 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.602 | 0 | 19.76 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:15 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.3 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 07:16 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 20.89 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:17 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.602 | 0 | 20.99 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:18 | 0 | 84.93 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.71 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:19 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20.35 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:20 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:21 | 0 | 84.95 | 85 | 0 | 39.92 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.79 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:22 | 0 | 84.98 | 85 | 0 | 39.9 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.7 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 07:23 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.58 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:24 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.41 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:25 | 0 | 84.95 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6665 | 23.3673 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.45 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:26 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6673 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.34 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 07:27 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.602 | 0 | 19.93 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 07:28 | 0 | 84.98 | 85 | 0 | 40.03 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.88 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6237 | 9.5034
2008-08-05 07:29 | 0 | 85.02 | 85 | 0 | 40.03 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 21.55 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:30 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 21.49 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 07:31 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.92 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:32 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.44 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:33 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20.08 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:34 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.86 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:35 | 0 | 84.97 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.79 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:36 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.67 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 07:37 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.5 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:38 | 0 | 84.93 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.38 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 07:39 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.37 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:40 | 0 | 85.01 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 19.25 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 07:41 | 0 | 84.97 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.34 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 07:42 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.602 | 0 | 19.37 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6237 | 9.5034
2008-08-05 07:43 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 19.53 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6237 | 9.5034
2008-08-05 07:44 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 20.7 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 07:45 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 21.45 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 07:46 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 21.39 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 07:47 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.85 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 07:48 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20.43 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:49 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.09 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 07:50 | 0 | 85.01 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.86 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:51 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.66 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 07:52 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.53 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:53 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.5 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 07:54 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20.33 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6237 | 9.5034
2008-08-05 07:55 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6673 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 21.11 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6237 | 9.5034
2008-08-05 07:56 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 21.62 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6237 | 9.5034
2008-08-05 07:57 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 21.57 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 07:58 | 0 | 85.01 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 21.08 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 07:59 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.54 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:00 | 0 | 85.01 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.2 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:01 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.98 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:02 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.78 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:03 | 0 | 84.96 | 85 | 0 | 40.06 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.77 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:04 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:05 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.57 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:06 | 0 | 84.99 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 21.03 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:07 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.85 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:08 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20.53 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:09 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.2 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:10 | 0 | 84.96 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.94 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:11 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.89 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:12 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.84 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:13 | 0 | 84.95 | 85 | 0 | 39.9 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.56 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:14 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.57 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:15 | 0 | 84.91 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6673 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20.02 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 08:16 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 20.58 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6237 | 9.5034
2008-08-05 08:17 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6673 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20.99 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6237 | 9.5034
2008-08-05 08:18 | 0 | 84.93 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.93 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:19 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20.63 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:20 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.28 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:21 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.02 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:22 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.84 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:23 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.5897 | 0 | 19.65 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:24 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.72 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:25 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6665 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 20.59 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:26 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 21.36 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 08:27 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 21.3 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:28 | 0 | 85.02 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.77 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:29 | 0 | 84.97 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20.29 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:30 | 0 | 85.01 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.08 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:31 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.89 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:32 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.75 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:33 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.66 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:34 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.75 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:35 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.1 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:36 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.602 | 0 | 21.03 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 08:37 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 21.5 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6237 | 9.5034
2008-08-05 08:38 | 0 | 84.96 | 85 | 0 | 40.06 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 21.39 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:39 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 20.77 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:40 | 0 | 85.04 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.37 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:41 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.14 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:42 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.94 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:43 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.84 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:44 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.66 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:45 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.66 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:46 | 0 | 84.93 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.54 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:47 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.72 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:48 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.602 | 0 | 20.09 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:49 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.602 | 0 | 20.43 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:50 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6665 | 23.3673 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.77 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:51 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.7 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:52 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6673 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.45 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:53 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.16 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:54 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6673 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.01 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6359 | 9.5034
2008-08-05 08:55 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.86 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:56 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.75 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:57 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.69 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 08:58 | 0 | 85.01 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6553 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.602 | 0 | 19.54 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 08:59 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 23.6673 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.97 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 09:00 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6673 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.602 | 0 | 20.4 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 09:01 | 0 | 85.02 | 85 | 0 | 39.94 | 40 | 105.4295 | 0 | 23.6665 | 23.3673 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 20.73 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6361 | 9.5034
2008-08-05 09:02 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98638 | 23.69 | 23.6142 | 0 | 20.53 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 14.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6369 | 9.5034
2008-08-05 09:03 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6673 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 20.26 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6237 | 9.5034
2008-08-05 09:04 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 23.6665 | 23.3665 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.98 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 09:05 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 23.6665 | 23.3673 | 0 |  | 0.98635 | 23.69 | 23.6142 | 0 | 19.74 | 10.1968 | 0 | 0.42575 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 15 | 15 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.0237 | 18.7923 | -1 | 20.0237 | 23.6245 | 9.5034
2008-08-05 09:06 | 0 | 84.71 | 85 | 0 | 39.83 | 40 | 100.8836 | 0 | 23.6665 | 30.5866 | 0 |  | 0.94395 | 31.01 | 23.6142 | 0 | 19.42 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 14.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7923 | -1 | 20.031 | 23.6369 | 0
2008-08-05 09:07 | 0 | 84.96 | 85 | 0 | 39.9 | 40 | 100.8836 | 0 | 29.2655 | 28.8601 | 0 |  | 0.93061 | 31.01 | 29.1941 | 0 | 19.21 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.1 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7979 | -1 | 20.031 | 30.9619 | 0
2008-08-05 09:08 | 0 | 85.06 | 85 | 0 | 39.96 | 40 | 100.9846 | 0 | 29.2384 | 28.8357 | 0 |  | 0.92982 | 31.01 | 29.1575 | 0 | 19.42 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.64 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7979 | -1 | 20.031 | 30.9422 | 0
2008-08-05 09:09 | 0 | 84.97 | 85 | 0 | 40 | 40 | 100.8836 | 0 | 29.1887 | 28.7882 | 0 |  | 0.92835 | 31.01 | 29.1209 | 0 | 20.18 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.81 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9504 | 0
2008-08-05 09:10 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 100.8836 | 0 | 29.1797 | 28.7794 | 0 |  | 0.92807 | 31.01 | 29.1209 | 0 | 20.54 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 19.89 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9353 | 0
2008-08-05 09:11 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 100.8836 | 0 | 29.1746 | 28.7745 | 0 |  | 0.92791 | 31.01 | 29.1087 | 0 | 20.17 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.94 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9278 | 0
2008-08-05 09:12 | 0 | 84.96 | 85 | 0 | 40 | 40 | 100.8836 | 0 | 29.1716 | 28.7715 | 0 |  | 0.92781 | 31.01 | 29.0965 | 0 | 19.96 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.97 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9299 | 0
2008-08-05 09:13 | 0 | 84.93 | 85 | 0 | 39.93 | 40 | 100.8836 | 0 | 29.1706 | 28.7993 | 0 |  | 0.92778 | 31.01 | 29.0965 | 0 | 19.69 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9459 | 0
2008-08-05 09:14 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 100.8836 | 0 | 29.1686 | 28.7686 | 0 |  | 0.92772 | 31.01 | 29.0965 | 0 | 19.45 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9341 | 0
2008-08-05 09:15 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 100.8836 | 0 | 29.1686 | 28.7676 | 0 |  | 0.92772 | 31.01 | 29.0965 | 0 | 19.37 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9341 | 0
2008-08-05 09:16 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 100.8836 | 0 | 29.1676 | 28.7676 | 0 |  | 0.92769 | 31.01 | 29.0965 | 0 | 19.5 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9351 | 0
2008-08-05 09:17 | 0 | 84.93 | 85 | 0 | 39.9 | 40 | 100.8836 | 0 | 29.1676 | 28.7964 | 0 |  | 0.92769 | 31.01 | 29.0965 | 0 | 19.76 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.917 | 0
2008-08-05 09:18 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 100.8836 | 0 | 29.1676 | 28.7676 | 0 |  | 0.92769 | 31.01 | 29.1087 | 0 | 20.06 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9362 | 0
2008-08-05 09:19 | 0 | 84.93 | 85 | 0 | 40 | 40 | 100.8836 | 0 | 29.1666 | 28.7666 | 0 |  | 0.92766 | 31.01 | 29.1087 | 0 | 20.4 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9491 | 0
2008-08-05 09:20 | 0 | 84.98 | 85 | 0 | 39.92 | 40 | 100.9846 | 0 | 29.1666 | 28.7954 | 0 |  | 0.92766 | 31.01 | 29.1087 | 0 | 20.71 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9491 | 0
2008-08-05 09:21 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 100.8836 | 0 | 29.1666 | 28.7954 | 0 |  | 0.92766 | 31.01 | 29.1087 | 0 | 20.57 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9053 | 0
2008-08-05 09:22 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 100.9846 | 0 | 29.1666 | 28.7666 | 0 |  | 0.92858 | 31.01 | 29.0965 | 0 | 20.16 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9491 | 0
2008-08-05 09:23 | 0 | 84.97 | 85 | 0 | 40 | 40 | 100.8836 | 0 | 29.1666 | 28.7954 | 0 |  | 0.92766 | 31.01 | 29.0965 | 0 | 19.83 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9362 | 0
2008-08-05 09:24 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 100.8836 | 0 | 29.1666 | 28.7954 | 0 |  | 0.92766 | 31.01 | 29.0965 | 0 | 19.62 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9619 | 0
2008-08-05 09:25 | 0 | 85.01 | 85 | 0 | 39.92 | 40 | 100.8836 | 0 | 29.1666 | 28.7954 | 0 |  | 0.92766 | 31.01 | 29.0965 | 0 | 19.44 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9438 | 0
2008-08-05 09:26 | 0 | 84.98 | 85 | 0 | 40 | 40 | 100.9846 | 0 | 29.1666 | 28.7954 | 0 |  | 0.92766 | 31.01 | 29.0965 | 0 | 19.4 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9619 | 0
2008-08-05 09:27 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 100.9846 | 0 | 29.1676 | 28.7964 | 0 |  | 0.92769 | 31.01 | 29.1087 | 0 | 19.34 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9299 | 0
2008-08-05 09:28 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 100.9846 | 0 | 29.1676 | 28.7676 | 0 |  | 0.92862 | 31.01 | 29.1087 | 0 | 19.5 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9608 | 0
2008-08-05 09:29 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 100.8836 | 0 | 29.1676 | 28.7964 | 0 |  | 0.92769 | 31.01 | 29.1087 | 0 | 20.48 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.917 | 0
2008-08-05 09:30 | 0 | 84.91 | 85 | 0 | 39.98 | 40 | 100.8836 | 0 | 29.1676 | 28.7676 | 0 |  | 0.92769 | 31.01 | 29.1087 | 0 | 21.45 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.948 | 0
2008-08-05 09:31 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 100.8836 | 0 | 29.1971 | 28.7676 | 0 |  | 0.92769 | 31.01 | 29.1209 | 0 | 21.58 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7979 | -1 | 20.031 | 30.948 | 0
2008-08-05 09:32 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 100.8836 | 0 | 29.1971 | 28.7676 | 0 |  | 0.92769 | 31.01 | 29.1087 | 0 | 21.01 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7979 | -1 | 20.031 | 30.9299 | 0
2008-08-05 09:33 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 100.8836 | 0 | 29.1971 | 28.7676 | 0 |  | 0.92769 | 31.01 | 29.1209 | 0 | 20.48 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7979 | -1 | 20.031 | 30.9351 | 0
2008-08-05 09:34 | 0 | 84.95 | 85 | 0 | 39.89 | 40 | 100.8836 | 0 | 29.1676 | 28.7964 | 0 |  | 0.92769 | 31.01 | 29.0965 | 0 | 20.04 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9351 | 0
2008-08-05 09:35 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 100.9846 | 0 | 29.1676 | 28.7676 | 0 |  | 0.92862 | 31.01 | 29.1087 | 0 | 19.77 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.948 | 0
2008-08-05 09:36 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 100.8836 | 0 | 29.1676 | 28.7676 | 0 |  | 0.92769 | 31.01 | 29.0965 | 0 | 19.67 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9351 | 0
2008-08-05 09:37 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 100.8836 | 0 | 29.1676 | 28.7964 | 0 |  | 0.92769 | 31.01 | 29.0965 | 0 | 19.54 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.948 | 0
2008-08-05 09:38 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 100.8836 | 0 | 29.1676 | 28.7676 | 0 |  | 0.92769 | 31.01 | 29.0965 | 0 | 19.54 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9288 | 0
2008-08-05 09:39 | 0 | 84.95 | 85 | 0 | 40 | 40 | 100.8836 | 0 | 29.1686 | 28.7676 | 0 |  | 0.92769 | 31.01 | 29.0965 | 0 | 20.35 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9341 | 0
2008-08-05 09:40 | 0 | 84.93 | 85 | 0 | 39.94 | 40 | 100.8836 | 0 | 29.1686 | 28.7974 | 0 |  | 0.92772 | 31.01 | 29.1087 | 0 | 21.18 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9351 | 0
2008-08-05 09:41 | 0 | 84.97 | 85 | 0 | 39.92 | 40 | 100.9846 | 0 | 29.1676 | 28.7676 | 0 |  | 0.92769 | 31.01 | 29.0965 | 0 | 21.32 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9598 | 0
2008-08-05 09:42 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 100.9846 | 0 | 29.1981 | 28.7974 | 0 |  | 0.92772 | 31.01 | 29.1209 | 0 | 20.95 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7979 | -1 | 20.031 | 30.9341 | 0
2008-08-05 09:43 | 0 | 84.96 | 85 | 0 | 40 | 40 | 100.8836 | 0 | 29.1676 | 28.7676 | 0 |  | 0.92769 | 31.01 | 29.1087 | 0 | 20.48 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.948 | 0
2008-08-05 09:44 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 100.8836 | 0 | 29.1676 | 28.7964 | 0 |  | 0.92769 | 31.01 | 29.1087 | 0 | 20.08 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.916 | 0
2008-08-05 09:45 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 100.8836 | 0 | 29.1686 | 28.7974 | 0 |  | 0.92865 | 31.01 | 29.1087 | 0 | 19.85 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.916 | 0
2008-08-05 09:46 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 100.8836 | 0 | 29.1676 | 28.7676 | 0 |  | 0.92769 | 31.01 | 29.1209 | 0 | 19.8 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9341 | 0
2008-08-05 09:47 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 100.8836 | 0 | 29.1686 | 28.7686 | 0 |  | 0.92772 | 31.01 | 29.0965 | 0 | 19.66 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9469 | 0
2008-08-05 09:48 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 100.8836 | 0 | 29.1686 | 28.7686 | 0 |  | 0.92772 | 31.01 | 29.0965 | 0 | 19.52 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9341 | 0
2008-08-05 09:49 | 0 | 84.98 | 85 | 0 | 40 | 40 | 100.8836 | 0 | 29.1686 | 28.7686 | 0 |  | 0.92772 | 31.01 | 29.1087 | 0 | 19.41 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9341 | 0
2008-08-05 09:50 | 0 | 84.98 | 85 | 0 | 40 | 40 | 100.9846 | 0 | 29.1686 | 28.7974 | 0 |  | 0.92772 | 31.01 | 29.0965 | 0 | 19.42 | 0 | 0 | 0 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7978 | -1 | 20.031 | 30.9031 | 0
2008-08-05 09:51 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 29.1981 | 29.9209 | 0 |  | 0.96952 | 31.01 | 29.1087 | 0 | 19.53 | 20.3989 | 0 | 0.85172 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7979 | -1 | 20.031 | 30.9469 | 19.0118
2008-08-05 09:52 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.96952 | 31.01 | 30.4274 | 0 | 20.24 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 09:53 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 21.03 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 09:54 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.4755 | 30.0947 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 21.07 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 09:55 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4151 | 0 | 20.66 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9178 | 19.0117
2008-08-05 09:56 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.3 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 09:57 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.96 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 09:58 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.75 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 09:59 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.6 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:00 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5017 | 30.0927 | 0 |  | 0.97042 | 31.01 | 30.4274 | 0 | 19.54 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9317 | 19.0117
2008-08-05 10:01 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.5 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:02 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.16 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:03 | 0 | 84.93 | 85 | 0 | 39.93 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4151 | 0 | 21.07 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:04 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 21.65 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:05 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97042 | 31.01 | 30.4274 | 0 | 21.31 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.917 | 19.0117
2008-08-05 10:06 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5168 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4396 | 0 | 20.72 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9298 | 19.0117
2008-08-05 10:07 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5017 | 30.0927 | 0 |  | 0.97042 | 31.01 | 30.4274 | 0 | 20.29 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9189 | 19.0117
2008-08-05 10:08 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.5306 | 0 | 30.5017 | 30.0927 | 0 |  | 0.96949 | 31.01 | 30.4274 | 0 | 20 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9317 | 19.0117
2008-08-05 10:09 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.5017 | 30.0639 | 0 |  | 0.97042 | 31.01 | 30.4274 | 0 | 19.74 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9317 | 19.0117
2008-08-05 10:10 | 0 | 85.02 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.5017 | 30.0927 | 0 |  | 0.97042 | 31.01 | 30.4274 | 0 | 19.72 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9614 | 19.0117
2008-08-05 10:11 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4151 | 0 | 19.53 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 10:12 | 0 | 84.98 | 85 | 0 | 39.92 | 40 | 105.4295 | 0 | 30.5027 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4151 | 0 | 19.9 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9475 | 19.0117
2008-08-05 10:13 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0649 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.63 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:14 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 30.4735 | 30.0927 | 0 |  | 0.96949 | 31.01 | 30.4274 | 0 | 21.03 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 10:15 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 30.4745 | 30.0927 | 0 |  | 0.96949 | 31.01 | 30.4151 | 0 | 21.08 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 10:16 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5017 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 20.62 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 10:17 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.5017 | 30.0639 | 0 |  | 0.96949 | 31.01 | 30.4274 | 0 | 20.3 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9475 | 19.0117
2008-08-05 10:18 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.4735 | 30.0639 | 0 |  | 0.96949 | 31.01 | 30.4029 | 0 | 19.96 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9317 | 19.0117
2008-08-05 10:19 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 30.5027 | 30.0649 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.8 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:20 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 30.5017 | 30.0639 | 0 |  | 0.96952 | 31.01 | 30.4274 | 0 | 19.65 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9317 | 19.0117
2008-08-05 10:21 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 105.5306 | 0 | 30.4745 | 30.0927 | 0 |  | 0.97042 | 31.01 | 30.4029 | 0 | 19.53 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 10:22 | 0 | 85.02 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.54 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.917 | 19.0117
2008-08-05 10:23 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5017 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.63 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9317 | 19.0117
2008-08-05 10:24 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.13 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:25 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 105.5306 | 0 | 30.5027 | 30.1081 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.64 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0116
2008-08-05 10:26 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.64 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:27 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.41 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9436 | 19.0117
2008-08-05 10:28 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.5811 | 0 | 30.5027 | 30.1081 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.05 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9159 | 19.0116
2008-08-05 10:29 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.77 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9436 | 19.0117
2008-08-05 10:30 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.5306 | 0 | 30.5168 | 30.0937 | 0 |  | 0.97092 | 31.01 | 30.4274 | 0 | 19.72 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:31 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 105.5811 | 0 | 30.5168 | 30.1081 | 0 |  | 0.97092 | 31.01 | 30.4396 | 0 | 19.52 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9287 | 19.0116
2008-08-05 10:32 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.5811 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97092 | 31.01 | 30.4274 | 0 | 19.5 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9436 | 19.0117
2008-08-05 10:33 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.3 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:34 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.4 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:35 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 105.5306 | 0 | 30.5178 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4396 | 0 | 19.9 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 10:36 | 0 | 85.01 | 85 | 0 | 39.94 | 40 | 105.5306 | 0 | 30.5037 | 30.1091 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 20.82 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9148 | 19.0116
2008-08-05 10:37 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.5306 | 0 | 30.5027 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 21.1 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 10:38 | 0 | 84.96 | 85 | 0 | 40 | 40 | 105.5306 | 0 | 30.4745 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 20.82 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 10:39 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.36 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:40 | 0 | 85.01 | 85 | 0 | 39.94 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4151 | 0 | 20.05 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 10:41 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 105.5306 | 0 | 30.4745 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 19.76 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 10:42 | 0 | 84.95 | 85 | 0 | 40 | 40 | 105.5306 | 0 | 30.4745 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 19.63 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 10:43 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.4745 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.3907 | 0 | 19.54 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:44 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0649 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.49 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9475 | 19.0117
2008-08-05 10:45 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.33 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 10:46 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 19.4 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 10:47 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 30.4755 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4151 | 0 | 19.56 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9039 | 19.0117
2008-08-05 10:48 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.96956 | 31.01 | 30.4274 | 0 | 19.81 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:49 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 105.5306 | 0 | 30.4755 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4151 | 0 | 20.25 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 10:50 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.5306 | 0 | 30.4755 | 30.0659 | 0 |  | 0.97049 | 31.01 | 30.4029 | 0 | 20.69 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 10:51 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 30.4755 | 30.0659 | 0 |  | 0.97049 | 31.01 | 30.4151 | 0 | 20.6 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9039 | 19.0117
2008-08-05 10:52 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.18 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:53 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 19.86 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:54 | 0 | 85.02 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.66 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 10:55 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.5 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 10:56 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 19.42 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:57 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.5306 | 0 | 30.5027 | 30.0947 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.28 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 10:58 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.36 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 10:59 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4151 | 0 | 19.26 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 11:00 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.5178 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4396 | 0 | 19.57 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9148 | 19.0117
2008-08-05 11:01 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 105.5811 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 20.06 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9148 | 19.0117
2008-08-05 11:02 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.5811 | 0 | 30.5178 | 30.1091 | 0 |  | 0.97095 | 31.01 | 30.4396 | 0 | 20.59 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9148 | 19.0116
2008-08-05 11:03 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 105.5306 | 0 | 30.5178 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4396 | 0 | 20.89 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 11:04 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.5306 | 0 | 30.5178 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4396 | 0 | 20.53 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9277 | 19.0117
2008-08-05 11:05 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.5811 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.25 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9159 | 19.0117
2008-08-05 11:06 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.89 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9436 | 19.0117
2008-08-05 11:07 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.72 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 11:08 | 0 | 85.03 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.58 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 11:09 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.4745 | 30.0937 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.48 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 11:10 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.5306 | 0 | 30.4745 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 19.38 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 11:11 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 105.5306 | 0 | 30.5037 | 30.0659 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 19.65 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9178 | 19.0117
2008-08-05 11:12 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 105.5306 | 0 | 30.5027 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 19.96 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 11:13 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.5306 | 0 | 30.4755 | 30.0947 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 20.53 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9335 | 19.0117
2008-08-05 11:14 | 0 | 84.93 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.96956 | 31.01 | 30.4274 | 0 | 20.88 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 11:15 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.3907 | 0 | 20.65 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 11:16 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 20.32 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:17 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 20.04 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9475 | 19.0117
2008-08-05 11:18 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.74 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:19 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.54 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:20 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.34 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:21 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 30.4755 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.37 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:22 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 105.5306 | 0 | 30.5027 | 30.0947 | 0 |  | 0.96956 | 31.01 | 30.4274 | 0 | 19.3 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 11:23 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 105.5306 | 0 | 30.5037 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4151 | 0 | 19.53 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9039 | 19.0117
2008-08-05 11:24 | 0 | 85.01 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 30.4755 | 30.0947 | 0 |  | 0.96956 | 31.01 | 30.4151 | 0 | 19.97 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9464 | 19.0117
2008-08-05 11:25 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4151 | 0 | 20.59 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9039 | 19.0117
2008-08-05 11:26 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 30.4755 | 30.0659 | 0 |  | 0.96952 | 31.01 | 30.4274 | 0 | 20.61 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 11:27 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 20.28 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9039 | 19.0117
2008-08-05 11:28 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.4745 | 30.0937 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.89 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9475 | 19.0117
2008-08-05 11:29 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 30.4755 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4151 | 0 | 19.68 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9475 | 19.0117
2008-08-05 11:30 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.5027 | 30.0937 | 0 |  | 0.96952 | 31.01 | 30.4151 | 0 | 19.77 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:31 | 0 | 85.01 | 85 | 0 | 40 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.63 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 11:32 | 0 | 84.99 | 85 | 0 | 40.04 | 40 | 105.4295 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 19.54 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 11:33 | 0 | 84.99 | 85 | 0 | 40 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4151 | 0 | 19.58 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 11:34 | 0 | 84.93 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 30.5037 | 30.0947 | 0 |  | 0.96956 | 31.01 | 30.4274 | 0 | 20.09 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 11:35 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 30.4755 | 30.0947 | 0 |  | 0.96956 | 31.01 | 30.4029 | 0 | 20.68 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 11:36 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 105.5306 | 0 | 30.5037 | 30.0659 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 20.93 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9039 | 19.0117
2008-08-05 11:37 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 20.58 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:38 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 20.09 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:39 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.84 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:40 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.68 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:41 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.58 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:42 | 0 | 84.95 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.83 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:43 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.4755 | 30.0659 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 20.73 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:44 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96956 | 31.01 | 30.4029 | 0 | 21.48 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:45 | 0 | 84.97 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 30.5027 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4151 | 0 | 21.32 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:46 | 0 | 85.01 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 20.88 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:47 | 0 | 84.95 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 20.37 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:48 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 30.4735 | 30.0639 | 0 |  | 0.96949 | 31.01 | 30.4151 | 0 | 20.08 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 11:49 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.76 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 11:50 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 30.5027 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4274 | 0 | 19.61 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:51 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.84 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 11:52 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.73 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 11:53 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 21.19 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 11:54 | 0 | 84.93 | 85 | 0 | 39.93 | 40 | 105.4295 | 0 | 30.4745 | 30.0937 | 0 |  | 0.96952 | 31.01 | 30.4151 | 0 | 21.1 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 11:55 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.59 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 11:56 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.4745 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 20.22 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9178 | 19.0117
2008-08-05 11:57 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5027 | 30.0649 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.97 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 11:58 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.83 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 11:59 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.4745 | 30.0649 | 0 |  | 0.97045 | 31.01 | 30.4151 | 0 | 19.58 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:00 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.61 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:01 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.4755 | 30.0937 | 0 |  | 0.96952 | 31.01 | 30.4151 | 0 | 19.56 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9593 | 19.0117
2008-08-05 12:02 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.4745 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 19.45 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:03 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.42 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:04 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 19.78 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9039 | 19.0117
2008-08-05 12:05 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.5037 | 30.0659 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 20.83 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 12:06 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 21.41 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 12:07 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.4755 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 21.14 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 12:08 | 0 | 84.98 | 85 | 0 | 40 | 40 | 105.5306 | 0 | 30.4755 | 30.0659 | 0 |  | 0.96952 | 31.01 | 30.4151 | 0 | 20.62 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9168 | 19.0117
2008-08-05 12:09 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.4745 | 30.0649 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 20.18 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 12:10 | 0 | 84.96 | 85 | 0 | 40 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 19.92 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9178 | 19.0117
2008-08-05 12:11 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.81 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 12:12 | 0 | 85.01 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.4745 | 30.0649 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 19.66 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 12:13 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.79 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 12:14 | 0 | 84.93 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 30.5037 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 20.69 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 12:15 | 0 | 85.02 | 85 | 0 | 40 | 40 | 105.5306 | 0 | 30.4745 | 30.0937 | 0 |  | 0.97049 | 31.01 | 30.4029 | 0 | 21.27 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 12:16 | 0 | 84.98 | 85 | 0 | 40.04 | 40 | 105.5306 | 0 | 30.5037 | 30.0659 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 21.22 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9346 | 19.0117
2008-08-05 12:17 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 30.4745 | 30.0937 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 20.74 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:18 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.5027 | 30.0937 | 0 |  | 0.96952 | 31.01 | 30.4274 | 0 | 20.39 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:19 | 0 | 84.99 | 85 | 0 | 40 | 40 | 105.5306 | 0 | 30.5027 | 30.0649 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9317 | 19.0117
2008-08-05 12:20 | 0 | 84.96 | 85 | 0 | 40 | 40 | 105.5306 | 0 | 30.4745 | 30.0649 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.85 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:21 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 30.5027 | 30.0937 | 0 |  | 0.96952 | 31.01 | 30.4274 | 0 | 19.71 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:22 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 19.69 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9178 | 19.0117
2008-08-05 12:23 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 105.5306 | 0 | 30.5027 | 30.0649 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.48 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 12:24 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 21.39 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 12:25 | 0 | 85.01 | 85 | 0 | 39.94 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 21.47 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:26 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.95 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:27 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.45 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 12:28 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 105.5306 | 0 | 30.4745 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.13 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 12:29 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.87 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:30 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.78 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:31 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.6 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:32 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 30.4745 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.7 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:33 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.06 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:34 | 0 | 85.01 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.47 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:35 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 20.49 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:36 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.16 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7473 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:37 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.93 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:38 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.61 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:39 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.96956 | 31.01 | 30.4274 | 0 | 19.56 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 12:40 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.96952 | 31.01 | 30.4274 | 0 | 19.54 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 12:41 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 105.5306 | 0 | 30.4755 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 19.56 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 12:42 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 19.57 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9464 | 19.0117
2008-08-05 12:43 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 19.56 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 12:44 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.96956 | 31.01 | 30.4396 | 0 | 19.94 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 12:45 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 105.5306 | 0 | 30.4755 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4151 | 0 | 20.93 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9039 | 19.0117
2008-08-05 12:46 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.4755 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4029 | 0 | 21.36 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9039 | 19.0117
2008-08-05 12:47 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 105.5306 | 0 | 30.4755 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 21.26 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 12:48 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.5027 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4274 | 0 | 20.59 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:49 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.04 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:50 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.96952 | 31.01 | 30.4151 | 0 | 19.75 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 12:51 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.58 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:52 | 0 | 84.95 | 85 | 0 | 39.96 | 40 | 105.4295 | 0 | 30.4745 | 30.0937 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.48 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:53 | 0 | 84.93 | 85 | 0 | 40 | 40 | 105.4295 | 0 | 30.5027 | 30.0649 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 19.45 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9168 | 19.0117
2008-08-05 12:54 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 105.4295 | 0 | 30.4755 | 30.0659 | 0 |  | 0.96956 | 31.01 | 30.4029 | 0 | 19.57 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9039 | 19.0117
2008-08-05 12:55 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5037 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4029 | 0 | 20.49 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7179 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9039 | 19.0117
2008-08-05 12:56 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 30.5037 | 30.0947 | 0 |  | 0.96956 | 31.01 | 30.4274 | 0 | 21.23 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7399 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9593 | 19.0117
2008-08-05 12:57 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 105.4295 | 0 | 30.4755 | 30.0947 | 0 |  | 0.96956 | 31.01 | 30.4151 | 0 | 21.28 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9296 | 19.0117
2008-08-05 12:58 | 0 | 84.98 | 85 | 0 | 39.9 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.72 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 12:59 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 20.3 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.905 | 19.0117
2008-08-05 13:00 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 105.4295 | 0 | 30.4745 | 30.0937 | 0 |  | 0.96952 | 31.01 | 30.4029 | 0 | 20 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 13:01 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 105.5306 | 0 | 30.5027 | 30.0937 | 0 |  | 0.97045 | 31.01 | 30.4274 | 0 | 19.8 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 13:02 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.5027 | 30.0649 | 0 |  | 0.97045 | 31.01 | 30.4029 | 0 | 19.58 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9307 | 19.0117
2008-08-05 13:03 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 105.4295 | 0 | 30.4745 | 30.0649 | 0 |  | 0.97049 | 31.01 | 30.4151 | 0 | 19.46 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 13:04 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 105.5306 | 0 | 30.5037 | 30.0659 | 0 |  | 0.97049 | 31.01 | 30.4151 | 0 | 20.2 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7253 | 0 | 0 | 0 | 20 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7992 | -1 | 20.031 | 30.9603 | 19.0117
2008-08-05 13:05 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 105.5306 | 0 | 30.4755 | 30.0947 | 0 |  | 0.97049 | 31.01 | 30.4274 | 0 | 21.35 | 20.3988 | 0 | 0.85171 | 20 | 18.7687 |  | 18.7326 | 0 | 0 | 0 | 20 | 19.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 20.031 | 18.7991 | -1 | 20.031 | 30.9039 | 19.0117
2008-08-05 13:06 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 101.2372 | 0 | 30.4755 | 29.9486 | 0 |  | 0.96952 | 30.89 | 30.4029 | 0 | 21.71 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.9524 | 0 | 0 | 0 | 25 | 20 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1162 | -1 | 15.0309 | 30.9464 | 0
2008-08-05 13:07 | 0 | 85.04 | 85 | 0 | 39.95 | 40 | 100.9846 | 0 | 28.6961 | 28.3036 | 0 |  | 0.91624 | 30.89 | 28.6447 | 0 | 21.2 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 23.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1144 | -1 | 15.0309 | 30.8269 | 0
2008-08-05 13:08 | 0 | 85.02 | 85 | 0 | 39.99 | 40 | 100.9846 | 0 | 28.6386 | 28.2504 | 0 |  | 0.91455 | 30.89 | 28.5592 | 0 | 20.44 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.53 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1144 | -1 | 15.0309 | 30.8131 | 0
2008-08-05 13:09 | 0 | 85.1 | 85 | 0 | 39.98 | 40 | 100.9846 | 0 | 28.6124 | 28.2257 | 0 |  | 0.91372 | 30.89 | 28.547 | 0 | 20.02 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.79 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8311 | 0
2008-08-05 13:10 | 0 | 85.06 | 85 | 0 | 39.98 | 40 | 100.9846 | 0 | 28.6027 | 28.2153 | 0 |  | 0.91341 | 30.89 | 28.5348 | 0 | 19.75 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.89 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8174 | 0
2008-08-05 13:11 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 100.9846 | 0 | 28.5979 | 28.2115 | 0 |  | 0.91326 | 30.89 | 28.5226 | 0 | 19.5 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.94 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8096 | 0
2008-08-05 13:12 | 0 | 84.98 | 85 | 0 | 40 | 40 | 100.9846 | 0 | 28.5949 | 28.2087 | 0 |  | 0.9132 | 30.89 | 28.5348 | 0 | 19.4 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8257 | 0
2008-08-05 13:13 | 0 | 85.03 | 85 | 0 | 39.99 | 40 | 100.9846 | 0 | 28.593 | 28.2068 | 0 |  | 0.91314 | 30.89 | 28.5226 | 0 | 19.17 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8147 | 0
2008-08-05 13:14 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 100.9846 | 0 | 28.5631 | 28.2058 | 0 |  | 0.91311 | 30.89 | 28.4982 | 0 | 19.17 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8205 | 0
2008-08-05 13:15 | 0 | 85.02 | 85 | 0 | 40.04 | 40 | 100.9846 | 0 | 28.5622 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.5226 | 0 | 19.1 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8205 | 0
2008-08-05 13:16 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 100.9846 | 0 | 28.5911 | 28.2049 | 0 |  | 0.91216 | 30.89 | 28.5226 | 0 | 19.17 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8085 | 0
2008-08-05 13:17 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 100.8836 | 0 | 28.5901 | 28.2039 | 0 |  | 0.91307 | 30.89 | 28.5226 | 0 | 19.24 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8168 | 0
2008-08-05 13:18 | 0 | 84.96 | 85 | 0 | 39.92 | 40 | 100.9846 | 0 | 28.5901 | 28.2039 | 0 |  | 0.91304 | 30.89 | 28.5226 | 0 | 20.18 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8178 | 0
2008-08-05 13:19 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 100.9846 | 0 | 28.5911 | 28.2049 | 0 |  | 0.91304 | 30.89 | 28.5226 | 0 | 21.07 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8487 | 0
2008-08-05 13:20 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 100.9846 | 0 | 28.5901 | 28.2039 | 0 |  | 0.91304 | 30.89 | 28.5226 | 0 | 21.18 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8309 | 0
2008-08-05 13:21 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 100.9846 | 0 | 28.5901 | 28.2039 | 0 |  | 0.91301 | 30.89 | 28.5226 | 0 | 20.75 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8178 | 0
2008-08-05 13:22 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 100.9846 | 0 | 28.5891 | 28.203 | 0 |  | 0.91301 | 30.89 | 28.5104 | 0 | 20.1 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8189 | 0
2008-08-05 13:23 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 100.9846 | 0 | 28.5891 | 28.203 | 0 |  | 0.91301 | 30.89 | 28.5226 | 0 | 19.85 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25.02 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8309 | 0
2008-08-05 13:24 | 0 | 85.01 | 85 | 0 | 39.94 | 40 | 100.9846 | 0 | 28.5901 | 28.2039 | 0 |  | 0.91304 | 30.89 | 28.5226 | 0 | 19.5 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8497 | 0
2008-08-05 13:25 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 100.9846 | 0 | 28.5901 | 28.203 | 0 |  | 0.91301 | 30.89 | 28.5226 | 0 | 19.38 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8178 | 0
2008-08-05 13:26 | 0 | 84.93 | 85 | 0 | 39.94 | 40 | 100.9846 | 0 | 28.5901 | 28.2039 | 0 |  | 0.91304 | 30.89 | 28.5348 | 0 | 19.45 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8178 | 0
2008-08-05 13:27 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 100.8836 | 0 | 28.5622 | 28.2039 | 0 |  | 0.91304 | 30.89 | 28.4982 | 0 | 20.25 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8178 | 0
2008-08-05 13:28 | 0 | 85.01 | 85 | 0 | 40.06 | 40 | 100.9846 | 0 | 28.5901 | 28.2039 | 0 |  | 0.91213 | 30.89 | 28.5226 | 0 | 21 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8168 | 0
2008-08-05 13:29 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 100.9846 | 0 | 28.5612 | 28.2039 | 0 |  | 0.91304 | 30.89 | 28.5104 | 0 | 20.99 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8178 | 0
2008-08-05 13:30 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 100.9846 | 0 | 28.5612 | 28.2039 | 0 |  | 0.91304 | 30.89 | 28.5226 | 0 | 20.44 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8487 | 0
2008-08-05 13:31 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 100.9846 | 0 | 28.5901 | 28.2039 | 0 |  | 0.91304 | 30.89 | 28.5348 | 0 | 20.04 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8477 | 0
2008-08-05 13:32 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 100.8836 | 0 | 28.5901 | 28.2049 | 0 |  | 0.91304 | 30.89 | 28.5226 | 0 | 19.69 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8487 | 0
2008-08-05 13:33 | 0 | 84.99 | 85 | 0 | 39.93 | 40 | 100.9846 | 0 | 28.5901 | 28.2039 | 0 |  | 0.91304 | 30.89 | 28.5226 | 0 | 19.57 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8168 | 0
2008-08-05 13:34 | 0 | 84.96 | 85 | 0 | 40 | 40 | 100.8836 | 0 | 28.5911 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.5226 | 0 | 19.37 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8168 | 0
2008-08-05 13:35 | 0 | 84.98 | 85 | 0 | 40.06 | 40 | 100.9846 | 0 | 28.5911 | 28.2039 | 0 |  | 0.91307 | 30.89 | 28.5226 | 0 | 19.36 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8168 | 0
2008-08-05 13:36 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 100.9846 | 0 | 28.5911 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.5226 | 0 | 19.92 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8168 | 0
2008-08-05 13:37 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 100.9846 | 0 | 28.5911 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.5226 | 0 | 20.57 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8158 | 0
2008-08-05 13:38 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 100.8836 | 0 | 28.5911 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.5226 | 0 | 20.95 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8168 | 0
2008-08-05 13:39 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 100.8836 | 0 | 28.5911 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.5226 | 0 | 20.41 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8168 | 0
2008-08-05 13:40 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 100.9846 | 0 | 28.5911 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.5226 | 0 | 20.06 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8168 | 0
2008-08-05 13:41 | 0 | 84.93 | 85 | 0 | 40.03 | 40 | 100.9846 | 0 | 28.5911 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.5226 | 0 | 19.85 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8168 | 0
2008-08-05 13:42 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 100.9846 | 0 | 28.5622 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.5104 | 0 | 19.49 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8477 | 0
2008-08-05 13:43 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 100.9846 | 0 | 28.5622 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.5226 | 0 | 19.44 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8168 | 0
2008-08-05 13:44 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 100.9846 | 0 | 28.5911 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.5226 | 0 | 19.34 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8299 | 0
2008-08-05 13:45 | 0 | 85.01 | 85 | 0 | 39.93 | 40 | 100.9846 | 0 | 28.5911 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.5226 | 0 | 19.25 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8158 | 0
2008-08-05 13:46 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 100.9846 | 0 | 28.592 | 28.2058 | 0 |  | 0.91311 | 30.89 | 28.5226 | 0 | 19.84 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8158 | 0
2008-08-05 13:47 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 100.8836 | 0 | 28.592 | 28.2058 | 0 |  | 0.91311 | 30.89 | 28.486 | 0 | 20.22 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8158 | 0
2008-08-05 13:48 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 100.9846 | 0 | 28.592 | 28.2058 | 0 |  | 0.91311 | 30.89 | 28.5226 | 0 | 20.62 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8158 | 0
2008-08-05 13:49 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 100.9846 | 0 | 28.592 | 28.2058 | 0 |  | 0.91311 | 30.89 | 28.5226 | 0 | 20.36 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8158 | 0
2008-08-05 13:50 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 100.9846 | 0 | 28.5911 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.4982 | 0 | 20.08 | 0 | 0 | 0 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8168 | 0
2008-08-05 13:51 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 103.7627 | 0 | 28.5911 | 28.2049 | 0 |  | 0.91307 | 30.89 | 28.5348 | 0 | 19.8 | 40.8073 | 0 | 1.7038 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1143 | -1 | 15.0309 | 30.8158 | 38.0324
2008-08-05 13:52 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 103.7627 | 0 | 29.3871 | 28.9818 | 0 |  | 0.93822 | 30.89 | 29.3162 | 0 | 19.49 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 13:53 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 103.7627 | 0 | 29.3871 | 28.9959 | 0 |  | 0.93822 | 30.89 | 29.304 | 0 | 19.3 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8474 | 38.0304
2008-08-05 13:54 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.3162 | 0 | 19.21 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 13:55 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.3162 | 0 | 19.09 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8163 | 38.0304
2008-08-05 13:56 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 103.7627 | 0 | 29.3871 | 28.9959 | 0 |  | 0.93865 | 30.89 | 29.3162 | 0 | 19.16 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 13:57 | 0 | 85.01 | 85 | 0 | 39.96 | 40 | 103.7627 | 0 | 29.3871 | 28.9818 | 0 |  | 0.93822 | 30.89 | 29.3162 | 0 | 19.14 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8024 | 38.0304
2008-08-05 13:58 | 0 | 84.93 | 85 | 0 | 39.93 | 40 | 103.7627 | 0 | 29.3871 | 28.9818 | 0 |  | 0.93822 | 30.89 | 29.3162 | 0 | 20 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8163 | 38.0304
2008-08-05 13:59 | 0 | 84.98 | 85 | 0 | 40.04 | 40 | 103.7627 | 0 | 29.3871 | 28.9818 | 0 |  | 0.93822 | 30.89 | 29.3162 | 0 | 20.72 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8036 | 38.0304
2008-08-05 14:00 | 0 | 84.95 | 85 | 0 | 40.06 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.3162 | 0 | 20.89 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8013 | 38.0304
2008-08-05 14:01 | 0 | 84.98 | 85 | 0 | 39.93 | 40 | 103.7627 | 0 | 29.3861 | 28.9949 | 0 |  | 0.93819 | 30.89 | 29.3162 | 0 | 20.33 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 14:02 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 103.8132 | 0 | 29.4005 | 28.9949 | 0 |  | 0.93865 | 30.89 | 29.3284 | 0 | 19.82 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8024 | 38.0304
2008-08-05 14:03 | 0 | 84.95 | 85 | 0 | 40 | 40 | 103.8132 | 0 | 29.4005 | 28.9949 | 0 |  | 0.93865 | 30.89 | 29.3284 | 0 | 19.58 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8301 | 38.0304
2008-08-05 14:04 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 103.7627 | 0 | 29.4005 | 28.9949 | 0 |  | 0.93865 | 30.89 | 29.3284 | 0 | 19.42 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8151 | 38.0304
2008-08-05 14:05 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 103.8132 | 0 | 29.4005 | 28.9949 | 0 |  | 0.93865 | 30.89 | 29.3162 | 0 | 19.21 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8301 | 38.0304
2008-08-05 14:06 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 103.8132 | 0 | 29.3861 | 28.9949 | 0 |  | 0.93865 | 30.89 | 29.3284 | 0 | 19.24 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8301 | 38.0304
2008-08-05 14:07 | 0 | 84.99 | 85 | 0 | 40.03 | 40 | 103.8132 | 0 | 29.3861 | 28.9949 | 0 |  | 0.93865 | 30.89 | 29.3162 | 0 | 19.17 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8024 | 38.0304
2008-08-05 14:08 | 0 | 85.01 | 85 | 0 | 40.02 | 40 | 103.7627 | 0 | 29.4005 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.3284 | 0 | 19.89 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8151 | 38.0304
2008-08-05 14:09 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 103.8132 | 0 | 29.4005 | 28.9949 | 0 |  | 0.93865 | 30.89 | 29.3284 | 0 | 20.77 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8151 | 38.0304
2008-08-05 14:10 | 0 | 85.01 | 85 | 0 | 40.02 | 40 | 103.8132 | 0 | 29.3861 | 28.9949 | 0 |  | 0.93865 | 30.89 | 29.3162 | 0 | 21.1 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8301 | 38.0304
2008-08-05 14:11 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 103.7627 | 0 | 29.3851 | 28.9939 | 0 |  | 0.93862 | 30.89 | 29.3162 | 0 | 20.61 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 14:12 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 103.8132 | 0 | 29.3995 | 28.9798 | 0 |  | 0.93862 | 30.89 | 29.3162 | 0 | 20.28 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8311 | 38.0304
2008-08-05 14:13 | 0 | 84.96 | 85 | 0 | 40 | 40 | 103.7627 | 0 | 29.3995 | 28.9798 | 0 |  | 0.93816 | 30.89 | 29.3162 | 0 | 19.86 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8034 | 38.0304
2008-08-05 14:14 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 103.8132 | 0 | 29.3995 | 28.9939 | 0 |  | 0.93862 | 30.89 | 29.3162 | 0 | 19.6 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8161 | 38.0304
2008-08-05 14:15 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 103.7627 | 0 | 29.3851 | 28.9798 | 0 |  | 0.93816 | 30.89 | 29.3162 | 0 | 19.46 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8311 | 38.0304
2008-08-05 14:16 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 103.7627 | 0 | 29.3861 | 28.9949 | 0 |  | 0.93819 | 30.89 | 29.3284 | 0 | 19.3 | 40.8051 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8184 | 38.0304
2008-08-05 14:17 | 0 | 84.99 | 85 | 0 | 39.9 | 40 | 103.7627 | 0 | 29.3851 | 28.9798 | 0 |  | 0.93816 | 30.89 | 29.3162 | 0 | 19.22 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.793 | 38.0304
2008-08-05 14:18 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 103.6617 | 0 | 29.3861 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.3162 | 0 | 19.58 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0305
2008-08-05 14:19 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.3162 | 0 | 20.24 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.822 | 38.0304
2008-08-05 14:20 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.304 | 0 | 20.75 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 14:21 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 103.7627 | 0 | 29.3851 | 28.9798 | 0 |  | 0.93728 | 30.89 | 29.3162 | 0 | 20.67 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 14:22 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 103.7627 | 0 | 29.3861 | 28.9526 | 0 |  | 0.93819 | 30.89 | 29.2796 | 0 | 20.35 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8301 | 38.0305
2008-08-05 14:23 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 103.7627 | 0 | 29.3562 | 28.9516 | 0 |  | 0.93816 | 30.89 | 29.3162 | 0 | 19.9 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8184 | 38.0305
2008-08-05 14:24 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 103.7627 | 0 | 29.3851 | 28.9798 | 0 |  | 0.93816 | 30.89 | 29.3162 | 0 | 19.52 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8184 | 38.0304
2008-08-05 14:25 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3851 | 28.9516 | 0 |  | 0.93725 | 30.89 | 29.304 | 0 | 19.37 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8184 | 38.0305
2008-08-05 14:26 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 103.7627 | 0 | 29.3572 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.2796 | 0 | 19.22 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.822 | 38.0304
2008-08-05 14:27 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93816 | 30.89 | 29.2796 | 0 | 19.1 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0304
2008-08-05 14:28 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3861 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.304 | 0 | 19.13 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8347 | 38.0305
2008-08-05 14:29 | 0 | 85.01 | 85 | 0 | 39.99 | 40 | 103.7627 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93819 | 30.89 | 29.2796 | 0 | 19.17 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 14:30 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3861 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.304 | 0 | 19.5 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7793 | 38.0305
2008-08-05 14:31 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.85 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 14:32 | 0 | 84.95 | 85 | 0 | 39.92 | 40 | 103.7627 | 0 | 29.3572 | 28.9808 | 0 |  | 0.93728 | 30.89 | 29.304 | 0 | 20.24 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7793 | 38.0304
2008-08-05 14:33 | 0 | 84.96 | 85 | 0 | 39.9 | 40 | 103.7627 | 0 | 29.3581 | 28.9818 | 0 |  | 0.93822 | 30.89 | 29.304 | 0 | 20.61 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 14:34 | 0 | 85.02 | 85 | 0 | 39.98 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.3162 | 0 | 20.28 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 14:35 | 0 | 84.98 | 85 | 0 | 40.03 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.3162 | 0 | 19.98 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 14:36 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.3162 | 0 | 19.56 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 14:37 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.3162 | 0 | 19.5 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 14:38 | 0 | 84.95 | 85 | 0 | 40.04 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.3162 | 0 | 19.24 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 14:39 | 0 | 84.93 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3871 | 28.9818 | 0 |  | 0.93731 | 30.89 | 29.3162 | 0 | 19.14 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8474 | 38.0304
2008-08-05 14:40 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93822 | 30.89 | 29.3162 | 0 | 19.09 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 14:41 | 0 | 85.01 | 85 | 0 | 40.03 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.3162 | 0 | 19.09 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7782 | 38.0304
2008-08-05 14:42 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 103.7627 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93822 | 30.89 | 29.3162 | 0 | 19.05 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8336 | 38.0305
2008-08-05 14:43 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 103.7627 | 0 | 29.3861 | 28.9818 | 0 |  | 0.93822 | 30.89 | 29.304 | 0 | 18.84 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8464 | 38.0304
2008-08-05 14:44 | 0 | 84.99 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 18.89 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8464 | 38.0305
2008-08-05 14:45 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 103.7627 | 0 | 29.3871 | 28.9818 | 0 |  | 0.93822 | 30.89 | 29.2796 | 0 | 19.24 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8464 | 38.0304
2008-08-05 14:46 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3581 | 28.9818 | 0 |  | 0.93822 | 30.89 | 29.304 | 0 | 20.25 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8464 | 38.0304
2008-08-05 14:47 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 103.6617 | 0 | 29.3581 | 28.9818 | 0 |  | 0.93731 | 30.89 | 29.2918 | 0 | 20.91 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0304
2008-08-05 14:48 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 103.6617 | 0 | 29.3581 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 20.53 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 14:49 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 20.08 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 14:50 | 0 | 84.98 | 85 | 0 | 40.04 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.58 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 14:51 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.28 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 14:52 | 0 | 84.93 | 85 | 0 | 39.92 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.1 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 14:53 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.02 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 14:54 | 0 | 84.98 | 85 | 0 | 40.04 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2918 | 0 | 18.97 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 14:55 | 0 | 84.93 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3581 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 18.96 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 14:56 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 18.98 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 14:57 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 19.38 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 14:58 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2918 | 0 | 20.37 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7955 | 38.0305
2008-08-05 14:59 | 0 | 84.99 | 85 | 0 | 39.94 | 40 | 103.7627 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 21.07 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:00 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 20.76 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:01 | 0 | 84.95 | 85 | 0 | 39.92 | 40 | 103.6617 | 0 | 29.3581 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 20.28 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:02 | 0 | 84.98 | 85 | 0 | 39.92 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.3162 | 0 | 19.7 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 15:03 | 0 | 84.97 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.304 | 0 | 19.5 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8474 | 38.0305
2008-08-05 15:04 | 0 | 85.01 | 85 | 0 | 39.99 | 40 | 103.7627 | 0 | 29.3572 | 28.9808 | 0 |  | 0.93728 | 30.89 | 29.304 | 0 | 19.13 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0304
2008-08-05 15:05 | 0 | 84.98 | 85 | 0 | 40.04 | 40 | 103.7627 | 0 | 29.3861 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.3162 | 0 | 19.13 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:06 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 103.6617 | 0 | 29.3581 | 28.9818 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 19.16 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0304
2008-08-05 15:07 | 0 | 85.01 | 85 | 0 | 39.94 | 40 | 103.7627 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.2796 | 0 | 20.16 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8464 | 38.0304
2008-08-05 15:08 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3572 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.2796 | 0 | 20.91 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0304
2008-08-05 15:09 | 0 | 84.98 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3861 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 20.66 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:10 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 103.7627 | 0 | 29.3861 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.3162 | 0 | 20.33 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8174 | 38.0305
2008-08-05 15:11 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.82 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.822 | 38.0304
2008-08-05 15:12 | 0 | 84.95 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3861 | 28.9808 | 0 |  | 0.93728 | 30.89 | 29.3162 | 0 | 19.57 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7793 | 38.0304
2008-08-05 15:13 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.29 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8474 | 38.0305
2008-08-05 15:14 | 0 | 84.95 | 85 | 0 | 39.9 | 40 | 103.7627 | 0 | 29.3572 | 28.9808 | 0 |  | 0.93819 | 30.89 | 29.2796 | 0 | 19.1 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0304
2008-08-05 15:15 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.304 | 0 | 19.21 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:16 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.08 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8347 | 38.0305
2008-08-05 15:17 | 0 | 84.98 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.32 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:18 | 0 | 84.98 | 85 | 0 | 40.04 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 19.93 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:19 | 0 | 84.97 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 20.3 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:20 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 103.6617 | 0 | 29.3581 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 20.26 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:21 | 0 | 84.98 | 85 | 0 | 39.94 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.96 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:22 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3581 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.72 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8209 | 38.0305
2008-08-05 15:23 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.5 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:24 | 0 | 85.01 | 85 | 0 | 39.95 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.34 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:25 | 0 | 84.97 | 85 | 0 | 40.02 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.14 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:26 | 0 | 84.98 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.1 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7965 | 38.0305
2008-08-05 15:27 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.01 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:28 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 103.6112 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 18.86 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:29 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 19.5 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:30 | 0 | 84.99 | 85 | 0 | 40.02 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 20.55 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:31 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 103.6112 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93685 | 30.89 | 29.2796 | 0 | 21.24 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:32 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 20.99 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:33 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 103.6617 | 0 | 29.3572 | 28.9818 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 20.41 | 40.8052 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0304
2008-08-05 15:34 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.85 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:35 | 0 | 84.95 | 85 | 0 | 40.03 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2918 | 0 | 19.6 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:36 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.34 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0733 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:37 | 0 | 84.98 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.28 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:38 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.21 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:39 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3581 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.74 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:40 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 20.63 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:41 | 0 | 84.97 | 85 | 0 | 40.03 | 40 | 103.6617 | 0 | 29.3581 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 21.41 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:42 | 0 | 84.96 | 85 | 0 | 39.95 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 21 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:43 | 0 | 84.98 | 85 | 0 | 40.03 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2918 | 0 | 20.41 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:44 | 0 | 85.01 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3562 | 28.9516 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.98 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:45 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93725 | 30.89 | 29.2796 | 0 | 19.6 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25.01 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:46 | 0 | 85.01 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.37 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:47 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 103.6617 | 0 | 29.3572 | 28.9385 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.32 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.837 | 38.0305
2008-08-05 15:48 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.24 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7965 | 38.0305
2008-08-05 15:49 | 0 | 84.97 | 85 | 0 | 39.95 | 40 | 103.6617 | 0 | 29.3427 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2674 | 0 | 19.12 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:50 | 0 | 84.97 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.72 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:51 | 0 | 84.96 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3581 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2674 | 0 | 20.05 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:52 | 0 | 84.95 | 85 | 0 | 40.04 | 40 | 103.6112 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93685 | 30.89 | 29.2796 | 0 | 20.33 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:53 | 0 | 84.96 | 85 | 0 | 39.94 | 40 | 103.6617 | 0 | 29.3572 | 28.9385 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 20.14 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 15:54 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.79 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7965 | 38.0305
2008-08-05 15:55 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.45 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.822 | 38.0305
2008-08-05 15:56 | 0 | 85.01 | 85 | 0 | 40 | 40 | 103.6112 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93682 | 30.89 | 29.2796 | 0 | 19.3 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:57 | 0 | 84.99 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2918 | 0 | 19.24 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:58 | 0 | 84.96 | 85 | 0 | 39.93 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 19.14 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 15:59 | 0 | 84.97 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 19.04 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:00 | 0 | 84.96 | 85 | 0 | 40.03 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93685 | 30.89 | 29.2796 | 0 | 19.38 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:01 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 103.6112 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93685 | 30.89 | 29.2674 | 0 | 20.14 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7955 | 38.0305
2008-08-05 16:02 | 0 | 84.93 | 85 | 0 | 39.9 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 20.61 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:03 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 20.21 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:04 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.86 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 16:05 | 0 | 84.93 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 19.56 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 16:06 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3581 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.42 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:07 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.13 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:08 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.01 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:09 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 19.06 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:10 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 19.13 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:11 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3437 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2674 | 0 | 19.76 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:12 | 0 | 84.99 | 85 | 0 | 40.04 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 20.56 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:13 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2674 | 0 | 20.61 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:14 | 0 | 85.06 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3581 | 28.9526 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 20.16 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:15 | 0 | 84.95 | 85 | 0 | 39.93 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.7 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:16 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3581 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.52 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 16:17 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 103.6617 | 0 | 29.3427 | 28.9394 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 19.17 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8243 | 38.0305
2008-08-05 16:18 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3581 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.09 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:19 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 103.6112 | 0 | 29.3572 | 28.9385 | 0 |  | 0.93682 | 30.89 | 29.2796 | 0 | 19.64 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7965 | 38.0305
2008-08-05 16:20 | 0 | 84.97 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3437 | 28.9394 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 20.62 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:21 | 0 | 84.97 | 85 | 0 | 40.06 | 40 | 103.6617 | 0 | 29.3437 | 28.9394 | 0 |  | 0.93685 | 30.89 | 29.2796 | 0 | 20.95 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7955 | 38.0305
2008-08-05 16:22 | 0 | 84.96 | 85 | 0 | 40 | 40 | 103.6112 | 0 | 29.3437 | 28.9526 | 0 |  | 0.93682 | 30.89 | 29.2674 | 0 | 20.75 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8232 | 38.0305
2008-08-05 16:23 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3427 | 28.9385 | 0 |  | 0.93682 | 30.89 | 29.2796 | 0 | 20.25 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7955 | 38.0305
2008-08-05 16:24 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 103.6112 | 0 | 29.3427 | 28.9385 | 0 |  | 0.93682 | 30.89 | 29.2796 | 0 | 19.72 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8243 | 38.0305
2008-08-05 16:25 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3427 | 28.9385 | 0 |  | 0.93728 | 30.89 | 29.2674 | 0 | 19.49 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8243 | 38.0305
2008-08-05 16:26 | 0 | 84.97 | 85 | 0 | 39.94 | 40 | 103.6112 | 0 | 29.3427 | 28.9385 | 0 |  | 0.93682 | 30.89 | 29.2674 | 0 | 19.37 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8243 | 38.0305
2008-08-05 16:27 | 0 | 84.95 | 85 | 0 | 39.95 | 40 | 103.6617 | 0 | 29.3427 | 28.9385 | 0 |  | 0.93682 | 30.89 | 29.2796 | 0 | 19.25 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7965 | 38.0305
2008-08-05 16:28 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 103.6112 | 0 | 29.3437 | 28.9385 | 0 |  | 0.93682 | 30.89 | 29.2796 | 0 | 19.22 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8243 | 38.0305
2008-08-05 16:29 | 0 | 84.91 | 85 | 0 | 40.02 | 40 | 103.6112 | 0 | 29.3427 | 28.9394 | 0 |  | 0.93685 | 30.89 | 29.2674 | 0 | 19.12 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 16:30 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 103.6112 | 0 | 29.3427 | 28.9385 | 0 |  | 0.93682 | 30.89 | 29.2674 | 0 | 19.42 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8116 | 38.0305
2008-08-05 16:31 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 103.6112 | 0 | 29.3427 | 28.9394 | 0 |  | 0.93685 | 30.89 | 29.2674 | 0 | 19.85 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7955 | 38.0305
2008-08-05 16:32 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 103.6112 | 0 | 29.3437 | 28.9394 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 20.33 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8105 | 38.0305
2008-08-05 16:33 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 103.6112 | 0 | 29.3581 | 28.9385 | 0 |  | 0.93682 | 30.89 | 29.2796 | 0 | 20.41 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8105 | 38.0305
2008-08-05 16:34 | 0 | 84.97 | 85 | 0 | 40.04 | 40 | 103.6112 | 0 | 29.3427 | 28.9385 | 0 |  | 0.93682 | 30.89 | 29.2674 | 0 | 20 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8116 | 38.0305
2008-08-05 16:35 | 0 | 84.93 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3427 | 28.9385 | 0 |  | 0.93682 | 30.89 | 29.2796 | 0 | 19.85 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.836 | 38.0305
2008-08-05 16:36 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3427 | 28.9385 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.48 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8243 | 38.0305
2008-08-05 16:37 | 0 | 84.99 | 85 | 0 | 39.95 | 40 | 103.6112 | 0 | 29.3427 | 28.9526 | 0 |  | 0.93682 | 30.89 | 29.2674 | 0 | 19.22 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8116 | 38.0305
2008-08-05 16:38 | 0 | 84.99 | 85 | 0 | 39.98 | 40 | 103.6112 | 0 | 29.3572 | 28.9385 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.13 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7965 | 38.0305
2008-08-05 16:39 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 103.6112 | 0 | 29.3572 | 28.9385 | 0 |  | 0.93682 | 30.89 | 29.2674 | 0 | 19.02 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7955 | 38.0305
2008-08-05 16:40 | 0 | 84.97 | 85 | 0 | 39.93 | 40 | 103.6112 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93685 | 30.89 | 29.2796 | 0 | 19.08 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8116 | 38.0305
2008-08-05 16:41 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 103.6112 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 19 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8105 | 38.0305
2008-08-05 16:42 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2674 | 0 | 19.02 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:43 | 0 | 85.02 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 20 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8232 | 38.0305
2008-08-05 16:44 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2674 | 0 | 20.97 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:45 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 103.6112 | 0 | 29.3581 | 28.9394 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 20.93 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7955 | 38.0305
2008-08-05 16:46 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93728 | 30.89 | 29.2918 | 0 | 20.51 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:47 | 0 | 84.96 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3427 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 20 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:48 | 0 | 84.97 | 85 | 0 | 39.9 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.68 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8243 | 38.0305
2008-08-05 16:49 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 103.6112 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2674 | 0 | 19.58 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8243 | 38.0305
2008-08-05 16:50 | 0 | 84.99 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.42 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 16:51 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.32 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 16:52 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.28 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 16:53 | 0 | 84.98 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 20.02 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 16:54 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93685 | 30.89 | 29.2796 | 0 | 20.97 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 16:55 | 0 | 84.99 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 20.97 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 16:56 | 0 | 84.97 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 20.61 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 16:57 | 0 | 84.97 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 20.17 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 16:58 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.81 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 16:59 | 0 | 85.01 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.56 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.822 | 38.0305
2008-08-05 17:00 | 0 | 84.93 | 85 | 0 | 39.94 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.5 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 17:01 | 0 | 84.99 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.26 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.822 | 38.0305
2008-08-05 17:02 | 0 | 84.98 | 85 | 0 | 40.03 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.3 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 17:03 | 0 | 84.96 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.36 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 17:04 | 0 | 85.06 | 85 | 0 | 39.97 | 40 | 103.6617 | 0 | 29.3427 | 28.9526 | 0 |  | 0.93731 | 30.89 | 29.2674 | 0 | 19.53 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 17:05 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 103.6617 | 0 | 29.3572 | 28.9385 | 0 |  | 0.93682 | 30.89 | 29.2674 | 0 | 19.61 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.822 | 38.0305
2008-08-05 17:06 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 103.6112 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.7 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 17:07 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3437 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.58 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7965 | 38.0305
2008-08-05 17:08 | 0 | 84.97 | 85 | 0 | 39.96 | 40 | 103.6112 | 0 | 29.3572 | 28.9385 | 0 |  | 0.93682 | 30.89 | 29.2674 | 0 | 19.36 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 17:09 | 0 | 84.93 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3437 | 28.9536 | 0 |  | 0.93685 | 30.89 | 29.2674 | 0 | 19.25 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7965 | 38.0305
2008-08-05 17:10 | 0 | 84.96 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3427 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2674 | 0 | 19.13 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8243 | 38.0305
2008-08-05 17:11 | 0 | 84.91 | 85 | 0 | 39.99 | 40 | 103.6112 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93685 | 30.89 | 29.2796 | 0 | 19.09 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8105 | 38.0305
2008-08-05 17:12 | 0 | 85.01 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 19.05 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8232 | 38.0305
2008-08-05 17:13 | 0 | 84.99 | 85 | 0 | 39.96 | 40 | 103.6617 | 0 | 29.3437 | 28.9394 | 0 |  | 0.93685 | 30.89 | 29.2674 | 0 | 19.08 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 17:14 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 103.6112 | 0 | 29.3437 | 28.9536 | 0 |  | 0.93685 | 30.89 | 29.2674 | 0 | 19.72 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8105 | 38.0305
2008-08-05 17:15 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3437 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2674 | 0 | 20.75 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 17:16 | 0 | 84.98 | 85 | 0 | 40.02 | 40 | 103.6617 | 0 | 29.3581 | 28.9394 | 0 |  | 0.93731 | 30.89 | 29.2796 | 0 | 21.27 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8232 | 38.0305
2008-08-05 17:17 | 0 | 84.95 | 85 | 0 | 40.02 | 40 | 103.6112 | 0 | 29.3437 | 28.9536 | 0 |  | 0.93685 | 30.89 | 29.2674 | 0 | 20.89 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8082 | 38.0305
2008-08-05 17:18 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 103.6112 | 0 | 29.3581 | 28.9394 | 0 |  | 0.93682 | 30.89 | 29.2796 | 0 | 20.29 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.7955 | 38.0305
2008-08-05 17:19 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 103.6112 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93685 | 30.89 | 29.2796 | 0 | 19.84 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.822 | 38.0305
2008-08-05 17:20 | 0 | 84.96 | 85 | 0 | 40 | 40 | 103.6617 | 0 | 29.3427 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2674 | 0 | 19.61 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.822 | 38.0305
2008-08-05 17:21 | 0 | 84.98 | 85 | 0 | 39.95 | 40 | 103.6617 | 0 | 29.3581 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.41 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 17:22 | 0 | 84.97 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3427 | 28.9385 | 0 |  | 0.93728 | 30.89 | 29.2674 | 0 | 19.26 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8243 | 38.0305
2008-08-05 17:23 | 0 | 84.95 | 85 | 0 | 39.98 | 40 | 103.6617 | 0 | 29.3572 | 28.9526 | 0 |  | 0.93728 | 30.89 | 29.2796 | 0 | 19.2 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0659 | 0 | 0 | 0 | 25 | 25 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.822 | 38.0305
2008-08-05 17:24 | 0 | 85.01 | 85 | 0 | 39.94 | 40 | 103.6617 | 0 | 29.3437 | 28.9536 | 0 |  | 0.93731 | 30.89 | 29.2674 | 0 | 19.1 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0586 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 17:25 | 0 | 85.02 | 85 | 0 | 40.02 | 40 | 103.6112 | 0 | 29.3581 | 28.9536 | 0 |  | 0.93685 | 30.89 | 29.2918 | 0 | 19.02 | 40.8053 | 0 | 1.7037 | 15 | 14.0857 |  | 14.0513 | 0 | 0 | 0 | 25 | 24.99 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 15.0309 | 14.1151 | -1 | 15.0309 | 30.8093 | 38.0305
2008-08-05 17:26 | 0 | 85.14 | 85 | 0 | 39.97 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.91501 | 0 | 0.31746 | 0 | 19.58 | 0 | 0 | 0 | 0 | 0 |  | 0.30769 | 0 | 0 | 0 | 4 | 22.98 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.34174 | 0
2008-08-05 17:27 | 0 | 85.33 | 85 | 0 | 39.94 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.92064 | 0 | 0.25641 | 0 | 19.66 | 0 | 0 | 0 | 0 | 0 |  | 0.30769 | 0 | 0 | 0 | 4 | 21.22 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.27701 | 0
2008-08-05 17:28 | 0 | 85.33 | 85 | 0 | 39.98 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.92516 | 0 | 0.23199 | 0 | 20.04 | 0 | 0 | 0 | 0 | 0 |  | 0.30037 | 0 | 0 | 0 | 4 | 19.78 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.23793 | 0
2008-08-05 17:29 | 0 | 85.32 | 85 | 0 | 40.02 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.92916 | 0 | 0.23199 | 0 | 20.4 | 0 | 0 | 0 | 0 | 0 |  | 0.29304 | 0 | 0 | 0 | 4 | 18.52 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.2119 | 0
2008-08-05 17:30 | 0 | 85.28 | 85 | 0 | 39.99 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.93277 | 0 | 0.20757 | 0 | 20.41 | 0 | 0 | 0 | 0 | 0 |  | 0.27839 | 0 | 0 | 0 | 4 | 17.39 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.22342 | 0
2008-08-05 17:31 | 0 | 85.19 | 85 | 0 | 39.92 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.93606 | 0 | 0.20757 | 0 | 20.36 | 0 | 0 | 0 | 0 | 0 |  | 0.27839 | 0 | 0 | 0 | 4 | 16.37 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.21042 | 0
2008-08-05 17:32 | 0 | 85.15 | 85 | 0 | 39.97 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.93911 | 0 | 0.20757 | 0 | 20.18 | 0 | 0 | 0 | 0 | 0 |  | 0.27106 | 0 | 0 | 0 | 4 | 15.43 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.19734 | 0
2008-08-05 17:33 | 0 | 85.1 | 85 | 0 | 39.98 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.94191 | 0 | 0.19536 | 0 | 19.94 | 0 | 0 | 0 | 0 | 0 |  | 0.27106 | 0 | 0 | 0 | 4 | 14.55 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.18441 | 0
2008-08-05 17:34 | 0 | 85.06 | 85 | 0 | 40 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.94461 | 0 | 0.19536 | 0 | 19.87 | 0 | 0 | 0 | 0 | 0 |  | 0.26374 | 0 | 0 | 0 | 4 | 13.74 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.18384 | 0
2008-08-05 17:35 | 0 | 85.03 | 85 | 0 | 39.95 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.94711 | 0 | 0.18315 | 0 | 19.75 | 0 | 0 | 0 | 0 | 0 |  | 0.26374 | 0 | 0 | 0 | 4 | 12.97 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.18344 | 0
2008-08-05 17:36 | 0 | 85.02 | 85 | 0 | 40 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.94954 | 0 | 0.18315 | 0 | 19.67 | 0 | 0 | 0 | 0 | 0 |  | 0.26374 | 0 | 0 | 0 | 4 | 12.26 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.18294 | 0
2008-08-05 17:37 | 0 | 85.06 | 85 | 0 | 39.98 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.95184 | 0 | 0.18315 | 0 | 19.62 | 0 | 0 | 0 | 0 | 0 |  | 0.26374 | 0 | 0 | 0 | 4 | 11.57 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.18248 | 0
2008-08-05 17:38 | 0 | 84.98 | 85 | 0 | 40.03 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.95402 | 0 | 0.18315 | 0 | 19.54 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 10.92 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.18202 | 0
2008-08-05 17:39 | 0 | 84.99 | 85 | 0 | 39.92 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.95607 | 0 | 0.18315 | 0 | 19.5 | 0 | 0 | 0 | 0 | 0 |  | 0.26374 | 0 | 0 | 0 | 4 | 10.3 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.18171 | 0
2008-08-05 17:40 | 0 | 84.93 | 85 | 0 | 39.98 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.95806 | 0 | 0.18315 | 0 | 19.86 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 9.71 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.1813 | 0
2008-08-05 17:41 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.96003 | 0 | 0.18315 | 0 | 20.91 | 0 | 0 | 0 | 0 | 0 |  | 0.26374 | 0 | 0 | 0 | 4 | 9.14 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.18091 | 0
2008-08-05 17:42 | 0 | 84.96 | 85 | 0 | 40.02 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.9618 | 0 | 0.18315 | 0 | 21.4 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 8.61 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.18054 | 0
2008-08-05 17:43 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.96355 | 0 | 0.18315 | 0 | 21.4 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 8.11 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.18028 | 0
2008-08-05 17:44 | 0 | 84.91 | 85 | 0 | 39.98 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.96516 | 0 | 0.18315 | 0 | 20.74 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 7.63 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17995 | 0
2008-08-05 17:45 | 0 | 84.98 | 85 | 0 | 40 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.96674 | 0 | 0.18315 | 0 | 20.39 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 7.18 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17964 | 0
2008-08-05 17:46 | 0 | 84.98 | 85 | 0 | 39.99 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.96823 | 0 | 0.18315 | 0 | 20.1 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 6.75 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17934 | 0
2008-08-05 17:47 | 0 | 84.91 | 85 | 0 | 39.99 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.96961 | 0 | 0.18315 | 0 | 19.97 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 6.34 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17913 | 0
2008-08-05 17:48 | 0 | 84.96 | 85 | 0 | 40.04 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.971 | 0 | 0.19536 | 0 | 19.84 | 0 | 0 | 0 | 0 | 0 |  | 0.26374 | 0 | 0 | 0 | 4 | 5.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17886 | 0
2008-08-05 17:49 | 0 | 84.96 | 85 | 0 | 40 | 40 | 100.4795 | 0 | 0 | 0 | 0 |  | 0.97184 | 0 | 0.18315 | 0 | 19.63 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 5.57 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17869 | 0
2008-08-05 17:50 | 0 | 84.93 | 85 | 0 | 40 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.97359 | 0 | 0.18315 | 0 | 19.71 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 5.21 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17835 | 0
2008-08-05 17:51 | 0 | 84.96 | 85 | 0 | 39.97 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.97481 | 0 | 0.18315 | 0 | 19.92 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 4.86 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17817 | 0
2008-08-05 17:52 | 0 | 84.97 | 85 | 0 | 39.97 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.97597 | 0 | 0.18315 | 0 | 20.12 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 4.53 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17803 | 0
2008-08-05 17:53 | 0 | 84.98 | 85 | 0 | 39.98 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.9771 | 0 | 0.18315 | 0 | 20.36 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 4.2 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17772 | 0
2008-08-05 17:54 | 0 | 84.93 | 85 | 0 | 40.02 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.97784 | 0 | 0.18315 | 0 | 20.63 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 4 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17756 | 0
2008-08-05 17:55 | 0 | 85.01 | 85 | 0 | 39.95 | 40 | 100.53 | 0 | 0 | 0 | 0 |  | 0.97794 | 0 | 0.18315 | 0 | 20.62 | 0 | 0 | 0 | 0 | 0 |  | 0.25641 | 0 | 0 | 0 | 4 | 3.97 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0.17754 | 0
2008-08-05 17:56 | 0 | 85.02 | 85 | 0 | 39.96 | 40 | 100.4795 | 0 | 4.2659 | 4.0958 | 1 |  | 0.97752 | 4.19 | 0.18315 | 0 | 20.29 | 0 | 0 | 0 | 25 | 23.5798 |  | 22.6007 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 0.17754 | 0
2008-08-05 17:57 | 0 | 84.39 | 85 | 0 | 39.98 | 40 | 100.53 | 0 | 4.2659 | 4.0958 | 0 |  | 0.97752 | 4.19 | 4.2491 | 0 | 20.16 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.173 | 0
2008-08-05 17:58 | 0 | 84.42 | 85 | 0 | 39.97 | 40 | 100.53 | 0 | 4.2659 | 4.098 | 0 |  | 0.97805 | 4.19 | 4.2491 | 0 | 19.75 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.94 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.173 | 0
2008-08-05 17:59 | 0 | 84.45 | 85 | 0 | 39.93 | 40 | 100.53 | 0 | 4.266 | 4.098 | 0 |  | 0.97805 | 4.19 | 4.2613 | 0 | 19.6 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.94 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1728 | 0
2008-08-05 18:00 | 0 | 84.55 | 85 | 0 | 39.93 | 40 | 100.53 | 0 | 4.266 | 4.098 | 0 |  | 0.97805 | 4.19 | 4.2613 | 0 | 19.58 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.94 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1728 | 0
2008-08-05 18:01 | 0 | 84.66 | 85 | 0 | 39.92 | 40 | 100.4795 | 0 | 4.2659 | 4.098 | 0 |  | 0.97805 | 4.19 | 4.2491 | 0 | 19.41 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.94 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1853 | 0
2008-08-05 18:02 | 0 | 84.71 | 85 | 0 | 40 | 40 | 100.53 | 0 | 4.2659 | 4.0979 | 0 |  | 0.97801 | 4.19 | 4.2613 | 0 | 19.42 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1853 | 0
2008-08-05 18:03 | 0 | 84.77 | 85 | 0 | 39.98 | 40 | 100.53 | 0 | 4.2659 | 4.0979 | 0 |  | 0.97801 | 4.19 | 4.2491 | 0 | 19.42 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1852 | 0
2008-08-05 18:04 | 0 | 84.88 | 85 | 0 | 39.96 | 40 | 100.53 | 0 | 4.2659 | 4.0958 | 0 |  | 0.97801 | 4.19 | 4.2613 | 0 | 19.34 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1853 | 0
2008-08-05 18:05 | 0 | 84.88 | 85 | 0 | 39.9 | 40 | 100.53 | 0 | 4.2659 | 4.0979 | 0 |  | 0.97801 | 4.19 | 4.2613 | 0 | 19.38 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1853 | 0
2008-08-05 18:06 | 0 | 84.93 | 85 | 0 | 39.96 | 40 | 100.53 | 0 | 4.2659 | 4.0979 | 0 |  | 0.97801 | 4.19 | 4.2491 | 0 | 19.36 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.95 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1853 | 0
2008-08-05 18:07 | 0 | 84.93 | 85 | 0 | 39.95 | 40 | 100.53 | 0 | 4.2657 | 4.0977 | 0 |  | 0.97798 | 4.19 | 4.2613 | 0 | 19.32 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1731 | 0
2008-08-05 18:08 | 0 | 84.95 | 85 | 0 | 39.94 | 40 | 100.53 | 0 | 4.2657 | 4.0977 | 0 |  | 0.97798 | 4.19 | 4.2491 | 0 | 19.38 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5238 | 0 | 0 | 0 | 4 | 3.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1855 | 0
2008-08-05 18:09 | 0 | 84.93 | 85 | 0 | 39.94 | 40 | 100.53 | 0 | 4.2657 | 4.0977 | 0 |  | 0.97798 | 4.19 | 4.2613 | 0 | 19.33 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5385 | 0 | 0 | 0 | 4 | 3.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1855 | 0
2008-08-05 18:10 | 0 | 84.95 | 85 | 0 | 39.99 | 40 | 100.53 | 0 | 4.2657 | 4.0977 | 0 |  | 0.97798 | 4.19 | 4.2613 | 0 | 19.37 | 0 | 0 | 0 | 25 | 23.5798 |  | 23.5311 | 0 | 0 | 0 | 4 | 3.96 |  | 0 | 0 | 0 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 25.0042 | 23.584 | -1 | 25.0042 | 4.1855 | 0

## Detector_Raw
Date | Time | Function | Cart | Flag | Error | ADTIM | Volume | Baseline | BL Dev. | MAXV | Area | Conc
2008-08-04 15:25 | 15:25 | CONT | B | NP | 0 | 150 | 1.2 | 0.18 | 0.025 | 0 | 0 | 0
2008-08-04 15:28 | 15:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.022 | 0.322 | 484595 | 1407
2008-08-04 15:30 | 15:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.18 | 0.017 | 0.369 | 617463 | 1772
2008-08-04 15:33 | 15:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.021 | 0.361 | 617028 | 1789
2008-08-04 15:35 | 15:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.18 | 0.018 | 0.369 | 618723 | 1776
2008-08-04 15:38 | 15:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.023 | 0.361 | 620453 | 1799
2008-08-04 15:40 | 15:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.18 | 0.027 | 0.369 | 619755 | 1779
2008-08-04 15:43 | 15:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.018 | 0.36 | 618600 | 1793
2008-08-04 15:45 | 15:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.18 | 0.019 | 0.37 | 619559 | 1778
2008-08-04 15:48 | 15:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.022 | 0.361 | 618923 | 1794
2008-08-04 15:50 | 15:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.18 | 0.021 | 0.294 | 396849 | 1139
2008-08-04 15:53 | 15:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.018 | 0.182 | 7012 | 20.325
2008-08-04 15:55 | 15:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.18 | 0.021 | 0.18 | 885 | 2.539
2008-08-04 15:58 | 15:58 | CONT | A | NP | 0 | 150 | 1.2 | 0.18 | 0.03 | 0 | 0 | 0
2008-08-04 16:00 | 16:00 | CONT | B | NP | 0 | 150 | 1.2 | 0.18 | 0.022 | 0 | 0 | 0
2008-08-04 16:03 | 16:03 | CONT | A | NP | 0 | 150 | 1.2 | 0.18 | 0.024 | 0 | 0 | 0
2008-08-04 16:05 | 16:05 | CONT | B | NP | 0 | 150 | 1.2 | 0.179 | 0.02 | 0 | 0 | 0
2008-08-04 16:08 | 16:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.025 | 0.196 | 55559 | 161.166
2008-08-04 16:10 | 16:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.24 | 196822 | 564.534
2008-08-04 16:13 | 16:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.24 | 201645 | 584.929
2008-08-04 16:15 | 16:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.016 | 0.242 | 201500 | 577.938
2008-08-04 16:18 | 16:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.031 | 0.24 | 202655 | 587.85
2008-08-04 16:20 | 16:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.031 | 0.242 | 201326 | 577.439
2008-08-04 16:23 | 16:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.238 | 198455 | 575.683
2008-08-04 16:25 | 16:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.18 | 0.021 | 0.238 | 188470 | 540.556
2008-08-04 16:28 | 16:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.021 | 0.236 | 188411 | 546.549
2008-08-04 16:30 | 16:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.18 | 0.027 | 0.238 | 188861 | 541.68
2008-08-04 16:33 | 16:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.027 | 0.236 | 189114 | 548.578
2008-08-04 16:35 | 16:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.238 | 188599 | 540.932
2008-08-04 16:38 | 16:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.238 | 197869 | 573.97
2008-08-04 16:40 | 16:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.18 | 0.017 | 0.243 | 203872 | 584.741
2008-08-04 16:43 | 16:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.028 | 0.24 | 203157 | 589.32
2008-08-04 16:45 | 16:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.242 | 202057 | 579.544
2008-08-04 16:48 | 16:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.023 | 0.239 | 202393 | 587.113
2008-08-04 16:50 | 16:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.243 | 203144 | 582.654
2008-08-04 16:53 | 16:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.238 | 197733 | 573.199
2008-08-04 16:55 | 16:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.238 | 190655 | 547.207
2008-08-04 16:58 | 16:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.236 | 191846 | 556.147
2008-08-04 17:00 | 17:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.027 | 0.239 | 189261 | 543.202
2008-08-04 17:03 | 17:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.024 | 0.235 | 190091 | 551.056
2008-08-04 17:05 | 17:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.239 | 191151 | 548.633
2008-08-04 17:08 | 17:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.032 | 0.237 | 196749 | 570.35
2008-08-04 17:10 | 17:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.242 | 202975 | 582.557
2008-08-04 17:13 | 17:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.042 | 0.239 | 203289 | 589.316
2008-08-04 17:15 | 17:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.032 | 0.243 | 200961 | 576.773
2008-08-04 17:18 | 17:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.019 | 0.239 | 204629 | 593.207
2008-08-04 17:20 | 17:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.027 | 0.243 | 203398 | 583.773
2008-08-04 17:23 | 17:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.023 | 0.238 | 198757 | 576.162
2008-08-04 17:25 | 17:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.18 | 0.023 | 0.238 | 189051 | 542.605
2008-08-04 17:28 | 17:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.017 | 0.235 | 189233 | 548.564
2008-08-04 17:30 | 17:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.237 | 187070 | 536.903
2008-08-04 17:33 | 17:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.235 | 189643 | 549.756
2008-08-04 17:35 | 17:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.238 | 190157 | 545.417
2008-08-04 17:38 | 17:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.025 | 0.237 | 195424 | 566.887
2008-08-04 17:40 | 17:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.016 | 0.242 | 202908 | 581.974
2008-08-04 17:43 | 17:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.029 | 0.239 | 202411 | 587.158
2008-08-04 17:45 | 17:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.031 | 0.242 | 204384 | 586.209
2008-08-04 17:48 | 17:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.24 | 204516 | 593.269
2008-08-04 17:50 | 17:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.244 | 203449 | 583.525
2008-08-04 17:53 | 17:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.238 | 199037 | 577.374
2008-08-04 17:55 | 17:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.028 | 0.239 | 190374 | 546.031
2008-08-04 17:58 | 17:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.236 | 190371 | 552.231
2008-08-04 18:00 | 18:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.017 | 0.238 | 189277 | 542.88
2008-08-04 18:03 | 18:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.18 | 0.029 | 0.236 | 190541 | 552.73
2008-08-04 18:05 | 18:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.238 | 188854 | 541.656
2008-08-04 18:08 | 18:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.237 | 195141 | 566.069
2008-08-04 18:10 | 18:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.242 | 202358 | 580.399
2008-08-04 18:13 | 18:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.24 | 202436 | 587.241
2008-08-04 18:15 | 18:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.242 | 202771 | 581.58
2008-08-04 18:18 | 18:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.239 | 202543 | 587.533
2008-08-04 18:20 | 18:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.018 | 0.242 | 202716 | 581.42
2008-08-04 18:23 | 18:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.031 | 0.238 | 199075 | 577.09
2008-08-04 18:25 | 18:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.018 | 0.238 | 188688 | 541.553
2008-08-04 18:28 | 18:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.029 | 0.235 | 189616 | 549.675
2008-08-04 18:30 | 18:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.027 | 0.239 | 188862 | 542.047
2008-08-04 18:33 | 18:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.235 | 189598 | 549.624
2008-08-04 18:35 | 18:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.237 | 189094 | 542.71
2008-08-04 18:38 | 18:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.03 | 0.237 | 195851 | 567.733
2008-08-04 18:40 | 18:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.031 | 0.243 | 203058 | 582.814
2008-08-04 18:43 | 18:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.24 | 204718 | 593.451
2008-08-04 18:45 | 18:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.243 | 202909 | 582.365
2008-08-04 18:48 | 18:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.017 | 0.24 | 204817 | 593.738
2008-08-04 18:50 | 18:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.244 | 203205 | 583.208
2008-08-04 18:53 | 18:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.037 | 0.238 | 200535 | 581.319
2008-08-04 18:55 | 18:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.239 | 189766 | 544.645
2008-08-04 18:58 | 18:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.235 | 190120 | 551.12
2008-08-04 19:00 | 19:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.239 | 190977 | 548.129
2008-08-04 19:03 | 19:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.034 | 0.236 | 191168 | 554.16
2008-08-04 19:05 | 19:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.239 | 191395 | 548.948
2008-08-04 19:08 | 19:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.027 | 0.237 | 195490 | 567.084
2008-08-04 19:10 | 19:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.243 | 204067 | 585.274
2008-08-04 19:13 | 19:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.24 | 206018 | 597.62
2008-08-04 19:15 | 19:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.243 | 204836 | 587.495
2008-08-04 19:18 | 19:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.018 | 0.24 | 205242 | 595.358
2008-08-04 19:20 | 19:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.243 | 204685 | 587.074
2008-08-04 19:23 | 19:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.29 | 376729 | 1093
2008-08-04 19:25 | 19:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.371 | 621272 | 1782
2008-08-04 19:28 | 19:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.363 | 622079 | 1805
2008-08-04 19:30 | 19:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.371 | 622819 | 1786
2008-08-04 19:33 | 19:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.363 | 624396 | 1811
2008-08-04 19:35 | 19:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.372 | 623671 | 1789
2008-08-04 19:38 | 19:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.36 | 614471 | 1782
2008-08-04 19:40 | 19:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.027 | 0.365 | 599711 | 1720
2008-08-04 19:43 | 19:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.356 | 599265 | 1738
2008-08-04 19:45 | 19:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.364 | 599893 | 1721
2008-08-04 19:48 | 19:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.018 | 0.357 | 599759 | 1740
2008-08-04 19:50 | 19:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.366 | 600816 | 1723
2008-08-04 19:53 | 19:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.359 | 610535 | 1770
2008-08-04 19:55 | 19:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.374 | 625097 | 1794
2008-08-04 19:58 | 19:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.364 | 626123 | 1815
2008-08-04 20:00 | 20:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.371 | 625395 | 1795
2008-08-04 20:03 | 20:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.029 | 0.364 | 624961 | 1812
2008-08-04 20:05 | 20:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.371 | 625122 | 1794
2008-08-04 20:08 | 20:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.36 | 617172 | 1789
2008-08-04 20:10 | 20:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.368 | 600230 | 1723
2008-08-04 20:13 | 20:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.356 | 601469 | 1744
2008-08-04 20:15 | 20:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.369 | 600628 | 1724
2008-08-04 20:18 | 20:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.356 | 599942 | 1739
2008-08-04 20:20 | 20:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.369 | 600698 | 1724
2008-08-04 20:23 | 20:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.029 | 0.358 | 610712 | 1770
2008-08-04 20:25 | 20:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.371 | 625959 | 1797
2008-08-04 20:28 | 20:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.362 | 625609 | 1814
2008-08-04 20:30 | 20:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.371 | 626367 | 1798
2008-08-04 20:33 | 20:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.027 | 0.362 | 626983 | 1818
2008-08-04 20:35 | 20:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.373 | 628020 | 1801
2008-08-04 20:38 | 20:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.36 | 616946 | 1790
2008-08-04 20:40 | 20:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.369 | 600110 | 1721
2008-08-04 20:43 | 20:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.032 | 0.355 | 599659 | 1739
2008-08-04 20:45 | 20:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.363 | 601692 | 1726
2008-08-04 20:48 | 20:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.03 | 0.355 | 600964 | 1743
2008-08-04 20:50 | 20:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.016 | 0.365 | 601256 | 1724
2008-08-04 20:53 | 20:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.027 | 0.358 | 611047 | 1773
2008-08-04 20:55 | 20:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.373 | 626618 | 1797
2008-08-04 20:58 | 20:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.363 | 626156 | 1816
2008-08-04 21:00 | 21:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.373 | 626259 | 1796
2008-08-04 21:03 | 21:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.364 | 627145 | 1819
2008-08-04 21:05 | 21:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.029 | 0.373 | 626830 | 1798
2008-08-04 21:08 | 21:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.36 | 617610 | 1792
2008-08-04 21:10 | 21:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.364 | 600429 | 1722
2008-08-04 21:13 | 21:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.356 | 599103 | 1738
2008-08-04 21:15 | 21:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.366 | 601005 | 1724
2008-08-04 21:18 | 21:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.356 | 600398 | 1742
2008-08-04 21:20 | 21:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.366 | 601203 | 1724
2008-08-04 21:23 | 21:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.359 | 612057 | 1774
2008-08-04 21:25 | 21:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.378 | 627139 | 1800
2008-08-04 21:28 | 21:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.362 | 626628 | 1816
2008-08-04 21:30 | 21:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.376 | 626325 | 1798
2008-08-04 21:33 | 21:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.03 | 0.363 | 626410 | 1816
2008-08-04 21:35 | 21:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.372 | 626642 | 1798
2008-08-04 21:38 | 21:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.361 | 616367 | 1787
2008-08-04 21:40 | 21:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.028 | 0.363 | 599976 | 1722
2008-08-04 21:43 | 21:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.017 | 0.355 | 601059 | 1742
2008-08-04 21:45 | 21:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.365 | 601100 | 1725
2008-08-04 21:48 | 21:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.353 | 601999 | 1745
2008-08-04 21:50 | 21:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.366 | 601205 | 1725
2008-08-04 21:53 | 21:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.03 | 0.359 | 609827 | 1768
2008-08-04 21:55 | 21:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.373 | 626328 | 1798
2008-08-04 21:58 | 21:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.363 | 626288 | 1815
2008-08-04 22:00 | 22:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.374 | 627165 | 1800
2008-08-04 22:03 | 22:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.364 | 627551 | 1819
2008-08-04 22:05 | 22:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.375 | 628113 | 1802
2008-08-04 22:08 | 22:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.362 | 617275 | 1791
2008-08-04 22:10 | 22:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.363 | 599222 | 1719
2008-08-04 22:13 | 22:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.018 | 0.355 | 599507 | 1739
2008-08-04 22:15 | 22:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.365 | 601204 | 1724
2008-08-04 22:18 | 22:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.027 | 0.355 | 600433 | 1742
2008-08-04 22:20 | 22:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.027 | 0.364 | 601246 | 1724
2008-08-04 22:23 | 22:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.358 | 609907 | 1769
2008-08-04 22:25 | 22:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.037 | 0.372 | 626870 | 1798
2008-08-04 22:28 | 22:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.362 | 627233 | 1819
2008-08-04 22:30 | 22:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.375 | 627262 | 1799
2008-08-04 22:33 | 22:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.362 | 626250 | 1817
2008-08-04 22:35 | 22:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.371 | 627617 | 1800
2008-08-04 22:38 | 22:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.288 | 373815 | 1084
2008-08-04 22:40 | 22:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.18 | 2151 | 6.168
2008-08-04 22:43 | 22:43 | CONT | A | NP | 0 | 150 | 1.2 | 0.179 | 0.019 | 0 | 0 | 0
2008-08-04 22:45 | 22:45 | CONT | B | NP | 0 | 150 | 1.2 | 0.179 | 0.018 | 0 | 0 | 0
2008-08-04 22:48 | 22:48 | CONT | A | NP | 0 | 150 | 1.2 | 0.179 | 0.022 | 0 | 0 | 0
2008-08-04 22:50 | 22:50 | CONT | B | NP | 0 | 150 | 1.2 | 0.179 | 0.024 | 0 | 0 | 0
2008-08-04 22:53 | 22:53 | CONT | A | NP | 0 | 150 | 1.2 | 0.179 | 0.026 | 0 | 0 | 0
2008-08-04 22:55 | 22:55 | CONT | B | NP | 0 | 150 | 1.2 | 0.179 | 0.022 | 0 | 0 | 0
2008-08-04 22:58 | 22:58 | CONT | A | NP | 0 | 150 | 1.2 | 0.179 | 0.024 | 0 | 0 | 0
2008-08-04 23:00 | 23:00 | CONT | B | NP | 0 | 150 | 1.2 | 0.179 | 0.029 | 0 | 0 | 0
2008-08-04 23:03 | 23:03 | CONT | A | NP | 0 | 150 | 1.2 | 0.179 | 0.028 | 0 | 0 | 0
2008-08-04 23:05 | 23:05 | CONT | B | NP | 0 | 150 | 1.2 | 0.179 | 0.021 | 0 | 0 | 0
2008-08-04 23:08 | 23:08 | CONT | A | NP | 0 | 150 | 1.2 | 0.179 | 0.022 | 0 | 0 | 0
2008-08-04 23:10 | 23:10 | CONT | B | NP | 0 | 150 | 1.2 | 0.179 | 0.03 | 0 | 0 | 0
2008-08-04 23:13 | 23:13 | CONT | A | NP | 0 | 150 | 1.2 | 0.179 | 0.019 | 0 | 0 | 0
2008-08-04 23:15 | 23:15 | CONT | B | NP | 0 | 150 | 1.2 | 0.179 | 0.02 | 0 | 0 | 0
2008-08-04 23:18 | 23:18 | CONT | A | NP | 0 | 150 | 1.2 | 0.179 | 0.022 | 0 | 0 | 0
2008-08-04 23:20 | 23:20 | CONT | B | NP | 0 | 150 | 1.2 | 0.178 | 0.031 | 0 | 0 | 0
2008-08-04 23:23 | 23:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.025 | 0.308 | 443650 | 1286
2008-08-04 23:25 | 23:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.509 | 1076460 | 3089
2008-08-04 23:28 | 23:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.496 | 1089916 | 3159
2008-08-04 23:30 | 23:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.513 | 1091232 | 3132
2008-08-04 23:33 | 23:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.031 | 0.495 | 1080767 | 3133
2008-08-04 23:35 | 23:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.021 | 0.502 | 1058606 | 3036
2008-08-04 23:38 | 23:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.029 | 0.489 | 1056676 | 3065
2008-08-04 23:40 | 23:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.021 | 0.505 | 1059868 | 3040
2008-08-04 23:43 | 23:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.024 | 0.487 | 1060248 | 3076
2008-08-04 23:45 | 23:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.506 | 1061075 | 3043
2008-08-04 23:48 | 23:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.491 | 1072560 | 3111
2008-08-04 23:50 | 23:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.027 | 0.515 | 1096364 | 3145
2008-08-04 23:53 | 23:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.026 | 0.498 | 1094874 | 3176
2008-08-04 23:55 | 23:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.515 | 1096376 | 3145
2008-08-04 23:58 | 23:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.035 | 0.498 | 1094421 | 3175
2008-08-05 | 0:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.516 | 1096910 | 3146
2008-08-05 00:03 | 0:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.495 | 1083707 | 3144
2008-08-05 00:05 | 0:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.511 | 1062045 | 3046
2008-08-05 00:08 | 0:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.489 | 1058021 | 3069
2008-08-05 00:10 | 0:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.031 | 0.506 | 1062386 | 3047
2008-08-05 00:13 | 0:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.488 | 1059758 | 3074
2008-08-05 00:15 | 0:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.51 | 1062622 | 3048
2008-08-05 00:18 | 0:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.494 | 1076053 | 3119
2008-08-05 00:20 | 0:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.028 | 0.517 | 1094633 | 3142
2008-08-05 00:23 | 0:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.019 | 0.499 | 1094354 | 3172
2008-08-05 00:25 | 0:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.025 | 0.516 | 1096026 | 3146
2008-08-05 00:28 | 0:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.499 | 1094221 | 3172
2008-08-05 00:30 | 0:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.517 | 1095099 | 3143
2008-08-05 00:33 | 0:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.496 | 1083704 | 3141
2008-08-05 00:35 | 0:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.019 | 0.506 | 1062160 | 3048
2008-08-05 00:38 | 0:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.49 | 1062503 | 3080
2008-08-05 00:40 | 0:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.026 | 0.505 | 1063800 | 3053
2008-08-05 00:43 | 0:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.02 | 0.488 | 1063647 | 3083
2008-08-05 00:45 | 0:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.505 | 1063693 | 3053
2008-08-05 00:48 | 0:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.028 | 0.493 | 1076022 | 3119
2008-08-05 00:50 | 0:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.519 | 1097750 | 3151
2008-08-05 00:53 | 0:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.023 | 0.498 | 1095706 | 3176
2008-08-05 00:55 | 0:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.021 | 0.514 | 1093882 | 3139
2008-08-05 00:58 | 0:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.027 | 0.499 | 1094273 | 3172
2008-08-05 01:00 | 1:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.02 | 0.522 | 1095596 | 3144
2008-08-05 01:03 | 1:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.031 | 0.497 | 1084530 | 3144
2008-08-05 01:05 | 1:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.038 | 0.505 | 1062305 | 3047
2008-08-05 01:08 | 1:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.022 | 0.488 | 1057900 | 3069
2008-08-05 01:10 | 1:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.022 | 0.509 | 1062993 | 3049
2008-08-05 01:13 | 1:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.029 | 0.488 | 1061351 | 3079
2008-08-05 01:15 | 1:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.023 | 0.511 | 1063523 | 3050
2008-08-05 01:18 | 1:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.022 | 0.494 | 1073968 | 3115
2008-08-05 01:20 | 1:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.018 | 0.52 | 1097685 | 3148
2008-08-05 01:23 | 1:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.026 | 0.499 | 1094714 | 3175
2008-08-05 01:25 | 1:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.032 | 0.515 | 1098056 | 3149
2008-08-05 01:28 | 1:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.024 | 0.5 | 1095858 | 3179
2008-08-05 01:30 | 1:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.022 | 0.522 | 1097547 | 3148
2008-08-05 01:33 | 1:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.028 | 0.492 | 1083198 | 3142
2008-08-05 01:35 | 1:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.032 | 0.504 | 1063725 | 3051
2008-08-05 01:38 | 1:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.018 | 0.488 | 1059055 | 3072
2008-08-05 01:40 | 1:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.024 | 0.506 | 1063160 | 3049
2008-08-05 01:43 | 1:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.028 | 0.488 | 1060379 | 3076
2008-08-05 01:45 | 1:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.018 | 0.512 | 1064261 | 3052
2008-08-05 01:48 | 1:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.028 | 0.494 | 1074305 | 3114
2008-08-05 01:50 | 1:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.019 | 0.517 | 1094594 | 3141
2008-08-05 01:53 | 1:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.032 | 0.499 | 1095656 | 3176
2008-08-05 01:55 | 1:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.018 | 0.518 | 1094831 | 3142
2008-08-05 01:58 | 1:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.021 | 0.501 | 1096053 | 3177
2008-08-05 02:00 | 2:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.02 | 0.525 | 1096290 | 3146
2008-08-05 02:03 | 2:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.021 | 0.499 | 1093982 | 3171
2008-08-05 02:05 | 2:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.022 | 0.515 | 1096441 | 3147
2008-08-05 02:08 | 2:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.026 | 0.499 | 1095276 | 3175
2008-08-05 02:10 | 2:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.029 | 0.516 | 1097060 | 3149
2008-08-05 02:13 | 2:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.027 | 0.5 | 1095073 | 3174
2008-08-05 02:15 | 2:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.023 | 0.516 | 1094479 | 3141
2008-08-05 02:18 | 2:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.179 | 0.019 | 0.5 | 1095929 | 3177
2008-08-05 02:20 | 2:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.022 | 0.519 | 1095610 | 3144
2008-08-05 02:23 | 2:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.018 | 0.5 | 1094957 | 3174
2008-08-05 02:25 | 2:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.02 | 0.517 | 1096126 | 3146
2008-08-05 02:28 | 2:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.02 | 0.5 | 1096471 | 3178
2008-08-05 02:30 | 2:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.022 | 0.515 | 1096609 | 3145
2008-08-05 02:33 | 2:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.022 | 0.501 | 1093764 | 3173
2008-08-05 02:35 | 2:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.027 | 0.525 | 1098385 | 3150
2008-08-05 02:38 | 2:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.024 | 0.643 | 1582553 | 4590
2008-08-05 02:40 | 2:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.024 | 0.901 | 2312250 | 6632
2008-08-05 02:43 | 2:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.026 | 0.852 | 2309231 | 6698
2008-08-05 02:45 | 2:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.027 | 0.889 | 2314124 | 6637
2008-08-05 02:48 | 2:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.022 | 0.852 | 2311062 | 6704
2008-08-05 02:50 | 2:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.042 | 0.897 | 2317433 | 6647
2008-08-05 02:53 | 2:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.019 | 0.847 | 2291252 | 6646
2008-08-05 02:55 | 2:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.021 | 0.876 | 2266869 | 6502
2008-08-05 02:58 | 2:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.025 | 0.836 | 2257317 | 6548
2008-08-05 03:00 | 3:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.026 | 0.87 | 2262357 | 6489
2008-08-05 03:03 | 3:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.024 | 0.838 | 2260706 | 6558
2008-08-05 03:05 | 3:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.036 | 0.876 | 2265437 | 6498
2008-08-05 03:08 | 3:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.029 | 0.845 | 2281776 | 6619
2008-08-05 03:10 | 3:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.029 | 0.891 | 2319342 | 6652
2008-08-05 03:13 | 3:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.026 | 0.851 | 2313548 | 6711
2008-08-05 03:15 | 3:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.03 | 0.902 | 2320743 | 6656
2008-08-05 03:18 | 3:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.03 | 0.853 | 2316368 | 6715
2008-08-05 03:20 | 3:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.033 | 0.89 | 2318500 | 6654
2008-08-05 03:23 | 3:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.027 | 0.849 | 2295091 | 6653
2008-08-05 03:25 | 3:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.038 | 0.87 | 2265750 | 6503
2008-08-05 03:28 | 3:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.024 | 0.837 | 2260897 | 6554
2008-08-05 03:30 | 3:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.04 | 0.869 | 2259332 | 6484
2008-08-05 03:33 | 3:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.022 | 0.831 | 2257118 | 6543
2008-08-05 03:35 | 3:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.032 | 0.865 | 2262248 | 6493
2008-08-05 03:38 | 3:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.029 | 0.845 | 2282977 | 6618
2008-08-05 03:40 | 3:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.027 | 0.885 | 2319986 | 6659
2008-08-05 03:43 | 3:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.035 | 0.853 | 2316993 | 6717
2008-08-05 03:45 | 3:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.024 | 0.887 | 2318316 | 6653
2008-08-05 03:48 | 3:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.033 | 0.848 | 2314369 | 6709
2008-08-05 03:50 | 3:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.036 | 0.886 | 2318615 | 6654
2008-08-05 03:53 | 3:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.022 | 0.838 | 2296077 | 6656
2008-08-05 03:55 | 3:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.029 | 0.875 | 2266417 | 6505
2008-08-05 03:58 | 3:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.034 | 0.838 | 2262438 | 6558
2008-08-05 04:00 | 4:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.037 | 0.873 | 2268526 | 6506
2008-08-05 04:03 | 4:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.035 | 0.837 | 2261426 | 6560
2008-08-05 04:05 | 4:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.041 | 0.873 | 2271900 | 6516
2008-08-05 04:08 | 4:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.03 | 0.845 | 2284119 | 6626
2008-08-05 04:10 | 4:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.037 | 0.897 | 2322451 | 6661
2008-08-05 04:13 | 4:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.033 | 0.849 | 2315774 | 6717
2008-08-05 04:15 | 4:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.027 | 0.881 | 2320317 | 6655
2008-08-05 04:18 | 4:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.034 | 0.855 | 2313453 | 6711
2008-08-05 04:20 | 4:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.037 | 0.9 | 2320443 | 6655
2008-08-05 04:23 | 4:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.023 | 0.845 | 2295530 | 6659
2008-08-05 04:25 | 4:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.027 | 0.865 | 2265592 | 6498
2008-08-05 04:28 | 4:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.042 | 0.833 | 2259928 | 6555
2008-08-05 04:30 | 4:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.031 | 0.881 | 2263837 | 6493
2008-08-05 04:33 | 4:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.043 | 0.838 | 2261861 | 6561
2008-08-05 04:35 | 4:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.028 | 0.87 | 2265705 | 6498
2008-08-05 04:38 | 4:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.035 | 0.842 | 2279119 | 6611
2008-08-05 04:40 | 4:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.028 | 0.883 | 2322820 | 6662
2008-08-05 04:43 | 4:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.042 | 0.849 | 2317003 | 6717
2008-08-05 04:45 | 4:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.049 | 0.885 | 2318292 | 6654
2008-08-05 04:48 | 4:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.032 | 0.851 | 2317628 | 6718
2008-08-05 04:50 | 4:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.034 | 0.888 | 2319402 | 6657
2008-08-05 04:53 | 4:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.033 | 0.844 | 2300757 | 6670
2008-08-05 04:55 | 4:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.021 | 0.882 | 2266782 | 6505
2008-08-05 04:58 | 4:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.032 | 0.834 | 2265538 | 6567
2008-08-05 05:00 | 5:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.021 | 0.867 | 2265410 | 6502
2008-08-05 05:03 | 5:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.038 | 0.835 | 2261044 | 6554
2008-08-05 05:05 | 5:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.024 | 0.867 | 2264259 | 6499
2008-08-05 05:08 | 5:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.027 | 0.84 | 2282925 | 6618
2008-08-05 05:10 | 5:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.025 | 0.897 | 2319966 | 6658
2008-08-05 05:13 | 5:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.024 | 0.847 | 2315926 | 6713
2008-08-05 05:15 | 5:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.027 | 0.886 | 2321631 | 6663
2008-08-05 05:18 | 5:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.029 | 0.851 | 2316352 | 6715
2008-08-05 05:20 | 5:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.031 | 0.884 | 2323062 | 6667
2008-08-05 05:23 | 5:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.029 | 0.848 | 2299908 | 6667
2008-08-05 05:25 | 5:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.033 | 0.866 | 2267005 | 6507
2008-08-05 05:28 | 5:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.027 | 0.833 | 2266165 | 6569
2008-08-05 05:30 | 5:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.034 | 0.88 | 2270950 | 6513
2008-08-05 05:33 | 5:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.022 | 0.835 | 2263124 | 6565
2008-08-05 05:35 | 5:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.044 | 0.866 | 2272907 | 6519
2008-08-05 05:38 | 5:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.043 | 0.838 | 2281319 | 6617
2008-08-05 05:40 | 5:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.03 | 0.884 | 2320754 | 6656
2008-08-05 05:43 | 5:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.034 | 0.852 | 2314912 | 6715
2008-08-05 05:45 | 5:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.029 | 0.883 | 2320623 | 6656
2008-08-05 05:48 | 5:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.048 | 0.847 | 2314470 | 6714
2008-08-05 05:50 | 5:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.033 | 0.881 | 2321492 | 6658
2008-08-05 05:53 | 5:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.043 | 1.013 | 2921632 | 8475
2008-08-05 05:55 | 5:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.046 | 1.325 | 3815320 | 10943
2008-08-05 05:58 | 5:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.056 | 1.275 | 3816025 | 11069
2008-08-05 06:00 | 6:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.039 | 1.34 | 3828127 | 10980
2008-08-05 06:03 | 6:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.035 | 1.273 | 3821082 | 11084
2008-08-05 06:05 | 6:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.056 | 1.329 | 3822851 | 10964
2008-08-05 06:08 | 6:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.045 | 1.267 | 3783030 | 10973
2008-08-05 06:10 | 6:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.042 | 1.303 | 3741432 | 10731
2008-08-05 06:13 | 6:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.054 | 1.248 | 3729408 | 10811
2008-08-05 06:15 | 6:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.039 | 1.294 | 3731744 | 10710
2008-08-05 06:18 | 6:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.039 | 1.248 | 3714543 | 10768
2008-08-05 06:20 | 6:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.029 | 1.3 | 3716731 | 10667
2008-08-05 06:23 | 6:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.042 | 1.254 | 3743221 | 10851
2008-08-05 06:25 | 6:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.049 | 1.323 | 3796023 | 10895
2008-08-05 06:28 | 6:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.06 | 1.262 | 3784000 | 10969
2008-08-05 06:30 | 6:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.042 | 1.314 | 3780126 | 10849
2008-08-05 06:33 | 6:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.054 | 1.26 | 3783798 | 10968
2008-08-05 06:35 | 6:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.041 | 1.315 | 3791530 | 10882
2008-08-05 06:38 | 6:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.038 | 1.256 | 3739883 | 10841
2008-08-05 06:40 | 6:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.045 | 1.291 | 3683550 | 10572
2008-08-05 06:43 | 6:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.046 | 1.234 | 3675584 | 10655
2008-08-05 06:45 | 6:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.039 | 1.281 | 3691333 | 10594
2008-08-05 06:48 | 6:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.047 | 1.234 | 3686153 | 10686
2008-08-05 06:50 | 6:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.05 | 1.29 | 3691709 | 10595
2008-08-05 06:53 | 6:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.045 | 1.247 | 3715125 | 10769
2008-08-05 06:55 | 6:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.031 | 1.319 | 3761041 | 10794
2008-08-05 06:58 | 6:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.038 | 1.255 | 3762612 | 10907
2008-08-05 07:00 | 7:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.066 | 1.313 | 3770953 | 10815
2008-08-05 07:03 | 7:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.044 | 1.25 | 3755749 | 10894
2008-08-05 07:05 | 7:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.035 | 1.321 | 3760389 | 10785
2008-08-05 07:08 | 7:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.062 | 1.239 | 3720999 | 10794
2008-08-05 07:10 | 7:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.043 | 1.3 | 3682540 | 10562
2008-08-05 07:13 | 7:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.046 | 1.224 | 3677229 | 10667
2008-08-05 07:15 | 7:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.047 | 1.28 | 3679539 | 10553
2008-08-05 07:18 | 7:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.035 | 1.23 | 3673844 | 10657
2008-08-05 07:20 | 7:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.032 | 1.298 | 3686836 | 10574
2008-08-05 07:23 | 7:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.044 | 1.239 | 3713628 | 10772
2008-08-05 07:25 | 7:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.055 | 1.317 | 3772052 | 10818
2008-08-05 07:28 | 7:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.056 | 1.25 | 3748729 | 10874
2008-08-05 07:30 | 7:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.04 | 1.324 | 3763384 | 10794
2008-08-05 07:33 | 7:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.042 | 1.251 | 3765214 | 10922
2008-08-05 07:35 | 7:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.049 | 1.304 | 3775499 | 10828
2008-08-05 07:38 | 7:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.05 | 1.241 | 3726283 | 10809
2008-08-05 07:40 | 7:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.051 | 1.281 | 3688979 | 10581
2008-08-05 07:43 | 7:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.039 | 1.225 | 3678450 | 10663
2008-08-05 07:45 | 7:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.042 | 1.269 | 3679175 | 10559
2008-08-05 07:48 | 7:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.041 | 1.231 | 3677586 | 10660
2008-08-05 07:50 | 7:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.042 | 1.268 | 3695734 | 10606
2008-08-05 07:53 | 7:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.059 | 1.237 | 3718031 | 10778
2008-08-05 07:55 | 7:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.061 | 1.304 | 3766103 | 10809
2008-08-05 07:58 | 7:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.04 | 1.236 | 3766970 | 10919
2008-08-05 08:00 | 8:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.056 | 1.307 | 3775788 | 10837
2008-08-05 08:03 | 8:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.05 | 1.25 | 3763496 | 10909
2008-08-05 08:05 | 8:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.043 | 1.308 | 3758773 | 10788
2008-08-05 08:08 | 8:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.045 | 1.239 | 3723755 | 10794
2008-08-05 08:10 | 8:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.038 | 1.27 | 3682852 | 10570
2008-08-05 08:13 | 8:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.044 | 1.23 | 3678069 | 10662
2008-08-05 08:15 | 8:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.058 | 1.276 | 3680389 | 10563
2008-08-05 08:18 | 8:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.044 | 1.222 | 3667830 | 10632
2008-08-05 08:20 | 8:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.048 | 1.302 | 3676545 | 10552
2008-08-05 08:23 | 8:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.047 | 1.235 | 3705392 | 10741
2008-08-05 08:25 | 8:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.042 | 1.303 | 3760661 | 10793
2008-08-05 08:28 | 8:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.048 | 1.246 | 3759685 | 10898
2008-08-05 08:30 | 8:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.044 | 1.3 | 3773780 | 10824
2008-08-05 08:33 | 8:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.044 | 1.247 | 3754620 | 10891
2008-08-05 08:35 | 8:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.037 | 1.299 | 3759659 | 10783
2008-08-05 08:38 | 8:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.052 | 1.236 | 3723264 | 10800
2008-08-05 08:40 | 8:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.03 | 1.29 | 3687547 | 10576
2008-08-05 08:43 | 8:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.048 | 1.221 | 3673390 | 10655
2008-08-05 08:45 | 8:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.047 | 1.274 | 3680619 | 10555
2008-08-05 08:48 | 8:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.052 | 1.219 | 3657230 | 10609
2008-08-05 08:50 | 8:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.037 | 1.275 | 3675725 | 10543
2008-08-05 08:53 | 8:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.049 | 1.227 | 3699975 | 10733
2008-08-05 08:55 | 8:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.041 | 1.304 | 3762551 | 10792
2008-08-05 08:58 | 8:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.043 | 1.248 | 3751593 | 10882
2008-08-05 09:00 | 9:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.036 | 1.309 | 3763023 | 10792
2008-08-05 09:03 | 9:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.057 | 1.255 | 3753122 | 10886
2008-08-05 09:05 | 9:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.046 | 1.313 | 3765503 | 10799
2008-08-05 09:08 | 9:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.042 | 0.829 | 2246255 | 6516
2008-08-05 09:10 | 9:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.03 | 0.184 | 17850 | 51.193
2008-08-05 09:13 | 9:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.044 | 0.181 | 10756 | 31.177
2008-08-05 09:15 | 9:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.019 | 0.181 | 6731 | 19.316
2008-08-05 09:18 | 9:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.029 | 0.18 | 4413 | 12.791
2008-08-05 09:20 | 9:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.028 | 0.18 | 4472 | 12.835
2008-08-05 09:23 | 9:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.025 | 0.179 | 3534 | 10.244
2008-08-05 09:25 | 9:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.023 | 0.179 | 2771 | 7.951
2008-08-05 09:28 | 9:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.028 | 0.179 | 2024 | 5.867
2008-08-05 09:30 | 9:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.029 | 0.179 | 1979 | 5.678
2008-08-05 09:33 | 9:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.02 | 0.179 | 2118 | 6.138
2008-08-05 09:35 | 9:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.028 | 0.179 | 1434 | 4.114
2008-08-05 09:38 | 9:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.02 | 0.179 | 2825 | 8.188
2008-08-05 09:40 | 9:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.031 | 0.179 | 1868 | 5.361
2008-08-05 09:43 | 9:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.026 | 0.179 | 1186 | 3.437
2008-08-05 09:45 | 9:45 | CONT | B | NP | 0 | 150 | 1.2 | 0.178 | 0.018 | 0 | 0 | 0
2008-08-05 09:48 | 9:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.021 | 0.179 | 1364 | 3.954
2008-08-05 09:50 | 9:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.028 | 0.179 | 1411 | 4.048
2008-08-05 09:53 | 9:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.023 | 0.919 | 2585162 | 7494
2008-08-05 09:55 | 9:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.037 | 2.298 | 7115629 | 20409
2008-08-05 09:58 | 9:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.052 | 2.17 | 7116442 | 20643
2008-08-05 10:00 | 10:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.055 | 2.31 | 7137315 | 20471
2008-08-05 10:03 | 10:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.064 | 2.174 | 7130148 | 20683
2008-08-05 10:05 | 10:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.067 | 2.269 | 7161087 | 20538
2008-08-05 10:08 | 10:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.075 | 2.193 | 7210184 | 20914
2008-08-05 10:10 | 10:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.063 | 2.336 | 7316407 | 20984
2008-08-05 10:13 | 10:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.065 | 2.216 | 7280084 | 21117
2008-08-05 10:15 | 10:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.083 | 2.326 | 7304580 | 20950
2008-08-05 10:18 | 10:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.065 | 2.232 | 7294962 | 21157
2008-08-05 10:20 | 10:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.063 | 2.326 | 7310795 | 20968
2008-08-05 10:23 | 10:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.082 | 2.209 | 7241661 | 21006
2008-08-05 10:25 | 10:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.065 | 2.288 | 7161510 | 20539
2008-08-05 10:28 | 10:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.079 | 2.184 | 7154281 | 20752
2008-08-05 10:30 | 10:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.06 | 2.298 | 7164851 | 20550
2008-08-05 10:33 | 10:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.071 | 2.184 | 7150253 | 20740
2008-08-05 10:35 | 10:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.084 | 2.292 | 7150426 | 20508
2008-08-05 10:38 | 10:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.071 | 2.192 | 7201276 | 20889
2008-08-05 10:40 | 10:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.075 | 2.345 | 7327054 | 21014
2008-08-05 10:43 | 10:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.081 | 2.22 | 7309726 | 21189
2008-08-05 10:45 | 10:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.067 | 2.332 | 7314879 | 20994
2008-08-05 10:48 | 10:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.068 | 2.222 | 7300319 | 21161
2008-08-05 10:50 | 10:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.068 | 2.352 | 7310719 | 20981
2008-08-05 10:53 | 10:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.063 | 2.206 | 7265561 | 21062
2008-08-05 10:55 | 10:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.07 | 2.283 | 7172079 | 20584
2008-08-05 10:58 | 10:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.086 | 2.183 | 7174600 | 20796
2008-08-05 11:00 | 11:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.08 | 2.308 | 7176587 | 20596
2008-08-05 11:03 | 11:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.072 | 2.194 | 7164474 | 20768
2008-08-05 11:05 | 11:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.042 | 2.308 | 7182651 | 20614
2008-08-05 11:08 | 11:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.064 | 2.198 | 7244248 | 20999
2008-08-05 11:10 | 11:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.063 | 2.328 | 7343591 | 21075
2008-08-05 11:13 | 11:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.069 | 2.221 | 7319223 | 21216
2008-08-05 11:15 | 11:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.084 | 2.344 | 7334262 | 21049
2008-08-05 11:18 | 11:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.08 | 2.233 | 7333501 | 21258
2008-08-05 11:20 | 11:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.073 | 2.331 | 7339372 | 21064
2008-08-05 11:23 | 11:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.066 | 2.21 | 7271267 | 21078
2008-08-05 11:25 | 11:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.068 | 2.303 | 7187914 | 20615
2008-08-05 11:28 | 11:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.085 | 2.172 | 7178950 | 20824
2008-08-05 11:30 | 11:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.079 | 2.304 | 7185794 | 20609
2008-08-05 11:33 | 11:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.078 | 2.215 | 7158614 | 20765
2008-08-05 11:35 | 11:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.072 | 2.325 | 7160129 | 20535
2008-08-05 11:38 | 11:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.07 | 2.213 | 7222598 | 20950
2008-08-05 11:40 | 11:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.066 | 2.349 | 7338187 | 21047
2008-08-05 11:43 | 11:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.062 | 2.23 | 7308885 | 21201
2008-08-05 11:45 | 11:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.077 | 2.379 | 7323902 | 21005
2008-08-05 11:48 | 11:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.082 | 2.234 | 7319941 | 21233
2008-08-05 11:50 | 11:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.09 | 2.373 | 7328313 | 21018
2008-08-05 11:53 | 11:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.086 | 2.218 | 7241046 | 21003
2008-08-05 11:55 | 11:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.064 | 2.304 | 7167940 | 20557
2008-08-05 11:58 | 11:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.08 | 2.195 | 7151126 | 20743
2008-08-05 12:00 | 12:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.082 | 2.32 | 7157615 | 20529
2008-08-05 12:03 | 12:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.089 | 2.189 | 7150909 | 20742
2008-08-05 12:05 | 12:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.078 | 2.307 | 7163662 | 20546
2008-08-05 12:08 | 12:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.075 | 2.226 | 7199507 | 20883
2008-08-05 12:10 | 12:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.074 | 2.349 | 7318198 | 20989
2008-08-05 12:13 | 12:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.076 | 2.238 | 7300239 | 21162
2008-08-05 12:15 | 12:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.079 | 2.353 | 7302262 | 20957
2008-08-05 12:18 | 12:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.079 | 2.238 | 7305648 | 21177
2008-08-05 12:20 | 12:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.067 | 2.384 | 7310525 | 20981
2008-08-05 12:23 | 12:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.079 | 2.233 | 7291249 | 21135
2008-08-05 12:25 | 12:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.088 | 2.373 | 7293603 | 20932
2008-08-05 12:28 | 12:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.081 | 2.24 | 7298306 | 21156
2008-08-05 12:30 | 12:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.064 | 2.354 | 7294223 | 20934
2008-08-05 12:33 | 12:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.094 | 2.225 | 7279676 | 21102
2008-08-05 12:35 | 12:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.065 | 2.34 | 7290589 | 20924
2008-08-05 12:38 | 12:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.097 | 2.238 | 7293377 | 21142
2008-08-05 12:40 | 12:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.064 | 2.353 | 7306382 | 20969
2008-08-05 12:43 | 12:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.082 | 2.222 | 7287388 | 21125
2008-08-05 12:45 | 12:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.08 | 2.37 | 7290580 | 20923
2008-08-05 12:48 | 12:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.068 | 2.228 | 7302465 | 21168
2008-08-05 12:50 | 12:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.087 | 2.375 | 7305823 | 20967
2008-08-05 12:53 | 12:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.064 | 2.232 | 7288673 | 21130
2008-08-05 12:55 | 12:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.086 | 2.341 | 7286274 | 20897
2008-08-05 12:58 | 12:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.097 | 2.227 | 7286505 | 21135
2008-08-05 13:00 | 13:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.087 | 2.341 | 7306710 | 20957
2008-08-05 13:03 | 13:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.093 | 2.229 | 7293164 | 21154
2008-08-05 13:05 | 13:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.077 | 2.334 | 7284966 | 20894
2008-08-05 13:08 | 13:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.074 | 1.463 | 4485950 | 13012
2008-08-05 13:10 | 13:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.056 | 0.19 | 38221 | 109.619
2008-08-05 13:13 | 13:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.044 | 0.185 | 22481 | 65.209
2008-08-05 13:15 | 13:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.042 | 0.183 | 15603 | 44.749
2008-08-05 13:18 | 13:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.026 | 0.182 | 12840 | 37.245
2008-08-05 13:20 | 13:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.033 | 0.181 | 10402 | 29.832
2008-08-05 13:23 | 13:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.046 | 0.181 | 8795 | 25.512
2008-08-05 13:25 | 13:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.035 | 0.18 | 7586 | 21.756
2008-08-05 13:28 | 13:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.028 | 0.18 | 7325 | 21.247
2008-08-05 13:30 | 13:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.029 | 0.18 | 5775 | 16.561
2008-08-05 13:33 | 13:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.034 | 0.18 | 6111 | 17.725
2008-08-05 13:35 | 13:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.025 | 0.179 | 4216 | 12.091
2008-08-05 13:38 | 13:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.017 | 0.179 | 3836 | 11.125
2008-08-05 13:40 | 13:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.029 | 0.179 | 3751 | 10.757
2008-08-05 13:43 | 13:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.036 | 0.179 | 4286 | 12.423
2008-08-05 13:45 | 13:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.025 | 0.179 | 3128 | 8.976
2008-08-05 13:48 | 13:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.027 | 0.179 | 4832 | 14.005
2008-08-05 13:50 | 13:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.028 | 0.179 | 3177 | 9.116
2008-08-05 13:53 | 13:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.03 | 1.797 | 5734504 | 16623
2008-08-05 13:55 | 13:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.041 | 4.111 | 13838000 | 39713
2008-08-05 13:58 | 13:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.11 | 3.927 | 13887000 | 40253
2008-08-05 14:00 | 14:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.117 | 4.152 | 13896000 | 39880
2008-08-05 14:03 | 14:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.107 | 3.932 | 13834000 | 40098
2008-08-05 14:05 | 14:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.128 | 4.108 | 13695000 | 39304
2008-08-05 14:08 | 14:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.138 | 3.92 | 13712000 | 39746
2008-08-05 14:10 | 14:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.125 | 4.094 | 13702000 | 39325
2008-08-05 14:13 | 14:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.139 | 3.9 | 13735000 | 39815
2008-08-05 14:15 | 14:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.118 | 4.086 | 13734000 | 39414
2008-08-05 14:18 | 14:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.136 | 3.948 | 13800000 | 40003
2008-08-05 14:20 | 14:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.136 | 4.127 | 13952000 | 40041
2008-08-05 14:23 | 14:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.137 | 3.971 | 13957000 | 40457
2008-08-05 14:25 | 14:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.143 | 4.129 | 13974000 | 40078
2008-08-05 14:28 | 14:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.15 | 3.975 | 13947000 | 40456
2008-08-05 14:30 | 14:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.157 | 4.122 | 13958000 | 40033
2008-08-05 14:33 | 14:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.141 | 3.922 | 13857000 | 40195
2008-08-05 14:35 | 14:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.136 | 4.092 | 13725000 | 39365
2008-08-05 14:38 | 14:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.143 | 3.902 | 13736000 | 39843
2008-08-05 14:40 | 14:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.128 | 4.081 | 13758000 | 39460
2008-08-05 14:43 | 14:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.139 | 3.912 | 13720000 | 39796
2008-08-05 14:45 | 14:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.134 | 4.079 | 13740000 | 39408
2008-08-05 14:48 | 14:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.158 | 3.934 | 13814000 | 40070
2008-08-05 14:50 | 14:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.151 | 4.142 | 13998000 | 40148
2008-08-05 14:53 | 14:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.139 | 3.964 | 13979000 | 40549
2008-08-05 14:55 | 14:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.139 | 4.191 | 13963000 | 40047
2008-08-05 14:58 | 14:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.139 | 3.969 | 13947000 | 40454
2008-08-05 15:00 | 15:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.139 | 4.144 | 13996000 | 40142
2008-08-05 15:03 | 15:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.152 | 3.971 | 13912000 | 40353
2008-08-05 15:05 | 15:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.134 | 4.121 | 13771000 | 39497
2008-08-05 15:08 | 15:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.137 | 3.911 | 13734000 | 39836
2008-08-05 15:10 | 15:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.144 | 4.078 | 13755000 | 39447
2008-08-05 15:13 | 15:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.139 | 3.919 | 13765000 | 39901
2008-08-05 15:15 | 15:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.129 | 4.069 | 13747000 | 39453
2008-08-05 15:18 | 15:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.155 | 3.872 | 13818000 | 40053
2008-08-05 15:20 | 15:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.146 | 4.138 | 13956000 | 40053
2008-08-05 15:23 | 15:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.146 | 3.99 | 13976000 | 40512
2008-08-05 15:25 | 15:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.139 | 4.159 | 13978000 | 40115
2008-08-05 15:28 | 15:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.144 | 3.981 | 13985000 | 40538
2008-08-05 15:30 | 15:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.139 | 4.148 | 13957000 | 40055
2008-08-05 15:33 | 15:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.14 | 3.963 | 13893000 | 40269
2008-08-05 15:35 | 15:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.148 | 4.062 | 13763000 | 39495
2008-08-05 15:38 | 15:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.135 | 3.909 | 13750000 | 39859
2008-08-05 15:40 | 15:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.134 | 4.082 | 13725000 | 39390
2008-08-05 15:43 | 15:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.143 | 3.902 | 13719000 | 39768
2008-08-05 15:45 | 15:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.164 | 4.105 | 13724000 | 39386
2008-08-05 15:48 | 15:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.154 | 3.94 | 13796000 | 39990
2008-08-05 15:50 | 15:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.158 | 4.133 | 13933000 | 39987
2008-08-05 15:53 | 15:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.15 | 3.973 | 13942000 | 40414
2008-08-05 15:55 | 15:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.147 | 4.143 | 13974000 | 40078
2008-08-05 15:58 | 15:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.156 | 3.968 | 13952000 | 40471
2008-08-05 16:00 | 16:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.151 | 4.17 | 13946000 | 40000
2008-08-05 16:03 | 16:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.143 | 3.938 | 13851000 | 40178
2008-08-05 16:05 | 16:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.134 | 4.08 | 13724000 | 39362
2008-08-05 16:08 | 16:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.145 | 3.91 | 13708000 | 39761
2008-08-05 16:10 | 16:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.148 | 4.074 | 13743000 | 39415
2008-08-05 16:13 | 16:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.149 | 3.908 | 13705000 | 39755
2008-08-05 16:15 | 16:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.127 | 4.087 | 13748000 | 39428
2008-08-05 16:18 | 16:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.134 | 3.936 | 13818000 | 40081
2008-08-05 16:20 | 16:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.144 | 4.12 | 13958000 | 40032
2008-08-05 16:23 | 16:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.159 | 3.976 | 13949000 | 40462
2008-08-05 16:25 | 16:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.143 | 4.128 | 13993000 | 40133
2008-08-05 16:28 | 16:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.148 | 3.953 | 13948000 | 40457
2008-08-05 16:30 | 16:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.164 | 4.141 | 13946000 | 39996
2008-08-05 16:33 | 16:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.152 | 3.96 | 13925000 | 40391
2008-08-05 16:35 | 16:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.147 | 4.155 | 13965000 | 40054
2008-08-05 16:38 | 16:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.16 | 3.974 | 13945000 | 40452
2008-08-05 16:40 | 16:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.178 | 0.143 | 4.173 | 13953000 | 40018
2008-08-05 16:43 | 16:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.178 | 0.162 | 3.976 | 13935000 | 40392
2008-08-05 16:45 | 16:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.154 | 4.162 | 13951000 | 40038
2008-08-05 16:48 | 16:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.149 | 3.978 | 13966000 | 40484
2008-08-05 16:50 | 16:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.15 | 4.145 | 13958000 | 40060
2008-08-05 16:53 | 16:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.15 | 3.903 | 13715000 | 39754
2008-08-05 16:55 | 16:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.154 | 4.07 | 13722000 | 39383
2008-08-05 16:58 | 16:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.149 | 3.918 | 13716000 | 39757
2008-08-05 17:00 | 17:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.152 | 4.11 | 13704000 | 39331
2008-08-05 17:03 | 17:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.152 | 3.845 | 13735000 | 39813
2008-08-05 17:05 | 17:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.13 | 4.13 | 13728000 | 39399
2008-08-05 17:08 | 17:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.166 | 3.972 | 13932000 | 40387
2008-08-05 17:10 | 17:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.159 | 4.121 | 13939000 | 40003
2008-08-05 17:13 | 17:13 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.145 | 3.967 | 13966000 | 40484
2008-08-05 17:15 | 17:15 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.139 | 4.131 | 14013000 | 40216
2008-08-05 17:18 | 17:18 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.146 | 3.989 | 13970000 | 40498
2008-08-05 17:20 | 17:20 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.163 | 3.003 | 10748000 | 30847
2008-08-05 17:23 | 17:23 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.112 | 0.214 | 125913 | 364.989
2008-08-05 17:25 | 17:25 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.051 | 0.191 | 42712 | 122.5
2008-08-05 17:28 | 17:28 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.053 | 0.186 | 28384 | 82.33
2008-08-05 17:30 | 17:30 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.042 | 0.184 | 22166 | 63.573
2008-08-05 17:33 | 17:33 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.035 | 0.183 | 17779 | 51.571
2008-08-05 17:35 | 17:35 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.033 | 0.182 | 14346 | 41.145
2008-08-05 17:38 | 17:38 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.024 | 0.182 | 13122 | 38.063
2008-08-05 17:40 | 17:40 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.032 | 0.181 | 11460 | 32.868
2008-08-05 17:43 | 17:43 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.029 | 0.181 | 11219 | 32.542
2008-08-05 17:45 | 17:45 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.04 | 0.181 | 9981 | 28.624
2008-08-05 17:48 | 17:48 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.038 | 0.18 | 10002 | 29.012
2008-08-05 17:50 | 17:50 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.035 | 0.181 | 7777 | 22.305
2008-08-05 17:53 | 17:53 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.031 | 0.18 | 8783 | 25.475
2008-08-05 17:55 | 17:55 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.042 | 0.18 | 7748 | 22.22
2008-08-05 17:58 | 17:58 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.041 | 0.18 | 7321 | 21.235
2008-08-05 18:00 | 18:00 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.025 | 0.18 | 5874 | 16.846
2008-08-05 18:03 | 18:03 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.047 | 0.18 | 6763 | 19.614
2008-08-05 18:05 | 18:05 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.033 | 0.18 | 5840 | 16.749
2008-08-05 18:08 | 18:08 | CONT | A | OK | 0 | 150 | 1.2 | 0.177 | 0.045 | 0.18 | 5563 | 16.125
2008-08-05 18:10 | 18:10 | CONT | B | OK | 0 | 150 | 1.2 | 0.177 | 0.046 | 0.179 | 5516 | 15.83

## AB-Split
"A-Channel" Detector Responses |  |  |  |  | "B-Channel" Detector Responses
Date & Time | Conc | Reference | Candidate | Zero | Date & Time | Conc | Reference | Candidate | Zero
2008-08-04 15:25 | 0 |  |  | 0 | 2008-08-04 15:25 | 0
2008-08-04 15:28 | 1407 |  |  |  | 2008-08-04 15:30 | 1772
2008-08-04 15:33 | 1789 |  |  |  | 2008-08-04 15:35 | 1776
2008-08-04 15:38 | 1799 |  |  |  | 2008-08-04 15:40 | 1779
2008-08-04 15:43 | 1793 |  |  |  | 2008-08-04 15:45 | 1778
2008-08-04 15:48 | 1794 |  |  |  | 2008-08-04 15:50 | 1139
2008-08-04 15:53 | 20.325 |  |  |  | 2008-08-04 15:55 | 2.539
2008-08-04 15:58 | 0 |  |  | 0 | 2008-08-04 16:00 | 0 |  |  | 0
2008-08-04 16:03 | 0 |  |  | 0 | 2008-08-04 16:05 | 0 |  |  | 0
2008-08-04 16:08 | 161.166 |  |  |  | 2008-08-04 16:10 | 564.534
2008-08-04 16:13 | 584.929 | 584.929 |  |  | 2008-08-04 16:15 | 577.938 | 577.938
2008-08-04 16:18 | 587.85 | 587.85 |  |  | 2008-08-04 16:20 | 577.439 | 577.439
2008-08-04 16:23 | 575.683 |  |  |  | 2008-08-04 16:25 | 540.556 |  | 540.556
2008-08-04 16:28 | 546.549 |  | 546.549 |  | 2008-08-04 16:30 | 541.68 |  | 541.68
2008-08-04 16:33 | 548.578 |  | 548.578 |  | 2008-08-04 16:35 | 540.932 |  | 540.932
2008-08-04 16:38 | 573.97 |  |  |  | 2008-08-04 16:40 | 584.741 | 584.741
2008-08-04 16:43 | 589.32 | 589.32 |  |  | 2008-08-04 16:45 | 579.544 | 579.544
2008-08-04 16:48 | 587.113 | 587.113 |  |  | 2008-08-04 16:50 | 582.654 | 582.654
2008-08-04 16:53 | 573.199 |  |  |  | 2008-08-04 16:55 | 547.207 |  | 547.207
2008-08-04 16:58 | 556.147 |  | 556.147 |  | 2008-08-04 17:00 | 543.202 |  | 543.202
2008-08-04 17:03 | 551.056 |  | 551.056 |  | 2008-08-04 17:05 | 548.633 |  | 548.633
2008-08-04 17:08 | 570.35 |  |  |  | 2008-08-04 17:10 | 582.557 | 582.557
2008-08-04 17:13 | 589.316 | 589.316 |  |  | 2008-08-04 17:15 | 576.773 | 576.773
2008-08-04 17:18 | 593.207 | 593.207 |  |  | 2008-08-04 17:20 | 583.773 | 583.773
2008-08-04 17:23 | 576.162 |  |  |  | 2008-08-04 17:25 | 542.605 |  | 542.605
2008-08-04 17:28 | 548.564 |  | 548.564 |  | 2008-08-04 17:30 | 536.903 |  | 536.903
2008-08-04 17:33 | 549.756 |  | 549.756 |  | 2008-08-04 17:35 | 545.417 |  | 545.417
2008-08-04 17:38 | 566.887 |  |  |  | 2008-08-04 17:40 | 581.974 | 581.974
2008-08-04 17:43 | 587.158 | 587.158 |  |  | 2008-08-04 17:45 | 586.209 | 586.209
2008-08-04 17:48 | 593.269 | 593.269 |  |  | 2008-08-04 17:50 | 583.525 | 583.525
2008-08-04 17:53 | 577.374 |  |  |  | 2008-08-04 17:55 | 546.031 |  | 546.031
2008-08-04 17:58 | 552.231 |  | 552.231 |  | 2008-08-04 18:00 | 542.88 |  | 542.88
2008-08-04 18:03 | 552.73 |  | 552.73 |  | 2008-08-04 18:05 | 541.656 |  | 541.656
2008-08-04 18:08 | 566.069 |  |  |  | 2008-08-04 18:10 | 580.399 | 580.399
2008-08-04 18:13 | 587.241 | 587.241 |  |  | 2008-08-04 18:15 | 581.58 | 581.58
2008-08-04 18:18 | 587.533 | 587.533 |  |  | 2008-08-04 18:20 | 581.42 | 581.42
2008-08-04 18:23 | 577.09 |  |  |  | 2008-08-04 18:25 | 541.553 |  | 541.553
2008-08-04 18:28 | 549.675 |  | 549.675 |  | 2008-08-04 18:30 | 542.047 |  | 542.047
2008-08-04 18:33 | 549.624 |  | 549.624 |  | 2008-08-04 18:35 | 542.71 |  | 542.71
2008-08-04 18:38 | 567.733 |  |  |  | 2008-08-04 18:40 | 582.814 | 582.814
2008-08-04 18:43 | 593.451 | 593.451 |  |  | 2008-08-04 18:45 | 582.365 | 582.365
2008-08-04 18:48 | 593.738 | 593.738 |  |  | 2008-08-04 18:50 | 583.208 | 583.208
2008-08-04 18:53 | 581.319 |  |  |  | 2008-08-04 18:55 | 544.645 |  | 544.645
2008-08-04 18:58 | 551.12 |  | 551.12 |  | 2008-08-04 19:00 | 548.129 |  | 548.129
2008-08-04 19:03 | 554.16 |  | 554.16 |  | 2008-08-04 19:05 | 548.948 |  | 548.948
2008-08-04 19:08 | 567.084 |  |  |  | 2008-08-04 19:10 | 585.274 | 585.274
2008-08-04 19:13 | 597.62 | 597.62 |  |  | 2008-08-04 19:15 | 587.495 | 587.495
2008-08-04 19:18 | 595.358 | 595.358 |  |  | 2008-08-04 19:20 | 587.074 | 587.074
2008-08-04 19:23 | 1093 |  |  |  | 2008-08-04 19:25 | 1782 | 1782
2008-08-04 19:28 | 1805 | 1805 |  |  | 2008-08-04 19:30 | 1786 | 1786
2008-08-04 19:33 | 1811 | 1811 |  |  | 2008-08-04 19:35 | 1789 | 1789
2008-08-04 19:38 | 1782 |  |  |  | 2008-08-04 19:40 | 1720 |  | 1720
2008-08-04 19:43 | 1738 |  | 1738 |  | 2008-08-04 19:45 | 1721 |  | 1721
2008-08-04 19:48 | 1740 |  | 1740 |  | 2008-08-04 19:50 | 1723 |  | 1723
2008-08-04 19:53 | 1770 |  |  |  | 2008-08-04 19:55 | 1794 | 1794
2008-08-04 19:58 | 1815 | 1815 |  |  | 2008-08-04 20:00 | 1795 | 1795
2008-08-04 20:03 | 1812 | 1812 |  |  | 2008-08-04 20:05 | 1794 | 1794
2008-08-04 20:08 | 1789 |  |  |  | 2008-08-04 20:10 | 1723 |  | 1723
2008-08-04 20:13 | 1744 |  | 1744 |  | 2008-08-04 20:15 | 1724 |  | 1724
2008-08-04 20:18 | 1739 |  | 1739 |  | 2008-08-04 20:20 | 1724 |  | 1724
2008-08-04 20:23 | 1770 |  |  |  | 2008-08-04 20:25 | 1797 | 1797
2008-08-04 20:28 | 1814 | 1814 |  |  | 2008-08-04 20:30 | 1798 | 1798
2008-08-04 20:33 | 1818 | 1818 |  |  | 2008-08-04 20:35 | 1801 | 1801
2008-08-04 20:38 | 1790 |  |  |  | 2008-08-04 20:40 | 1721 |  | 1721
2008-08-04 20:43 | 1739 |  | 1739 |  | 2008-08-04 20:45 | 1726 |  | 1726
2008-08-04 20:48 | 1743 |  | 1743 |  | 2008-08-04 20:50 | 1724 |  | 1724
2008-08-04 20:53 | 1773 |  |  |  | 2008-08-04 20:55 | 1797 | 1797
2008-08-04 20:58 | 1816 | 1816 |  |  | 2008-08-04 21:00 | 1796 | 1796
2008-08-04 21:03 | 1819 | 1819 |  |  | 2008-08-04 21:05 | 1798 | 1798
2008-08-04 21:08 | 1792 |  |  |  | 2008-08-04 21:10 | 1722 |  | 1722
2008-08-04 21:13 | 1738 |  | 1738 |  | 2008-08-04 21:15 | 1724 |  | 1724
2008-08-04 21:18 | 1742 |  | 1742 |  | 2008-08-04 21:20 | 1724 |  | 1724
2008-08-04 21:23 | 1774 |  |  |  | 2008-08-04 21:25 | 1800 | 1800
2008-08-04 21:28 | 1816 | 1816 |  |  | 2008-08-04 21:30 | 1798 | 1798
2008-08-04 21:33 | 1816 | 1816 |  |  | 2008-08-04 21:35 | 1798 | 1798
2008-08-04 21:38 | 1787 |  |  |  | 2008-08-04 21:40 | 1722 |  | 1722
2008-08-04 21:43 | 1742 |  | 1742 |  | 2008-08-04 21:45 | 1725 |  | 1725
2008-08-04 21:48 | 1745 |  | 1745 |  | 2008-08-04 21:50 | 1725 |  | 1725
2008-08-04 21:53 | 1768 |  |  |  | 2008-08-04 21:55 | 1798 | 1798
2008-08-04 21:58 | 1815 | 1815 |  |  | 2008-08-04 22:00 | 1800 | 1800
2008-08-04 22:03 | 1819 | 1819 |  |  | 2008-08-04 22:05 | 1802 | 1802
2008-08-04 22:08 | 1791 |  |  |  | 2008-08-04 22:10 | 1719 |  | 1719
2008-08-04 22:13 | 1739 |  | 1739 |  | 2008-08-04 22:15 | 1724 |  | 1724
2008-08-04 22:18 | 1742 |  | 1742 |  | 2008-08-04 22:20 | 1724 |  | 1724
2008-08-04 22:23 | 1769 |  |  |  | 2008-08-04 22:25 | 1798 | 1798
2008-08-04 22:28 | 1819 | 1819 |  |  | 2008-08-04 22:30 | 1799 | 1799
2008-08-04 22:33 | 1817 | 1817 |  |  | 2008-08-04 22:35 | 1800 | 1800
2008-08-04 22:38 | 1084 |  |  |  | 2008-08-04 22:40 | 6.168
2008-08-04 22:43 | 0 |  |  |  | 2008-08-04 22:45 | 0
2008-08-04 22:48 | 0 |  |  |  | 2008-08-04 22:50 | 0
2008-08-04 22:53 | 0 |  |  |  | 2008-08-04 22:55 | 0
2008-08-04 22:58 | 0 |  |  |  | 2008-08-04 23:00 | 0
2008-08-04 23:03 | 0 |  |  |  | 2008-08-04 23:05 | 0
2008-08-04 23:08 | 0 |  |  |  | 2008-08-04 23:10 | 0
2008-08-04 23:13 | 0 |  |  |  | 2008-08-04 23:15 | 0
2008-08-04 23:18 | 0 |  |  |  | 2008-08-04 23:20 | 0
2008-08-04 23:23 | 1286 |  |  |  | 2008-08-04 23:25 | 3089
2008-08-04 23:28 | 3159 | 3159 |  |  | 2008-08-04 23:30 | 3132 | 3132
2008-08-04 23:33 | 3133 |  |  |  | 2008-08-04 23:35 | 3036 |  | 3036
2008-08-04 23:38 | 3065 |  | 3065 |  | 2008-08-04 23:40 | 3040 |  | 3040
2008-08-04 23:43 | 3076 |  | 3076 |  | 2008-08-04 23:45 | 3043 |  | 3043
2008-08-04 23:48 | 3111 |  |  |  | 2008-08-04 23:50 | 3145 | 3145
2008-08-04 23:53 | 3176 | 3176 |  |  | 2008-08-04 23:55 | 3145 | 3145
2008-08-04 23:58 | 3175 | 3175 |  |  | 2008-08-05 | 3146 | 3146
2008-08-05 00:03 | 3144 |  |  |  | 2008-08-05 00:05 | 3046 |  | 3046
2008-08-05 00:08 | 3069 |  | 3069 |  | 2008-08-05 00:10 | 3047 |  | 3047
2008-08-05 00:13 | 3074 |  | 3074 |  | 2008-08-05 00:15 | 3048 |  | 3048
2008-08-05 00:18 | 3119 |  |  |  | 2008-08-05 00:20 | 3142 | 3142
2008-08-05 00:23 | 3172 | 3172 |  |  | 2008-08-05 00:25 | 3146 | 3146
2008-08-05 00:28 | 3172 | 3172 |  |  | 2008-08-05 00:30 | 3143 | 3143
2008-08-05 00:33 | 3141 |  |  |  | 2008-08-05 00:35 | 3048 |  | 3048
2008-08-05 00:38 | 3080 |  | 3080 |  | 2008-08-05 00:40 | 3053 |  | 3053
2008-08-05 00:43 | 3083 |  | 3083 |  | 2008-08-05 00:45 | 3053 |  | 3053
2008-08-05 00:48 | 3119 |  |  |  | 2008-08-05 00:50 | 3151 | 3151
2008-08-05 00:53 | 3176 | 3176 |  |  | 2008-08-05 00:55 | 3139 | 3139
2008-08-05 00:58 | 3172 | 3172 |  |  | 2008-08-05 01:00 | 3144 | 3144
2008-08-05 01:03 | 3144 |  |  |  | 2008-08-05 01:05 | 3047 |  | 3047
2008-08-05 01:08 | 3069 |  | 3069 |  | 2008-08-05 01:10 | 3049 |  | 3049
2008-08-05 01:13 | 3079 |  | 3079 |  | 2008-08-05 01:15 | 3050 |  | 3050
2008-08-05 01:18 | 3115 |  |  |  | 2008-08-05 01:20 | 3148 | 3148
2008-08-05 01:23 | 3175 | 3175 |  |  | 2008-08-05 01:25 | 3149 | 3149
2008-08-05 01:28 | 3179 | 3179 |  |  | 2008-08-05 01:30 | 3148 | 3148
2008-08-05 01:33 | 3142 |  |  |  | 2008-08-05 01:35 | 3051 |  | 3051
2008-08-05 01:38 | 3072 |  | 3072 |  | 2008-08-05 01:40 | 3049 |  | 3049
2008-08-05 01:43 | 3076 |  | 3076 |  | 2008-08-05 01:45 | 3052 |  | 3052
2008-08-05 01:48 | 3114 |  |  |  | 2008-08-05 01:50 | 3141 | 3141
2008-08-05 01:53 | 3176 | 3176 |  |  | 2008-08-05 01:55 | 3142 | 3142
2008-08-05 01:58 | 3177 | 3177 |  |  | 2008-08-05 02:00 | 3146 | 3146
2008-08-05 02:03 | 3171 | 3171 |  |  | 2008-08-05 02:05 | 3147 | 3147
2008-08-05 02:08 | 3175 | 3175 |  |  | 2008-08-05 02:10 | 3149 | 3149
2008-08-05 02:13 | 3174 | 3174 |  |  | 2008-08-05 02:15 | 3141 | 3141
2008-08-05 02:18 | 3177 | 3177 |  |  | 2008-08-05 02:20 | 3144 | 3144
2008-08-05 02:23 | 3174 | 3174 |  |  | 2008-08-05 02:25 | 3146 | 3146
2008-08-05 02:28 | 3178 | 3178 |  |  | 2008-08-05 02:30 | 3145 | 3145
2008-08-05 02:33 | 3173 |  |  |  | 2008-08-05 02:35 | 3150
2008-08-05 02:38 | 4590 |  |  |  | 2008-08-05 02:40 | 6632 | 6632
2008-08-05 02:43 | 6698 | 6698 |  |  | 2008-08-05 02:45 | 6637 | 6637
2008-08-05 02:48 | 6704 | 6704 |  |  | 2008-08-05 02:50 | 6647 | 6647
2008-08-05 02:53 | 6646 |  |  |  | 2008-08-05 02:55 | 6502 |  | 6502
2008-08-05 02:58 | 6548 |  | 6548 |  | 2008-08-05 03:00 | 6489 |  | 6489
2008-08-05 03:03 | 6558 |  | 6558 |  | 2008-08-05 03:05 | 6498 |  | 6498
2008-08-05 03:08 | 6619 |  |  |  | 2008-08-05 03:10 | 6652 | 6652
2008-08-05 03:13 | 6711 | 6711 |  |  | 2008-08-05 03:15 | 6656 | 6656
2008-08-05 03:18 | 6715 | 6715 |  |  | 2008-08-05 03:20 | 6654 | 6654
2008-08-05 03:23 | 6653 |  |  |  | 2008-08-05 03:25 | 6503 |  | 6503
2008-08-05 03:28 | 6554 |  | 6554 |  | 2008-08-05 03:30 | 6484 |  | 6484
2008-08-05 03:33 | 6543 |  | 6543 |  | 2008-08-05 03:35 | 6493 |  | 6493
2008-08-05 03:38 | 6618 |  |  |  | 2008-08-05 03:40 | 6659 | 6659
2008-08-05 03:43 | 6717 | 6717 |  |  | 2008-08-05 03:45 | 6653 | 6653
2008-08-05 03:48 | 6709 | 6709 |  |  | 2008-08-05 03:50 | 6654 | 6654
2008-08-05 03:53 | 6656 |  |  |  | 2008-08-05 03:55 | 6505 |  | 6505
2008-08-05 03:58 | 6558 |  | 6558 |  | 2008-08-05 04:00 | 6506 |  | 6506
2008-08-05 04:03 | 6560 |  | 6560 |  | 2008-08-05 04:05 | 6516 |  | 6516
2008-08-05 04:08 | 6626 |  |  |  | 2008-08-05 04:10 | 6661 | 6661
2008-08-05 04:13 | 6717 | 6717 |  |  | 2008-08-05 04:15 | 6655 | 6655
2008-08-05 04:18 | 6711 | 6711 |  |  | 2008-08-05 04:20 | 6655 | 6655
2008-08-05 04:23 | 6659 |  |  |  | 2008-08-05 04:25 | 6498 |  | 6498
2008-08-05 04:28 | 6555 |  | 6555 |  | 2008-08-05 04:30 | 6493 |  | 6493
2008-08-05 04:33 | 6561 |  | 6561 |  | 2008-08-05 04:35 | 6498 |  | 6498
2008-08-05 04:38 | 6611 |  |  |  | 2008-08-05 04:40 | 6662 | 6662
2008-08-05 04:43 | 6717 | 6717 |  |  | 2008-08-05 04:45 | 6654 | 6654
2008-08-05 04:48 | 6718 | 6718 |  |  | 2008-08-05 04:50 | 6657 | 6657
2008-08-05 04:53 | 6670 |  |  |  | 2008-08-05 04:55 | 6505 |  | 6505
2008-08-05 04:58 | 6567 |  | 6567 |  | 2008-08-05 05:00 | 6502 |  | 6502
2008-08-05 05:03 | 6554 |  | 6554 |  | 2008-08-05 05:05 | 6499 |  | 6499
2008-08-05 05:08 | 6618 |  |  |  | 2008-08-05 05:10 | 6658 | 6658
2008-08-05 05:13 | 6713 | 6713 |  |  | 2008-08-05 05:15 | 6663 | 6663
2008-08-05 05:18 | 6715 | 6715 |  |  | 2008-08-05 05:20 | 6667 | 6667
2008-08-05 05:23 | 6667 |  |  |  | 2008-08-05 05:25 | 6507 |  | 6507
2008-08-05 05:28 | 6569 |  | 6569 |  | 2008-08-05 05:30 | 6513 |  | 6513
2008-08-05 05:33 | 6565 |  | 6565 |  | 2008-08-05 05:35 | 6519 |  | 6519
2008-08-05 05:38 | 6617 |  |  |  | 2008-08-05 05:40 | 6656 | 6656
2008-08-05 05:43 | 6715 | 6715 |  |  | 2008-08-05 05:45 | 6656 | 6656
2008-08-05 05:48 | 6714 | 6714 |  |  | 2008-08-05 05:50 | 6658 | 6658
2008-08-05 05:53 | 8475 |  |  |  | 2008-08-05 05:55 | 10943 | 10943
2008-08-05 05:58 | 11069 | 11069 |  |  | 2008-08-05 06:00 | 10980 | 10980
2008-08-05 06:03 | 11084 | 11084 |  |  | 2008-08-05 06:05 | 10964 | 10964
2008-08-05 06:08 | 10973 |  |  |  | 2008-08-05 06:10 | 10731 |  | 10731
2008-08-05 06:13 | 10811 |  | 10811 |  | 2008-08-05 06:15 | 10710 |  | 10710
2008-08-05 06:18 | 10768 |  | 10768 |  | 2008-08-05 06:20 | 10667 |  | 10667
2008-08-05 06:23 | 10851 |  |  |  | 2008-08-05 06:25 | 10895 | 10895
2008-08-05 06:28 | 10969 | 10969 |  |  | 2008-08-05 06:30 | 10849 | 10849
2008-08-05 06:33 | 10968 | 10968 |  |  | 2008-08-05 06:35 | 10882 | 10882
2008-08-05 06:38 | 10841 |  |  |  | 2008-08-05 06:40 | 10572 |  | 10572
2008-08-05 06:43 | 10655 |  | 10655 |  | 2008-08-05 06:45 | 10594 |  | 10594
2008-08-05 06:48 | 10686 |  | 10686 |  | 2008-08-05 06:50 | 10595 |  | 10595
2008-08-05 06:53 | 10769 |  |  |  | 2008-08-05 06:55 | 10794 | 10794
2008-08-05 06:58 | 10907 | 10907 |  |  | 2008-08-05 07:00 | 10815 | 10815
2008-08-05 07:03 | 10894 | 10894 |  |  | 2008-08-05 07:05 | 10785 | 10785
2008-08-05 07:08 | 10794 |  |  |  | 2008-08-05 07:10 | 10562 |  | 10562
2008-08-05 07:13 | 10667 |  | 10667 |  | 2008-08-05 07:15 | 10553 |  | 10553
2008-08-05 07:18 | 10657 |  | 10657 |  | 2008-08-05 07:20 | 10574 |  | 10574
2008-08-05 07:23 | 10772 |  |  |  | 2008-08-05 07:25 | 10818 | 10818
2008-08-05 07:28 | 10874 | 10874 |  |  | 2008-08-05 07:30 | 10794 | 10794
2008-08-05 07:33 | 10922 | 10922 |  |  | 2008-08-05 07:35 | 10828 | 10828
2008-08-05 07:38 | 10809 |  |  |  | 2008-08-05 07:40 | 10581 |  | 10581
2008-08-05 07:43 | 10663 |  | 10663 |  | 2008-08-05 07:45 | 10559 |  | 10559
2008-08-05 07:48 | 10660 |  | 10660 |  | 2008-08-05 07:50 | 10606 |  | 10606
2008-08-05 07:53 | 10778 |  |  |  | 2008-08-05 07:55 | 10809 | 10809
2008-08-05 07:58 | 10919 | 10919 |  |  | 2008-08-05 08:00 | 10837 | 10837
2008-08-05 08:03 | 10909 | 10909 |  |  | 2008-08-05 08:05 | 10788 | 10788
2008-08-05 08:08 | 10794 |  |  |  | 2008-08-05 08:10 | 10570 |  | 10570
2008-08-05 08:13 | 10662 |  | 10662 |  | 2008-08-05 08:15 | 10563 |  | 10563
2008-08-05 08:18 | 10632 |  | 10632 |  | 2008-08-05 08:20 | 10552 |  | 10552
2008-08-05 08:23 | 10741 |  |  |  | 2008-08-05 08:25 | 10793 | 10793
2008-08-05 08:28 | 10898 | 10898 |  |  | 2008-08-05 08:30 | 10824 | 10824
2008-08-05 08:33 | 10891 | 10891 |  |  | 2008-08-05 08:35 | 10783 | 10783
2008-08-05 08:38 | 10800 |  |  |  | 2008-08-05 08:40 | 10576 |  | 10576
2008-08-05 08:43 | 10655 |  | 10655 |  | 2008-08-05 08:45 | 10555 |  | 10555
2008-08-05 08:48 | 10609 |  | 10609 |  | 2008-08-05 08:50 | 10543 |  | 10543
2008-08-05 08:53 | 10733 |  |  |  | 2008-08-05 08:55 | 10792 | 10792
2008-08-05 08:58 | 10882 | 10882 |  |  | 2008-08-05 09:00 | 10792 | 10792
2008-08-05 09:03 | 10886 | 10886 |  |  | 2008-08-05 09:05 | 10799 | 10799
2008-08-05 09:08 | 6516 |  |  |  | 2008-08-05 09:10 | 51.193
2008-08-05 09:13 | 31.177 |  |  |  | 2008-08-05 09:15 | 19.316
2008-08-05 09:18 | 12.791 |  |  |  | 2008-08-05 09:20 | 12.835
2008-08-05 09:23 | 10.244 |  |  |  | 2008-08-05 09:25 | 7.951
2008-08-05 09:28 | 5.867 |  |  |  | 2008-08-05 09:30 | 5.678
2008-08-05 09:33 | 6.138 |  |  |  | 2008-08-05 09:35 | 4.114
2008-08-05 09:38 | 8.188 |  |  |  | 2008-08-05 09:40 | 5.361
2008-08-05 09:43 | 3.437 |  |  |  | 2008-08-05 09:45 | 0
2008-08-05 09:48 | 3.954 |  |  |  | 2008-08-05 09:50 | 4.048
2008-08-05 09:53 | 7494 |  |  |  | 2008-08-05 09:55 | 20409
2008-08-05 09:58 | 20643 |  | 20643 |  | 2008-08-05 10:00 | 20471
2008-08-05 10:03 | 20683 |  | 20683 |  | 2008-08-05 10:05 | 20538 | 20538
2008-08-05 10:08 | 20914 |  |  |  | 2008-08-05 10:10 | 20984 | 20984
2008-08-05 10:13 | 21117 | 21117 |  |  | 2008-08-05 10:15 | 20950 | 20950
2008-08-05 10:18 | 21157 | 21157 |  |  | 2008-08-05 10:20 | 20968 | 20968
2008-08-05 10:23 | 21006 |  |  |  | 2008-08-05 10:25 | 20539 |  | 20539
2008-08-05 10:28 | 20752 |  | 20752 |  | 2008-08-05 10:30 | 20550 |  | 20550
2008-08-05 10:33 | 20740 |  | 20740 |  | 2008-08-05 10:35 | 20508 |  | 20508
2008-08-05 10:38 | 20889 |  |  |  | 2008-08-05 10:40 | 21014 | 21014
2008-08-05 10:43 | 21189 | 21189 |  |  | 2008-08-05 10:45 | 20994 | 20994
2008-08-05 10:48 | 21161 | 21161 |  |  | 2008-08-05 10:50 | 20981 | 20981
2008-08-05 10:53 | 21062 |  |  |  | 2008-08-05 10:55 | 20584 |  | 20584
2008-08-05 10:58 | 20796 |  | 20796 |  | 2008-08-05 11:00 | 20596 |  | 20596
2008-08-05 11:03 | 20768 |  | 20768 |  | 2008-08-05 11:05 | 20614 |  | 20614
2008-08-05 11:08 | 20999 |  |  |  | 2008-08-05 11:10 | 21075 | 21075
2008-08-05 11:13 | 21216 | 21216 |  |  | 2008-08-05 11:15 | 21049 | 21049
2008-08-05 11:18 | 21258 | 21258 |  |  | 2008-08-05 11:20 | 21064 | 21064
2008-08-05 11:23 | 21078 |  |  |  | 2008-08-05 11:25 | 20615 |  | 20615
2008-08-05 11:28 | 20824 |  | 20824 |  | 2008-08-05 11:30 | 20609 |  | 20609
2008-08-05 11:33 | 20765 |  | 20765 |  | 2008-08-05 11:35 | 20535 |  | 20535
2008-08-05 11:38 | 20950 |  |  |  | 2008-08-05 11:40 | 21047 | 21047
2008-08-05 11:43 | 21201 | 21201 |  |  | 2008-08-05 11:45 | 21005 | 21005
2008-08-05 11:48 | 21233 | 21233 |  |  | 2008-08-05 11:50 | 21018 | 21018
2008-08-05 11:53 | 21003 |  |  |  | 2008-08-05 11:55 | 20557 |  | 20557
2008-08-05 11:58 | 20743 |  | 20743 |  | 2008-08-05 12:00 | 20529 |  | 20529
2008-08-05 12:03 | 20742 |  | 20742 |  | 2008-08-05 12:05 | 20546 |  | 20546
2008-08-05 12:08 | 20883 |  |  |  | 2008-08-05 12:10 | 20989 | 20989
2008-08-05 12:13 | 21162 | 21162 |  |  | 2008-08-05 12:15 | 20957 | 20957
2008-08-05 12:18 | 21177 | 21177 |  |  | 2008-08-05 12:20 | 20981 | 20981
2008-08-05 12:23 | 21135 | 21135 |  |  | 2008-08-05 12:25 | 20932 | 20932
2008-08-05 12:28 | 21156 | 21156 |  |  | 2008-08-05 12:30 | 20934 | 20934
2008-08-05 12:33 | 21102 | 21102 |  |  | 2008-08-05 12:35 | 20924 | 20924
2008-08-05 12:38 | 21142 | 21142 |  |  | 2008-08-05 12:40 | 20969 | 20969
2008-08-05 12:43 | 21125 | 21125 |  |  | 2008-08-05 12:45 | 20923 | 20923
2008-08-05 12:48 | 21168 | 21168 |  |  | 2008-08-05 12:50 | 20967 | 20967
2008-08-05 12:53 | 21130 | 21130 |  |  | 2008-08-05 12:55 | 20897 | 20897
2008-08-05 12:58 | 21135 | 21135 |  |  | 2008-08-05 13:00 | 20957 | 20957
2008-08-05 13:03 | 21154 | 21154 |  |  | 2008-08-05 13:05 | 20894 | 20894
2008-08-05 13:08 | 13012 |  |  |  | 2008-08-05 13:10 | 109.619
2008-08-05 13:13 | 65.209 |  |  |  | 2008-08-05 13:15 | 44.749
2008-08-05 13:18 | 37.245 |  |  |  | 2008-08-05 13:20 | 29.832
2008-08-05 13:23 | 25.512 |  |  |  | 2008-08-05 13:25 | 21.756
2008-08-05 13:28 | 21.247 |  |  |  | 2008-08-05 13:30 | 16.561
2008-08-05 13:33 | 17.725 |  |  |  | 2008-08-05 13:35 | 12.091
2008-08-05 13:38 | 11.125 |  |  |  | 2008-08-05 13:40 | 10.757
2008-08-05 13:43 | 12.423 |  |  |  | 2008-08-05 13:45 | 8.976
2008-08-05 13:48 | 14.005 |  |  |  | 2008-08-05 13:50 | 9.116
2008-08-05 13:53 | 16623 |  |  |  | 2008-08-05 13:55 | 39713
2008-08-05 13:58 | 40253 | 40253 |  |  | 2008-08-05 14:00 | 39880
2008-08-05 14:03 | 40098 |  |  |  | 2008-08-05 14:05 | 39304 |  | 39304
2008-08-05 14:08 | 39746 |  | 39746 |  | 2008-08-05 14:10 | 39325 |  | 39325
2008-08-05 14:13 | 39815 |  | 39815 |  | 2008-08-05 14:15 | 39414 |  | 39414
2008-08-05 14:18 | 40003 |  |  |  | 2008-08-05 14:20 | 40041 | 40041
2008-08-05 14:23 | 40457 | 40457 |  |  | 2008-08-05 14:25 | 40078 | 40078
2008-08-05 14:28 | 40456 | 40456 |  |  | 2008-08-05 14:30 | 40033 | 40033
2008-08-05 14:33 | 40195 |  |  |  | 2008-08-05 14:35 | 39365 |  | 39365
2008-08-05 14:38 | 39843 |  | 39843 |  | 2008-08-05 14:40 | 39460 |  | 39460
2008-08-05 14:43 | 39796 |  | 39796 |  | 2008-08-05 14:45 | 39408 |  | 39408
2008-08-05 14:48 | 40070 |  |  |  | 2008-08-05 14:50 | 40148 | 40148
2008-08-05 14:53 | 40549 | 40549 |  |  | 2008-08-05 14:55 | 40047 | 40047
2008-08-05 14:58 | 40454 | 40454 |  |  | 2008-08-05 15:00 | 40142 | 40142
2008-08-05 15:03 | 40353 |  |  |  | 2008-08-05 15:05 | 39497 |  | 39497
2008-08-05 15:08 | 39836 |  | 39836 |  | 2008-08-05 15:10 | 39447 |  | 39447
2008-08-05 15:13 | 39901 |  | 39901 |  | 2008-08-05 15:15 | 39453 |  | 39453
2008-08-05 15:18 | 40053 |  |  |  | 2008-08-05 15:20 | 40053 | 40053
2008-08-05 15:23 | 40512 | 40512 |  |  | 2008-08-05 15:25 | 40115 | 40115
2008-08-05 15:28 | 40538 | 40538 |  |  | 2008-08-05 15:30 | 40055 | 40055
2008-08-05 15:33 | 40269 |  |  |  | 2008-08-05 15:35 | 39495 |  | 39495
2008-08-05 15:38 | 39859 |  | 39859 |  | 2008-08-05 15:40 | 39390 |  | 39390
2008-08-05 15:43 | 39768 |  | 39768 |  | 2008-08-05 15:45 | 39386 |  | 39386
2008-08-05 15:48 | 39990 |  |  |  | 2008-08-05 15:50 | 39987 | 39987
2008-08-05 15:53 | 40414 | 40414 |  |  | 2008-08-05 15:55 | 40078 | 40078
2008-08-05 15:58 | 40471 | 40471 |  |  | 2008-08-05 16:00 | 40000 | 40000
2008-08-05 16:03 | 40178 |  |  |  | 2008-08-05 16:05 | 39362 |  | 39362
2008-08-05 16:08 | 39761 |  | 39761 |  | 2008-08-05 16:10 | 39415 |  | 39415
2008-08-05 16:13 | 39755 |  | 39755 |  | 2008-08-05 16:15 | 39428 |  | 39428
2008-08-05 16:18 | 40081 |  |  |  | 2008-08-05 16:20 | 40032 | 40032
2008-08-05 16:23 | 40462 | 40462 |  |  | 2008-08-05 16:25 | 40133 | 40133
2008-08-05 16:28 | 40457 | 40457 |  |  | 2008-08-05 16:30 | 39996 | 39996
2008-08-05 16:33 | 40391 | 40391 |  |  | 2008-08-05 16:35 | 40054 | 40054
2008-08-05 16:38 | 40452 | 40452 |  |  | 2008-08-05 16:40 | 40018 | 40018
2008-08-05 16:43 | 40392 | 40392 |  |  | 2008-08-05 16:45 | 40038 | 40038
2008-08-05 16:48 | 40484 |  |  |  | 2008-08-05 16:50 | 40060 | 40060
2008-08-05 16:53 | 39754 |  | 39754 |  | 2008-08-05 16:55 | 39383 |  | 39383
2008-08-05 16:58 | 39757 |  | 39757 |  | 2008-08-05 17:00 | 39331 |  | 39331
2008-08-05 17:03 | 39813 |  |  |  | 2008-08-05 17:05 | 39399 |  | 39399
2008-08-05 17:08 | 40387 |  |  |  | 2008-08-05 17:10 | 40003 | 40003
2008-08-05 17:13 | 40484 | 40484 |  |  | 2008-08-05 17:15 | 40216 | 40216
2008-08-05 17:18 | 40498 | 40498 |  |  | 2008-08-05 17:20 | 30847
2008-08-05 17:23 | 364.989 |  |  |  | 2008-08-05 17:25 | 122.5
2008-08-05 17:28 | 82.33 |  |  |  | 2008-08-05 17:30 | 63.573
2008-08-05 17:33 | 51.571 |  |  |  | 2008-08-05 17:35 | 41.145
2008-08-05 17:38 | 38.063 |  |  |  | 2008-08-05 17:40 | 32.868
2008-08-05 17:43 | 32.542 |  |  |  | 2008-08-05 17:45 | 28.624
2008-08-05 17:48 | 29.012 |  |  |  | 2008-08-05 17:50 | 22.305
2008-08-05 17:53 | 25.475 |  |  |  | 2008-08-05 17:55 | 22.22
2008-08-05 17:58 | 21.235 |  |  |  | 2008-08-05 18:00 | 16.846
2008-08-05 18:03 | 19.614 |  |  | 19.614 | 2008-08-05 18:05 | 16.749 |  |  | 16.749
2008-08-05 18:08 | 16.125 |  |  | 16.125 | 2008-08-05 18:10 | 15.83 |  |  | 15.83
2008-08-05 18:13 | 556.162 |  |  |  | 2008-08-05 18:15 | 1776
2008-08-05 18:18 | 1804 | 1804 |  |  | 2008-08-05 18:20 | 1790 | 1790
2008-08-05 18:23 | 1812 | 1812 |  |  | 2008-08-05 18:25 | 1797 | 1797
2008-08-05 18:28 | 1791 |  |  |  | 2008-08-05 18:30 | 1727 |  | 1727
2008-08-05 18:33 | 1749 |  | 1749 |  | 2008-08-05 18:35 | 1728 |  | 1728
2008-08-05 18:38 | 1745 |  | 1745 |  | 2008-08-05 18:40 | 1726 |  | 1726
2008-08-05 18:43 | 1778 |  |  |  | 2008-08-05 18:45 | 1799 | 1799
2008-08-05 18:48 | 1817 | 1817 |  |  | 2008-08-05 18:50 | 1795 | 1795
2008-08-05 18:53 | 1818 | 1818 |  |  | 2008-08-05 18:55 | 1800 | 1800
2008-08-05 18:58 | 1788 |  |  |  | 2008-08-05 19:00 | 1726 |  | 1726
2008-08-05 19:03 | 1744 |  | 1744 |  | 2008-08-05 19:05 | 1731 |  | 1731
2008-08-05 19:08 | 1747 |  | 1747 |  | 2008-08-05 19:10 | 1730 |  | 1730
2008-08-05 19:13 | 1765 |  |  |  | 2008-08-05 19:15 | 1799 | 1799
2008-08-05 19:18 | 1813 | 1813 |  |  | 2008-08-05 19:20 | 1799 | 1799
2008-08-05 19:23 | 1820 | 1820 |  |  | 2008-08-05 19:25 | 1797 | 1797
2008-08-05 19:28 | 1789 |  |  |  | 2008-08-05 19:30 | 1724 |  | 1724
2008-08-05 19:33 | 1741 |  | 1741 |  | 2008-08-05 19:35 | 1727 |  | 1727
2008-08-05 19:38 | 1742 |  | 1742 |  | 2008-08-05 19:40 | 1725 |  | 1725
2008-08-05 19:43 | 1767 |  |  |  | 2008-08-05 19:45 | 1800 | 1800
2008-08-05 19:48 | 1813 | 1813 |  |  | 2008-08-05 19:50 | 1799 | 1799
2008-08-05 19:53 | 1809 | 1809 |  |  | 2008-08-05 19:55 | 1797 | 1797
2008-08-05 19:58 | 1791 |  |  |  | 2008-08-05 20:00 | 1724 |  | 1724
2008-08-05 20:03 | 1744 |  | 1744 |  | 2008-08-05 20:05 | 1723 |  | 1723
2008-08-05 20:08 | 1744 |  | 1744 |  | 2008-08-05 20:10 | 1724 |  | 1724
2008-08-05 20:13 | 1768 |  |  |  | 2008-08-05 20:15 | 1797 | 1797
2008-08-05 20:18 | 1810 | 1810 |  |  | 2008-08-05 20:20 | 1796 | 1796
2008-08-05 20:23 | 1813 | 1813 |  |  | 2008-08-05 20:25 | 1793 | 1793
2008-08-05 20:28 | 1788 |  |  |  | 2008-08-05 20:30 | 1723 |  | 1723
2008-08-05 20:33 | 1742 |  | 1742 |  | 2008-08-05 20:35 | 1723 |  | 1723
2008-08-05 20:38 | 1743 |  | 1743 |  | 2008-08-05 20:40 | 1725 |  | 1725
2008-08-05 20:43 | 1764 |  |  |  | 2008-08-05 20:45 | 1793 | 1793
2008-08-05 20:48 | 1812 | 1812 |  |  | 2008-08-05 20:50 | 1796 | 1796
2008-08-05 20:53 | 1809 | 1809 |  |  | 2008-08-05 20:55 | 1796 | 1796
2008-08-05 20:58 | 1789 |  |  |  | 2008-08-05 21:00 | 1717 |  | 1717
2008-08-05 21:03 | 1740 |  | 1740 |  | 2008-08-05 21:05 | 1722 |  | 1722
2008-08-05 21:08 | 1739 |  | 1739 |  | 2008-08-05 21:10 | 1723 |  | 1723
2008-08-05 21:13 | 1761 |  |  |  | 2008-08-05 21:15 | 1792 | 1792
2008-08-05 21:18 | 1811 | 1811 |  |  | 2008-08-05 21:20 | 1796 | 1796
2008-08-05 21:23 | 1810 | 1810 |  |  | 2008-08-05 21:25 | 1796 | 1796
2008-08-05 21:28 | 1146 |  |  |  | 2008-08-05 21:30 | 10.926
2008-08-05 21:33 | 0 |  |  |  | 2008-08-05 21:35 | 6.102
2008-08-05 21:38 | 0 |  |  |  | 2008-08-05 21:40 | 5.47
2008-08-05 21:43 | 0 |  |  |  | 2008-08-05 21:45 | 3.134
2008-08-05 21:48 | 0 |  |  |  | 2008-08-05 21:50 | 3.434
2008-08-05 21:53 | 4.329 |  |  |  | 2008-08-05 21:55 | 0
2008-08-05 21:58 | 0 |  |  |  | 2008-08-05 22:00 | 3.565
2008-08-05 22:03 | 0 |  |  |  | 2008-08-05 22:05 | 0
2008-08-05 22:08 | 0 |  |  |  | 2008-08-05 22:10 | 0
2008-08-05 22:13 | 1199 |  |  |  | 2008-08-05 22:15 | 3080
2008-08-05 22:18 | 3133 | 3133 |  |  | 2008-08-05 22:20 | 3112 | 3112
2008-08-05 22:23 | 3150 | 3150 |  |  | 2008-08-05 22:25 | 3124 | 3124
2008-08-05 22:28 | 3115 |  |  |  | 2008-08-05 22:30 | 3023 |  | 3023
2008-08-05 22:33 | 3050 |  | 3050 |  | 2008-08-05 22:35 | 3027 |  | 3027
2008-08-05 22:38 | 3054 |  | 3054 |  | 2008-08-05 22:40 | 3031 |  | 3031
2008-08-05 22:43 | 3098 |  |  |  | 2008-08-05 22:45 | 3127 | 3127
2008-08-05 22:48 | 3153 | 3153 |  |  | 2008-08-05 22:50 | 3127 | 3127
2008-08-05 22:53 | 3157 | 3157 |  |  | 2008-08-05 22:55 | 3133 | 3133
2008-08-05 22:58 | 3121 |  |  |  | 2008-08-05 23:00 | 3022 |  | 3022
2008-08-05 23:03 | 3051 |  | 3051 |  | 2008-08-05 23:05 | 3025 |  | 3025
2008-08-05 23:08 | 3056 |  | 3056 |  | 2008-08-05 23:10 | 3027 |  | 3027
2008-08-05 23:13 | 3094 |  |  |  | 2008-08-05 23:15 | 3133 | 3133
2008-08-05 23:18 | 3159 | 3159 |  |  | 2008-08-05 23:20 | 3134 | 3134
2008-08-05 23:23 | 3156 | 3156 |  |  | 2008-08-05 23:25 | 3129 | 3129
2008-08-05 23:28 | 3125 |  |  |  | 2008-08-05 23:30 | 3023 |  | 3023
2008-08-05 23:33 | 3056 |  | 3056 |  | 2008-08-05 23:35 | 3028 |  | 3028
2008-08-05 23:38 | 3060 |  | 3060 |  | 2008-08-05 23:40 | 3029 |  | 3029
2008-08-05 23:43 | 3098 |  |  |  | 2008-08-05 23:45 | 3131 | 3131
2008-08-05 23:48 | 3154 | 3154 |  |  | 2008-08-05 23:50 | 3131 | 3131
2008-08-05 23:53 | 3153 | 3153 |  |  | 2008-08-05 23:55 | 3127 | 3127
2008-08-05 23:58 | 3127 |  |  |  | 2008-08-06 | 3028 |  | 3028
2008-08-06 00:03 | 3058 |  | 3058 |  | 2008-08-06 00:05 | 3031 |  | 3031
2008-08-06 00:08 | 3063 |  | 3063 |  | 2008-08-06 00:10 | 3034 |  | 3034
2008-08-06 00:13 | 3101 |  |  |  | 2008-08-06 00:15 | 3134 | 3134
2008-08-06 00:18 | 3160 | 3160 |  |  | 2008-08-06 00:20 | 3137 | 3137
2008-08-06 00:23 | 3155 | 3155 |  |  | 2008-08-06 00:25 | 3133 | 3133
2008-08-06 00:28 | 3136 |  |  |  | 2008-08-06 00:30 | 3026 |  | 3026
2008-08-06 00:33 | 3058 |  | 3058 |  | 2008-08-06 00:35 | 3031 |  | 3031
2008-08-06 00:38 | 3061 |  | 3061 |  | 2008-08-06 00:40 | 3031 |  | 3031
2008-08-06 00:43 | 3096 |  |  |  | 2008-08-06 00:45 | 3127 | 3127
2008-08-06 00:48 | 3159 | 3159 |  |  | 2008-08-06 00:50 | 3137 | 3137
2008-08-06 00:53 | 3158 | 3158 |  |  | 2008-08-06 00:55 | 3128 | 3128
2008-08-06 00:58 | 3131 |  |  |  | 2008-08-06 01:00 | 3038 |  | 3038
2008-08-06 01:03 | 3063 |  | 3063 |  | 2008-08-06 01:05 | 3032 |  | 3032
2008-08-06 01:08 | 3060 |  | 3060 |  | 2008-08-06 01:10 | 3033 |  | 3033
2008-08-06 01:13 | 3094 |  |  |  | 2008-08-06 01:15 | 3129 | 3129
2008-08-06 01:18 | 3156 | 3156 |  |  | 2008-08-06 01:20 | 3125 | 3125
2008-08-06 01:23 | 3158 | 3158 |  |  | 2008-08-06 01:25 | 3122 | 3122
2008-08-06 01:28 | 4468 |  |  |  | 2008-08-06 01:30 | 6594 | 6594
2008-08-06 01:33 | 6662 | 6662 |  |  | 2008-08-06 01:35 | 6602 | 6602
2008-08-06 01:38 | 6668 | 6668 |  |  | 2008-08-06 01:40 | 6608 | 6608
2008-08-06 01:43 | 6607 |  |  |  | 2008-08-06 01:45 | 6465 |  | 6465
2008-08-06 01:48 | 6528 |  | 6528 |  | 2008-08-06 01:50 | 6475 |  | 6475
2008-08-06 01:53 | 6526 |  | 6526 |  | 2008-08-06 01:55 | 6467 |  | 6467
2008-08-06 01:58 | 6575 |  |  |  | 2008-08-06 02:00 | 6617 | 6617
2008-08-06 02:03 | 6680 | 6680 |  |  | 2008-08-06 02:05 | 6615 | 6615
2008-08-06 02:08 | 6673 | 6673 |  |  | 2008-08-06 02:10 | 6614 | 6614
2008-08-06 02:13 | 6626 |  |  |  | 2008-08-06 02:15 | 6459 |  | 6459
2008-08-06 02:18 | 6522 |  | 6522 |  | 2008-08-06 02:20 | 6462 |  | 6462
2008-08-06 02:23 | 6521 |  | 6521 |  | 2008-08-06 02:25 | 6469 |  | 6469
2008-08-06 02:28 | 6581 |  |  |  | 2008-08-06 02:30 | 6610 | 6610
2008-08-06 02:33 | 6677 | 6677 |  |  | 2008-08-06 02:35 | 6605 | 6605
2008-08-06 02:38 | 6674 | 6674 |  |  | 2008-08-06 02:40 | 6615 | 6615
2008-08-06 02:43 | 6625 |  |  |  | 2008-08-06 02:45 | 6472 |  | 6472
2008-08-06 02:48 | 6533 |  | 6533 |  | 2008-08-06 02:50 | 6473 |  | 6473
2008-08-06 02:53 | 6526 |  | 6526 |  | 2008-08-06 02:55 | 6476 |  | 6476
2008-08-06 02:58 | 6581 |  |  |  | 2008-08-06 03:00 | 6619 | 6619
2008-08-06 03:03 | 6676 | 6676 |  |  | 2008-08-06 03:05 | 6621 | 6621
2008-08-06 03:08 | 6673 | 6673 |  |  | 2008-08-06 03:10 | 6613 | 6613
2008-08-06 03:13 | 6624 |  |  |  | 2008-08-06 03:15 | 6470 |  | 6470
2008-08-06 03:18 | 6525 |  | 6525 |  | 2008-08-06 03:20 | 6466 |  | 6466
2008-08-06 03:23 | 6520 |  | 6520 |  | 2008-08-06 03:25 | 6462 |  | 6462
2008-08-06 03:28 | 6569 |  |  |  | 2008-08-06 03:30 | 6618 | 6618
2008-08-06 03:33 | 6675 | 6675 |  |  | 2008-08-06 03:35 | 6615 | 6615
2008-08-06 03:38 | 6676 | 6676 |  |  | 2008-08-06 03:40 | 6616 | 6616
2008-08-06 03:43 | 6633 |  |  |  | 2008-08-06 03:45 | 6472 |  | 6472
2008-08-06 03:48 | 6527 |  | 6527 |  | 2008-08-06 03:50 | 6481 |  | 6481
2008-08-06 03:53 | 6537 |  | 6537 |  | 2008-08-06 03:55 | 6475 |  | 6475
2008-08-06 03:58 | 6579 |  |  |  | 2008-08-06 04:00 | 6618 | 6618
2008-08-06 04:03 | 6663 | 6663 |  |  | 2008-08-06 04:05 | 6617 | 6617
2008-08-06 04:08 | 6677 | 6677 |  |  | 2008-08-06 04:10 | 6614 | 6614
2008-08-06 04:13 | 6625 |  |  |  | 2008-08-06 04:15 | 6458 |  | 6458
2008-08-06 04:18 | 6513 |  | 6513 |  | 2008-08-06 04:20 | 6460 |  | 6460
2008-08-06 04:23 | 6513 |  | 6513 |  | 2008-08-06 04:25 | 6451 |  | 6451
2008-08-06 04:28 | 6562 |  |  |  | 2008-08-06 04:30 | 6620 | 6620
2008-08-06 04:33 | 6670 | 6670 |  |  | 2008-08-06 04:35 | 6615 | 6615
2008-08-06 04:38 | 6667 | 6667 |  |  | 2008-08-06 04:40 | 6614 | 6614
2008-08-06 04:43 | 8305 |  |  |  | 2008-08-06 04:45 | 10875
2008-08-06 04:48 | 11000 | 11000 |  |  | 2008-08-06 04:50 | 10923 | 10923
2008-08-06 04:53 | 11023 | 11023 |  |  | 2008-08-06 04:55 | 10930 | 10930
2008-08-06 04:58 | 10927 |  |  |  | 2008-08-06 05:00 | 10679 |  | 10679
2008-08-06 05:03 | 10766 |  | 10766 |  | 2008-08-06 05:05 | 10687 |  | 10687
2008-08-06 05:08 | 10768 |  | 10768 |  | 2008-08-06 05:10 | 10664 |  | 10664
2008-08-06 05:13 | 10856 |  |  |  | 2008-08-06 05:15 | 10926 | 10926
2008-08-06 05:18 | 11019 | 11019 |  |  | 2008-08-06 05:20 | 10917 | 10917
2008-08-06 05:23 | 11020 | 11020 |  |  | 2008-08-06 05:25 | 10916 | 10916
2008-08-06 05:28 | 10937 |  |  |  | 2008-08-06 05:30 | 10678 |  | 10678
2008-08-06 05:33 | 10772 |  | 10772 |  | 2008-08-06 05:35 | 10683 |  | 10683
2008-08-06 05:38 | 10775 |  | 10775 |  | 2008-08-06 05:40 | 10686 |  | 10686
2008-08-06 05:43 | 10873 |  |  |  | 2008-08-06 05:45 | 10930 | 10930
2008-08-06 05:48 | 11021 | 11021 |  |  | 2008-08-06 05:50 | 10937 | 10937
2008-08-06 05:53 | 11030 | 11030 |  |  | 2008-08-06 05:55 | 10931 | 10931
2008-08-06 05:58 | 10931 |  |  |  | 2008-08-06 06:00 | 10673 |  | 10673
2008-08-06 06:03 | 10771 |  | 10771 |  | 2008-08-06 06:05 | 10667 |  | 10667
2008-08-06 06:08 | 10749 |  | 10749 |  | 2008-08-06 06:10 | 10676 |  | 10676
2008-08-06 06:13 | 10849 |  |  |  | 2008-08-06 06:15 | 10870 | 10870
2008-08-06 06:18 | 10958 | 10958 |  |  | 2008-08-06 06:20 | 10834 | 10834
2008-08-06 06:23 | 10921 | 10921 |  |  | 2008-08-06 06:25 | 10845 | 10845
2008-08-06 06:28 | 10854 |  |  |  | 2008-08-06 06:30 | 10579 |  | 10579
2008-08-06 06:33 | 10641 |  | 10641 |  | 2008-08-06 06:35 | 10543 |  | 10543
2008-08-06 06:38 | 10641 |  | 10641 |  | 2008-08-06 06:40 | 10550 |  | 10550
2008-08-06 06:43 | 10741 |  |  |  | 2008-08-06 06:45 | 10773 | 10773
2008-08-06 06:48 | 10863 | 10863 |  |  | 2008-08-06 06:50 | 10799 | 10799
2008-08-06 06:53 | 10894 | 10894 |  |  | 2008-08-06 06:55 | 10782 | 10782
2008-08-06 06:58 | 10766 |  |  |  | 2008-08-06 07:00 | 10499 |  | 10499
2008-08-06 07:03 | 10613 |  | 10613 |  | 2008-08-06 07:05 | 10548 |  | 10548
2008-08-06 07:08 | 10641 |  | 10641 |  | 2008-08-06 07:10 | 10527 |  | 10527
2008-08-06 07:13 | 10673 |  |  |  | 2008-08-06 07:15 | 10771 | 10771
2008-08-06 07:18 | 10845 | 10845 |  |  | 2008-08-06 07:20 | 10736 | 10736
2008-08-06 07:23 | 10825 | 10825 |  |  | 2008-08-06 07:25 | 10756 | 10756
2008-08-06 07:28 | 10778 |  |  |  | 2008-08-06 07:30 | 10525 |  | 10525
2008-08-06 07:33 | 10600 |  | 10600 |  | 2008-08-06 07:35 | 10483 |  | 10483
2008-08-06 07:38 | 10574 |  | 10574 |  | 2008-08-06 07:40 | 10502 |  | 10502
2008-08-06 07:43 | 10671 |  |  |  | 2008-08-06 07:45 | 10720 | 10720
2008-08-06 07:48 | 10823 | 10823 |  |  | 2008-08-06 07:50 | 10750 | 10750
2008-08-06 07:53 | 10836 | 10836 |  |  | 2008-08-06 07:55 | 10740 | 10740
2008-08-06 07:58 | 6765 |  |  |  | 2008-08-06 08:00 | 47.809
2008-08-06 08:03 | 24.807 |  |  |  | 2008-08-06 08:05 | 17.259
2008-08-06 08:08 | 11.931 |  |  |  | 2008-08-06 08:10 | 10.939
2008-08-06 08:13 | 10.087 |  |  |  | 2008-08-06 08:15 | 7.305
2008-08-06 08:18 | 7.14 |  |  |  | 2008-08-06 08:20 | 7.226
2008-08-06 08:23 | 5.487 |  |  |  | 2008-08-06 08:25 | 5.439
2008-08-06 08:28 | 0 |  |  |  | 2008-08-06 08:30 | 0
2008-08-06 08:33 | 0 |  |  |  | 2008-08-06 08:35 | 4.616
2008-08-06 08:38 | 0 |  |  |  | 2008-08-06 08:40 | 4.858
2008-08-06 08:43 | 7802 |  |  |  | 2008-08-06 08:45 | 20633
2008-08-06 08:48 | 20897 |  |  |  | 2008-08-06 08:50 | 20782
2008-08-06 08:53 | 21000 | 21000 |  |  | 2008-08-06 08:55 | 20837 | 20837
2008-08-06 08:58 | 20844 |  |  |  | 2008-08-06 09:00 | 20433 |  | 20433
2008-08-06 09:03 | 20664 |  | 20664 |  | 2008-08-06 09:05 | 20472 |  | 20472
2008-08-06 09:08 | 20655 |  | 20655 |  | 2008-08-06 09:10 | 20455 |  | 20455
2008-08-06 09:13 | 20793 |  |  |  | 2008-08-06 09:15 | 20903 | 20903
2008-08-06 09:18 | 21074 | 21074 |  |  | 2008-08-06 09:20 | 20834 | 20834
2008-08-06 09:23 | 21037 | 21037 |  |  | 2008-08-06 09:25 | 20890 | 20890
2008-08-06 09:28 | 20919 |  |  |  | 2008-08-06 09:30 | 20466 |  | 20466
2008-08-06 09:33 | 20626 |  | 20626 |  | 2008-08-06 09:35 | 20451 |  | 20451
2008-08-06 09:38 | 20631 |  | 20631 |  | 2008-08-06 09:40 | 20421 |  | 20421
2008-08-06 09:43 | 20752 |  |  |  | 2008-08-06 09:45 | 20868 | 20868
2008-08-06 09:48 | 21079 | 21079 |  |  | 2008-08-06 09:50 | 20883 | 20883
2008-08-06 09:53 | 21058 | 21058 |  |  | 2008-08-06 09:55 | 20851 | 20851
2008-08-06 09:58 | 20928 |  |  |  | 2008-08-06 10:00 | 20438 |  | 20438
2008-08-06 10:03 | 20592 |  | 20592 |  | 2008-08-06 10:05 | 20416 |  | 20416
2008-08-06 10:08 | 20643 |  | 20643 |  | 2008-08-06 10:10 | 20476 |  | 20476
2008-08-06 10:13 | 20787 |  |  |  | 2008-08-06 10:15 | 20879 | 20879
2008-08-06 10:18 | 21102 | 21102 |  |  | 2008-08-06 10:20 | 20875 | 20875
2008-08-06 10:23 | 21054 | 21054 |  |  | 2008-08-06 10:25 | 20830 | 20830
2008-08-06 10:28 | 20934 |  |  |  | 2008-08-06 10:30 | 20452 |  | 20452
2008-08-06 10:33 | 20648 |  | 20648 |  | 2008-08-06 10:35 | 20456 |  | 20456
2008-08-06 10:38 | 20650 |  | 20650 |  | 2008-08-06 10:40 | 20468 |  | 20468
2008-08-06 10:43 | 20771 |  |  |  | 2008-08-06 10:45 | 20827 | 20827
2008-08-06 10:48 | 21056 | 21056 |  |  | 2008-08-06 10:50 | 20876 | 20876
2008-08-06 10:53 | 21050 | 21050 |  |  | 2008-08-06 10:55 | 20846 | 20846
2008-08-06 10:58 | 20888 |  |  |  | 2008-08-06 11:00 | 20422 |  | 20422
2008-08-06 11:03 | 20633 |  | 20633 |  | 2008-08-06 11:05 | 20447 |  | 20447
2008-08-06 11:08 | 20587 |  | 20587 |  | 2008-08-06 11:10 | 20402 |  | 20402
2008-08-06 11:13 | 20761 |  |  |  | 2008-08-06 11:15 | 20873 | 20873
2008-08-06 11:18 | 21001 | 21001 |  |  | 2008-08-06 11:20 | 20828 | 20828
2008-08-06 11:23 | 21050 | 21050 |  |  | 2008-08-06 11:25 | 20874 | 20874
2008-08-06 11:28 | 20875 |  |  |  | 2008-08-06 11:30 | 20369 |  | 20369
2008-08-06 11:33 | 20564 |  | 20564 |  | 2008-08-06 11:35 | 20426 |  | 20426
2008-08-06 11:38 | 20630 |  | 20630 |  | 2008-08-06 11:40 | 20423 |  | 20423
2008-08-06 11:43 | 20760 |  |  |  | 2008-08-06 11:45 | 20836 | 20836
2008-08-06 11:48 | 20988 | 20988 |  |  | 2008-08-06 11:50 | 20837 | 20837
2008-08-06 11:53 | 21034 | 21034 |  |  | 2008-08-06 11:55 | 20815 | 20815
2008-08-06 11:58 | 14010 |  |  |  | 2008-08-06 12:00 | 104.473
2008-08-06 12:03 | 58.645 |  |  |  | 2008-08-06 12:05 | 41.711
2008-08-06 12:08 | 34.706 |  |  |  | 2008-08-06 12:10 | 27.859
2008-08-06 12:13 | 28.011 |  |  |  | 2008-08-06 12:15 | 23.074
2008-08-06 12:18 | 22.185 |  |  |  | 2008-08-06 12:20 | 18.126
2008-08-06 12:23 | 17.304 |  |  |  | 2008-08-06 12:25 | 15.18
2008-08-06 12:28 | 15.24 |  |  |  | 2008-08-06 12:30 | 13.777
2008-08-06 12:33 | 14.463 |  |  |  | 2008-08-06 12:35 | 11.985
2008-08-06 12:38 | 12.845 |  |  |  | 2008-08-06 12:40 | 8.944
2008-08-06 12:43 | 15297 |  |  |  | 2008-08-06 12:45 | 39624
2008-08-06 12:48 | 40220 | 40220 |  |  | 2008-08-06 12:50 | 39801 | 39801
2008-08-06 12:53 | 40178 | 40178 |  |  | 2008-08-06 12:55 | 39840 | 39840
2008-08-06 12:58 | 40058 |  |  |  | 2008-08-06 13:00 | 39249 |  | 39249
2008-08-06 13:03 | 39588 |  | 39588 |  | 2008-08-06 13:05 | 39189 |  | 39189
2008-08-06 13:08 | 39611 |  | 39611 |  | 2008-08-06 13:10 | 39179 |  | 39179
2008-08-06 13:13 | 39815 |  |  |  | 2008-08-06 13:15 | 39836 | 39836
2008-08-06 13:18 | 40310 | 40310 |  |  | 2008-08-06 13:20 | 39859 | 39859
2008-08-06 13:23 | 40230 | 40230 |  |  | 2008-08-06 13:25 | 39773 | 39773
2008-08-06 13:28 | 39992 |  |  |  | 2008-08-06 13:30 | 39217 |  | 39217
2008-08-06 13:33 | 39626 |  | 39626 |  | 2008-08-06 13:35 | 39193 |  | 39193
2008-08-06 13:38 | 39536 |  | 39536 |  | 2008-08-06 13:40 | 39164 |  | 39164
2008-08-06 13:43 | 39745 |  |  |  | 2008-08-06 13:45 | 39801 | 39801
2008-08-06 13:48 | 40153 | 40153 |  |  | 2008-08-06 13:50 | 39763 | 39763
2008-08-06 13:53 | 40235 | 40235 |  |  | 2008-08-06 13:55 | 39897 | 39897
2008-08-06 13:58 | 39972 |  |  |  | 2008-08-06 14:00 | 39086 |  | 39086
2008-08-06 14:03 | 39502 |  | 39502 |  | 2008-08-06 14:05 | 39153 |  | 39153
2008-08-06 14:08 | 39502 |  | 39502 |  | 2008-08-06 14:10 | 39181 |  | 39181
2008-08-06 14:13 | 39782 |  |  |  | 2008-08-06 14:15 | 39735 | 39735
2008-08-06 14:18 | 40198 | 40198 |  |  | 2008-08-06 14:20 | 39850 | 39850
2008-08-06 14:23 | 40245 | 40245 |  |  | 2008-08-06 14:25 | 39727 | 39727
2008-08-06 14:28 | 39988 |  |  |  | 2008-08-06 14:30 | 39232 |  | 39232
2008-08-06 14:33 | 39567 |  | 39567 |  | 2008-08-06 14:35 | 39224 |  | 39224
2008-08-06 14:38 | 39507 |  | 39507 |  | 2008-08-06 14:40 | 39224 |  | 39224
2008-08-06 14:43 | 39856 |  |  |  | 2008-08-06 14:45 | 39835 | 39835
2008-08-06 14:48 | 40144 | 40144 |  |  | 2008-08-06 14:50 | 39882 | 39882
2008-08-06 14:53 | 40207 | 40207 |  |  | 2008-08-06 14:55 | 39821 | 39821
2008-08-06 14:58 | 39932 |  |  |  | 2008-08-06 15:00 | 39260 |  | 39260
2008-08-06 15:03 | 39601 |  | 39601 |  | 2008-08-06 15:05 | 39231 |  | 39231
2008-08-06 15:08 | 39586 |  | 39586 |  | 2008-08-06 15:10 | 39169 |  | 39169
2008-08-06 15:13 | 39694 |  |  |  | 2008-08-06 15:15 | 39893 | 39893
2008-08-06 15:18 | 40220 | 40220 |  |  | 2008-08-06 15:20 | 39776 | 39776
2008-08-06 15:23 | 40108 | 40108 |  |  | 2008-08-06 15:25 | 39778 | 39778
2008-08-06 15:28 | 39941 |  |  |  | 2008-08-06 15:30 | 39095 |  | 39095
2008-08-06 15:33 | 39437 |  | 39437 |  | 2008-08-06 15:35 | 39117 |  | 39117
2008-08-06 15:38 | 39460 |  | 39460 |  | 2008-08-06 15:40 | 39157 |  | 39157
2008-08-06 15:43 | 39644 |  |  |  | 2008-08-06 15:45 | 39695 | 39695
2008-08-06 15:48 | 40142 | 40142 |  |  | 2008-08-06 15:50 | 39750 | 39750
2008-08-06 15:53 | 40006 |  |  |  | 2008-08-06 15:55 | 39668 | 39668
2008-08-06 15:58 | 26078 |  |  |  | 2008-08-06 16:00 | 199.358
2008-08-06 16:03 | 112.236 |  |  |  | 2008-08-06 16:05 | 79.735
2008-08-06 16:08 | 63.268 |  |  |  | 2008-08-06 16:10 | 51.911
2008-08-06 16:13 | 44.362 |  |  |  | 2008-08-06 16:15 | 35.276
2008-08-06 16:18 | 35.262 |  |  |  | 2008-08-06 16:20 | 26.409
2008-08-06 16:23 | 29.22 |  |  |  | 2008-08-06 16:25 | 25.154
2008-08-06 16:28 | 25.375 |  |  |  | 2008-08-06 16:30 | 23.714
2008-08-06 16:33 | 23.458 |  |  |  | 2008-08-06 16:35 | 18.521
2008-08-06 16:38 | 18.375 |  |  |  | 2008-08-06 16:40 | 17.658
2008-08-06 16:43 | 16.275 |  |  |  | 2008-08-06 16:45 | 15.738
2008-08-06 16:48 | 17.004 |  |  |  | 2008-08-06 16:50 | 15.856
2008-08-06 16:53 | 14.859 |  |  |  | 2008-08-06 16:55 | 12.23
2008-08-06 16:58 | 11.369 |  |  |  | 2008-08-06 17:00 | 12.008
2008-08-06 17:03 | 6.615 |  |  | 6.615 | 2008-08-06 17:05 | 10.79 |  |  | 10.79
2008-08-06 17:08 | 11.53 |  |  | 11.53 | 2008-08-06 17:10 | 11.192 |  |  | 11.192
2008-08-06 17:13 | 10.514 |  |  | 10.514 | 2008-08-06 17:15 | 9.624 |  |  | 9.624
2008-08-06 17:18 | 10.177 |  |  | 10.177 | 2008-08-06 17:20 | 6.602 |  |  | 6.602
2008-08-06 17:23 | 7.633 |  |  | 7.633 | 2008-08-06 17:25 | 7.258 |  |  | 7.258
2008-08-06 17:28 | 77.088 |  |  |  | 2008-08-06 17:30 | 46.858
2008-08-06 17:33 | 8.855 |  |  |  | 2008-08-06 17:35 | 11.548
2008-08-06 17:38 | 10.937 |  |  |  | 2008-08-06 17:40 | 10.47
2008-08-06 17:43 | 9.204 |  |  |  | 2008-08-06 17:45 | 1130
2008-08-06 17:48 | 1706 |  |  |  | 2008-08-06 17:50 | 1691
2008-08-06 17:53 | 1713 |  |  |  | 2008-08-06 17:55 | 1687
2008-08-06 17:58 | 1716 |  |  |  | 2008-08-06 18:00 | 1700
2008-08-06 18:03 | 1719 |  |  |  | 2008-08-06 18:05 | 1696
2008-08-06 18:08 | 1717 |  |  |  | 2008-08-06 18:10 | 1700
2008-08-06 18:13 | 1718 |  |  |  | 2008-08-06 18:15 | 1702
2008-08-06 18:18 | 1716 |  |  |  | 2008-08-06 18:20 | 1696
2008-08-06 18:23 | 1715 |  |  |  | 2008-08-06 18:25 | 1698
2008-08-06 18:28 | 1725 |  |  |  | 2008-08-06 18:30 | 1700
2008-08-06 18:33 | 1720 |  |  |  | 2008-08-06 18:35 | 1706
2008-08-06 18:38 | 1718 |  |  |  | 2008-08-06 18:40 | 1703
2008-08-06 18:43 | 1721 |  |  |  | 2008-08-06 18:45 | 1703
2008-08-06 18:48 | 1724 |  |  |  | 2008-08-06 18:50 | 1705
2008-08-06 18:53 | 1721 |  |  |  | 2008-08-06 18:55 | 1702
2008-08-06 18:58 | 1724 |  |  |  | 2008-08-06 19:00 | 1711
2008-08-06 19:03 | 1726 |  |  |  | 2008-08-06 19:05 | 1702
2008-08-06 19:08 | 1720 |  |  |  | 2008-08-06 19:10 | 1703
2008-08-06 19:13 | 1720 |  |  |  | 2008-08-06 19:15 | 1696
2008-08-06 19:18 | 1722 |  |  |  | 2008-08-06 19:20 | 1704
2008-08-06 19:23 | 1725 |  |  |  | 2008-08-06 19:25 | 1705
2008-08-06 19:28 | 1725 |  |  |  | 2008-08-06 19:30 | 1702
2008-08-06 19:33 | 1721 |  |  |  | 2008-08-06 19:35 | 1704
2008-08-06 19:38 | 1722 |  |  |  | 2008-08-06 19:40 | 1706
2008-08-06 19:43 | 1724 |  |  |  | 2008-08-06 19:45 | 1708
2008-08-06 19:48 | 1723 |  |  |  | 2008-08-06 19:50 | 1700
2008-08-06 19:53 | 1725 |  |  |  | 2008-08-06 19:55 | 1708
2008-08-06 19:58 | 1730 |  |  |  | 2008-08-06 20:00 | 1715
2008-08-06 20:03 | 1725 |  |  |  | 2008-08-06 20:05 | 1707
2008-08-06 20:08 | 1727 |  |  |  | 2008-08-06 20:10 | 1707
2008-08-06 20:13 | 1729 |  |  |  | 2008-08-06 20:15 | 1711
2008-08-06 20:18 | 1729 |  |  |  | 2008-08-06 20:20 | 1711
2008-08-06 20:23 | 1731 |  |  |  | 2008-08-06 20:25 | 1711
2008-08-06 20:28 | 1725 |  |  |  | 2008-08-06 20:30 | 1714
2008-08-06 20:33 | 1729 |  |  |  | 2008-08-06 20:35 | 1712
2008-08-06 20:38 | 1732 |  |  |  | 2008-08-06 20:40 | 1711
2008-08-06 20:43 | 1726 |  |  |  | 2008-08-06 20:45 | 1711
2008-08-06 20:48 | 1727 |  |  |  | 2008-08-06 20:50 | 1709
2008-08-06 20:53 | 1729 |  |  |  | 2008-08-06 20:55 | 1710
2008-08-06 20:58 | 1727 |  |  |  | 2008-08-06 21:00 | 271.459
2008-08-06 21:03 | 10.209 |  |  |  | 2008-08-06 21:05 | 6.782
2008-08-06 21:08 | 0 |  |  |  | 2008-08-06 21:10 | 5.471
2008-08-06 21:13 | 0 |  |  |  | 2008-08-06 21:15 | 3.694
2008-08-06 21:18 | 0 |  |  |  | 2008-08-06 21:20 | 4.325
2008-08-06 21:23 | 2.991 |  |  |  | 2008-08-06 21:25 | 0
2008-08-06 21:28 | 0 |  |  |  | 2008-08-06 21:30 | 3.713
2008-08-06 21:33 | 0 |  |  |  | 2008-08-06 21:35 | 0
2008-08-06 21:38 | 3.956 |  |  |  | 2008-08-06 21:40 | 3.236
2008-08-06 21:43 | 2.615 |  |  |  | 2008-08-06 21:45 | 2468
[10 empty] | 653 | 152 | 190 | 653 | 225 | 271
Formulas:
Date & Time = IF([Date & Time, row above]<Test_Info!$G9,IF(MIN(IF(Detector_AB="A",IF(Detector_Time>[Date & Time, row above],Detector_Time,""),""))=0,NA(),MIN(IF(Detector_AB="A",IF(Detector_Time>[Date & Time, row above],Detector_Time,""),""))),NA())
Date & Time = IF([Date & Time, row above]<Test_Info!$G9,IF(MIN(IF(Detector_AB="B",IF(Detector_Time>[Date & Time, row above],Detector_Time,""),""))=0,NA(),MIN(IF(Detector_AB="B",IF(Detector_Time>[Date & Time, row above],Detector_Time,""),""))),NA())
Recalculate

## Point-1
Single-Point Bracketing Data
Reference Setpoint |  |  | 0.5 | mg/m3 |  | Bracketing Procedure Start |  | 2008-08-04 15:25
Reference Concentration |  |  | 0.53524 | mg/m3
[6 empty] | Bracketing Procedure End |  | 2008-08-05 18:10
Candidate Setpoint |  |  | 0.5 | mg/m3
Date/Time |  | Meas. Type | "A-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "A-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 15:25 |  | Zero | 0 | [7 empty] | 2008-08-04 15:25 | 2008-08-04 15:25
2008-08-04 16:15 |  | Ref | 586.3895 | 585.8311 | 2 | 1.9454 |  |  |  |  | 2008-08-04 16:13 | 2008-08-04 16:18
2008-08-04 16:30 |  | Cand | 547.5635 | 546.8376 | 2 | 1.8617 | 586.5771 | 0.93225 | 0.0038539 |  | 2008-08-04 16:28 | 2008-08-04 16:33
2008-08-04 16:45 |  | Ref | 588.2165 | 587.323 | 2 | 1.9454 |  |  |  |  | 2008-08-04 16:43 | 2008-08-04 16:48
2008-08-04 17:00 |  | Cand | 553.6015 | 552.5405 | 2 | 1.8617 | 588.678 | 0.93861 | 0.0038486 |  | 2008-08-04 16:58 | 2008-08-04 17:03
2008-08-04 17:15 |  | Ref | 591.2615 | 590.033 | 2 | 1.9454 |  |  |  |  | 2008-08-04 17:13 | 2008-08-04 17:18
2008-08-04 17:30 |  | Cand | 549.16 | 547.7639 | 2 | 1.8617 | 589.3414 | 0.92945 | 0.0038321 |  | 2008-08-04 17:28 | 2008-08-04 17:33
2008-08-04 17:45 |  | Ref | 590.2135 | 588.6499 | 2 | 1.9454 |  |  |  |  | 2008-08-04 17:43 | 2008-08-04 17:48
2008-08-04 18:00 |  | Cand | 552.4805 | 550.7494 | 2 | 1.8617 | 587.0691 | 0.93813 | 0.0038585 |  | 2008-08-04 17:58 | 2008-08-04 18:03
2008-08-04 18:15 |  | Ref | 587.387 | 585.4884 | 2 | 1.9454 |  |  |  |  | 2008-08-04 18:13 | 2008-08-04 18:18
2008-08-04 18:30 |  | Cand | 549.6495 | 547.5833 | 2 | 1.8617 | 588.4246 | 0.93059 | 0.0038397 |  | 2008-08-04 18:28 | 2008-08-04 18:33
2008-08-04 18:45 |  | Ref | 593.5945 | 591.3608 | 2 | 1.9454 |  |  |  |  | 2008-08-04 18:43 | 2008-08-04 18:48
2008-08-04 19:00 |  | Cand | 552.64 | 550.2388 | 2 | 1.8617 | 592.6405 | 0.92845 | 0.0038095 |  | 2008-08-04 18:58 | 2008-08-04 19:03
2008-08-04 19:15 |  | Ref | 596.489 | 593.9203 | 2 | 1.9454 |  |  |  |  | 2008-08-04 19:13 | 2008-08-04 19:18
2008-08-05 18:05 |  | Zero | 17.8695 | [7 empty] | 2008-08-05 18:03 | 2008-08-05 18:08
Bracketing Intervals
Detector Linearity | Linearity Factor at 0.5 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 2.6328 | 2008-08-04 16:28 | 2008-08-04 19:03
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 2.7511 | 2008-08-04 16:13 | 2008-08-04 19:18
| Degrees of Freedom |  |  | 5
Date/Time |  | Meas. Type | "B-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "B-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 16:03 |  | Zero | 0 | [7 empty] | 2008-08-04 16:00 | 2008-08-04 16:05
2008-08-04 16:18 |  | Ref | 577.6885 | 577.5324 | 2 | 1.7313 |  |  |  |  | 2008-08-04 16:15 | 2008-08-04 16:20
2008-08-04 16:30 |  | Cand | 541.056 | 540.7698 | 3 | 1.8098 | 579.5043 | 0.93316 | 0.0036253 |  | 2008-08-04 16:25 | 2008-08-04 16:35
2008-08-04 16:45 |  | Ref | 582.313 | 581.8706 | 3 | 1.4136 |  |  |  |  | 2008-08-04 16:40 | 2008-08-04 16:50
2008-08-04 17:00 |  | Cand | 546.3473 | 545.7488 | 3 | 1.8098 | 581.0752 | 0.93921 | 0.0035086 |  | 2008-08-04 16:55 | 2008-08-04 17:05
2008-08-04 17:15 |  | Ref | 581.0343 | 580.2797 | 3 | 1.4136 |  |  |  |  | 2008-08-04 17:10 | 2008-08-04 17:20
2008-08-04 17:30 |  | Cand | 541.6417 | 540.7309 | 3 | 1.8098 | 581.5577 | 0.9298 | 0.0034983 |  | 2008-08-04 17:25 | 2008-08-04 17:35
2008-08-04 17:45 |  | Ref | 583.9027 | 582.8358 | 3 | 1.4136 |  |  |  |  | 2008-08-04 17:40 | 2008-08-04 17:50
2008-08-04 18:00 |  | Cand | 543.5223 | 542.2993 | 3 | 1.8098 | 581.2948 | 0.93292 | 0.0035023 |  | 2008-08-04 17:55 | 2008-08-04 18:05
2008-08-04 18:15 |  | Ref | 581.133 | 579.7539 | 3 | 1.4136 |  |  |  |  | 2008-08-04 18:10 | 2008-08-04 18:20
2008-08-04 18:30 |  | Cand | 542.1033 | 540.5681 | 3 | 1.8098 | 580.4291 | 0.93132 | 0.0035063 |  | 2008-08-04 18:25 | 2008-08-04 18:35
2008-08-04 18:45 |  | Ref | 582.7957 | 581.1043 | 3 | 1.4136 |  |  |  |  | 2008-08-04 18:40 | 2008-08-04 18:50
2008-08-04 19:00 |  | Cand | 547.2407 | 545.3931 | 3 | 1.8098 |  | 0 |  |  | 2008-08-04 18:55 | 2008-08-04 19:05
|  | Ref |  |  | 0
2008-08-05 18:08 |  | Zero | 16.2895 | [7 empty] | 2008-08-05 18:05 | 2008-08-05 18:10
Bracketing Intervals
Detector Linearity | Linearity Factor at 0.5 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 3.1346 | 2008-08-04 16:25 | 2008-08-04 19:05
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 2.4485 | 2008-08-04 16:15 | 2008-08-04 19:05
| Degrees of Freedom |  |  | 4
| Ratio Average Uncertainty Components, 1s |  |  | Pooled Detector Linearity |  |  |  | 0
|  |  |  | Measurement Stability |  |  |  | 0.0011162
|  |  |  | Ratio Repeatability |  |  |  | 0.0010878
|  | Set Average Candidate/Reference Ratio |  |  |  |  | 0.93308
|  | Number of Brackets in this Set |  |  |  |  | 11
|  | Uncertainty for This Set's Average, 1s |  |  |  |  | 0.0015586
|  | Nested "Level-2" StdErr, s2 |  |  |  |  | 0.0038602
Formulas:
First Record = MIN(IF(A_Ref="","",IF(A_Time>[First Record, row above],A_Time,"")))
First Record = MIN(IF(A_Cand="","",IF(A_Time>[First Record, row above],A_Time,"")))
Last Record = MIN(IF(A_Zero="",IF(A_Time>[First Record],A_Time1,""),""))
Last Record = MIN(IF(A_Ref="",IF(A_Time>[First Record],A_Time1,""),""))
Note on A8: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on A31: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on D3: Ed Brown: The "nominal" reference setpoint, used for identification, not for calculation
Note on D4: Ed Brown: Actual certified reference concentration
Note on D6: Ed Brown: The "nominal" candidate setpoint, used for identification, not for calculation
Note on D9: Ed Brown: Average response for measurement interval
Note on D32: Ed Brown: Average response for measurement interval
Note on E9: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E27: Ed Brown: Value from Linearity Chack reponse table
Note on E29: Ed Brown: Degrees of Freedom used for pooling below
Note on E32: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E50: Ed Brown: Value from Linearity Chack reponse table
Note on E52: Ed Brown: Degrees of Freedom used for pooling below
Note on F10: Ed Brown: Number of individual readings in measurement interval
Note on F33: Ed Brown: Number of individual readings in measurement interval
Note on G10: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on G33: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on H9: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H32: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H60: Ed Brown: The "End Result" of this bracketing procedure
Note on H62: Ed Brown: Uncertainty of the ratio result of this bracketing procedure
Note on I9: Ed Brown: Calculated measurement ratio for bracket
Note on I32: Ed Brown: Calculated measurement ratio for bracket
Note on J9: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on J32: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on K27: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K28: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on K50: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K51: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on M26: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Note on M49: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Populate Parse Table

## Point-2
Single-Point Bracketing Data
Reference Setpoint |  |  | 1.5 | mg/m3 |  | Bracketing Procedure Start |  | 2008-08-04 15:25
Reference Concentration |  |  | 1.6557 | mg/m3
[6 empty] | Bracketing Procedure End |  | 2008-08-05 18:10
Candidate Setpoint |  |  | 1.5 | mg/m3
Date/Time |  | Meas. Type | "A-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "A-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 16:00 |  | Zero | 0 | [7 empty] | 2008-08-04 15:58 | 2008-08-04 16:03
2008-08-04 19:30 |  | Ref | 1808 | 1805.6022 | 2 | 1.9305 |  |  |  |  | 2008-08-04 19:28 | 2008-08-04 19:33
2008-08-04 19:45 |  | Cand | 1739 | 1736.4309 | 2 | 1.6998 | 1808.1809 | 0.96032 | 0.0011871 |  | 2008-08-04 19:43 | 2008-08-04 19:48
2008-08-04 20:00 |  | Ref | 1813.5 | 1810.7596 | 2 | 1.9305 |  |  |  |  | 2008-08-04 19:58 | 2008-08-04 20:03
2008-08-04 20:15 |  | Cand | 1741.5 | 1738.5884 | 2 | 1.6998 | 1811.8384 | 0.95957 | 0.0011844 |  | 2008-08-04 20:13 | 2008-08-04 20:18
2008-08-04 20:30 |  | Ref | 1816 | 1812.9171 | 2 | 1.9305 |  |  |  |  | 2008-08-04 20:28 | 2008-08-04 20:33
2008-08-04 20:45 |  | Cand | 1741 | 1737.7458 | 2 | 1.6998 | 1813.4958 | 0.95823 | 0.0011827 |  | 2008-08-04 20:43 | 2008-08-04 20:48
2008-08-04 21:00 |  | Ref | 1817.5 | 1814.0745 | 2 | 1.9305 |  |  |  |  | 2008-08-04 20:58 | 2008-08-04 21:03
2008-08-04 21:15 |  | Cand | 1740 | 1736.4033 | 2 | 1.6998 | 1813.1533 | 0.95767 | 0.0011827 |  | 2008-08-04 21:13 | 2008-08-04 21:18
2008-08-04 21:30 |  | Ref | 1816 | 1812.232 | 2 | 1.9305 |  |  |  |  | 2008-08-04 21:28 | 2008-08-04 21:33
2008-08-04 21:45 |  | Cand | 1743.5 | 1739.5607 | 2 | 1.6998 | 1812.5607 | 0.95973 | 0.001184 |  | 2008-08-04 21:43 | 2008-08-04 21:48
2008-08-04 22:00 |  | Ref | 1817 | 1812.8894 | 2 | 1.9305 |  |  |  |  | 2008-08-04 21:58 | 2008-08-04 22:03
2008-08-04 22:15 |  | Cand | 1740.5 | 1736.2182 | 2 | 1.6998 | 1813.2182 | 0.95753 | 0.0011826 |  | 2008-08-04 22:13 | 2008-08-04 22:18
2008-08-04 22:30 |  | Ref | 1818 | 1813.5469 | 2 | 1.9305 |  |  |  |  | 2008-08-04 22:28 | 2008-08-04 22:33
2008-08-05 18:05 |  | Zero | 17.8695 | [7 empty] | 2008-08-05 18:03 | 2008-08-05 18:08
Bracketing Intervals
Detector Linearity | Linearity Factor at 1.5 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 2.4039 | 2008-08-04 19:43 | 2008-08-04 22:18
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 2.7301 | 2008-08-04 19:28 | 2008-08-04 22:33
| Degrees of Freedom |  |  | 5
Date/Time |  | Meas. Type | "B-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "B-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 16:03 |  | Zero | 0 | [7 empty] | 2008-08-04 16:00 | 2008-08-04 16:05
2008-08-04 19:30 |  | Ref | 1785.6667 | 1783.5069 | 3 | 1.9029 |  |  |  |  | 2008-08-04 19:25 | 2008-08-04 19:35
2008-08-04 19:45 |  | Cand | 1721.3333 | 1719.0174 | 3 | 1.0851 | 1787.6841 | 0.96159 | 0.00094458 |  | 2008-08-04 19:40 | 2008-08-04 19:50
2008-08-04 20:00 |  | Ref | 1794.3333 | 1791.8613 | 3 | 1.9029 |  |  |  |  | 2008-08-04 19:55 | 2008-08-04 20:05
2008-08-04 20:15 |  | Cand | 1723.6667 | 1721.0385 | 3 | 1.0851 | 1793.8718 | 0.9594 | 0.00094006 |  | 2008-08-04 20:10 | 2008-08-04 20:20
2008-08-04 20:30 |  | Ref | 1798.6667 | 1795.8824 | 3 | 1.9029 |  |  |  |  | 2008-08-04 20:25 | 2008-08-04 20:35
2008-08-04 20:45 |  | Cand | 1723.6667 | 1720.7262 | 3 | 1.0851 | 1794.8929 | 0.95868 | 0.00093911 |  | 2008-08-04 20:40 | 2008-08-04 20:50
2008-08-04 21:00 |  | Ref | 1797 | 1793.9034 | 3 | 1.9029 |  |  |  |  | 2008-08-04 20:55 | 2008-08-04 21:05
2008-08-04 21:15 |  | Cand | 1723.3333 | 1720.0806 | 3 | 1.0851 | 1794.5806 | 0.95849 | 0.00093916 |  | 2008-08-04 21:10 | 2008-08-04 21:20
2008-08-04 21:30 |  | Ref | 1798.6667 | 1795.2578 | 3 | 1.9029 |  |  |  |  | 2008-08-04 21:25 | 2008-08-04 21:35
2008-08-04 21:45 |  | Cand | 1724 | 1720.435 | 3 | 1.0851 | 1795.7684 | 0.95805 | 0.00093829 |  | 2008-08-04 21:40 | 2008-08-04 21:50
2008-08-04 22:00 |  | Ref | 1800 | 1796.2789 | 3 | 1.9029 |  |  |  |  | 2008-08-04 21:55 | 2008-08-04 22:05
2008-08-04 22:15 |  | Cand | 1722.3333 | 1718.4561 | 3 | 1.0851 | 1795.6228 | 0.95703 | 0.00093778 |  | 2008-08-04 22:10 | 2008-08-04 22:20
2008-08-04 22:30 |  | Ref | 1799 | 1794.9667 | 3 | 1.9029 |  |  |  |  | 2008-08-04 22:25 | 2008-08-04 22:35
2008-08-05 18:08 |  | Zero | 16.2895 | [7 empty] | 2008-08-05 18:05 | 2008-08-05 18:10
Bracketing Intervals
Detector Linearity | Linearity Factor at 1.5 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 1.8794 | 2008-08-04 19:40 | 2008-08-04 22:20
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 3.2958 | 2008-08-04 19:25 | 2008-08-04 22:35
| Degrees of Freedom |  |  | 5
| Ratio Average Uncertainty Components, 1s |  |  | Pooled Detector Linearity |  |  |  | 0
|  |  |  | Measurement Stability |  |  |  | 0.00030855
|  |  |  | Ratio Repeatability |  |  |  | 0.00036111
|  | Set Average Candidate/Reference Ratio |  |  |  |  | 0.95886
|  | Number of Brackets in this Set |  |  |  |  | 12
|  | Uncertainty for This Set's Average, 1s |  |  |  |  | 0.00047498
|  | Nested "Level-2" StdErr, s2 |  |  |  |  | 0.0013104
Formulas:
First Record = MIN(IF(A_Cand="","",IF(A_Time>[First Record, row above],A_Time,"")))
First Record = MIN(IF(A_Ref="","",IF(A_Time>[First Record, row above],A_Time,"")))
Last Record = MIN(IF(A_Zero="",IF(A_Time>[First Record],A_Time1,""),""))
Last Record = MIN(IF(A_Ref="",IF(A_Time>[First Record],A_Time1,""),""))
Note on A8: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on A31: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on D3: Ed Brown: The "nominal" reference setpoint, used for identification, not for calculation
Note on D4: Ed Brown: Actual certified reference concentration
Note on D6: Ed Brown: The "nominal" candidate setpoint, used for identification, not for calculation
Note on D9: Ed Brown: Average response for measurement interval
Note on D32: Ed Brown: Average response for measurement interval
Note on E9: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E27: Ed Brown: Value from Linearity Chack reponse table
Note on E29: Ed Brown: Degrees of Freedom used for pooling below
Note on E32: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E50: Ed Brown: Value from Linearity Chack reponse table
Note on E52: Ed Brown: Degrees of Freedom used for pooling below
Note on F10: Ed Brown: Number of individual readings in measurement interval
Note on F33: Ed Brown: Number of individual readings in measurement interval
Note on G10: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on G33: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on H9: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H32: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H60: Ed Brown: The "End Result" of this bracketing procedure
Note on H62: Ed Brown: Uncertainty of the ratio result of this bracketing procedure
Note on I9: Ed Brown: Calculated measurement ratio for bracket
Note on I32: Ed Brown: Calculated measurement ratio for bracket
Note on J9: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on J32: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on K27: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K28: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on K50: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K51: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on M26: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Note on M49: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Populate Parse Table

## Point-3
Single-Point Bracketing Data
Reference Setpoint |  |  | 2.7 | mg/m3 |  | Bracketing Procedure Start |  | 2008-08-04 15:25
Reference Concentration |  |  | 2.8581 | mg/m3
[6 empty] | Bracketing Procedure End |  | 2008-08-05 18:10
Candidate Setpoint |  |  | 2.7 | mg/m3
Date/Time |  | Meas. Type | "A-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "A-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 16:00 |  | Zero | 0 | [7 empty] | 2008-08-04 15:58 | 2008-08-04 16:03
2008-08-04 23:28 |  | Ref | 3159 | 3153.8904 | 1 | 3.9259 |  |  |  |  | 2008-08-04 23:28 | 2008-08-04 23:28
2008-08-04 23:40 |  | Cand | 3070.5 | 3065.2476 | 2 | 4.0352 | 3161.2476 | 0.96963 | 0.0014872 |  | 2008-08-04 23:38 | 2008-08-04 23:43
2008-08-04 23:55 |  | Ref | 3175.5 | 3170.0763 | 2 | 2.776 |  |  |  |  | 2008-08-04 23:53 | 2008-08-04 23:58
2008-08-05 00:10 |  | Cand | 3071.5 | 3065.9051 | 2 | 4.0352 | 3168.1551 | 0.96773 | 0.0014078 |  | 2008-08-05 00:08 | 2008-08-05 00:13
2008-08-05 00:25 |  | Ref | 3172 | 3166.2338 | 2 | 2.776 |  |  |  |  | 2008-08-05 00:23 | 2008-08-05 00:28
2008-08-05 00:40 |  | Cand | 3081.5 | 3075.5625 | 2 | 4.0352 | 3167.0625 | 0.97111 | 0.0014091 |  | 2008-08-05 00:38 | 2008-08-05 00:43
2008-08-05 00:55 |  | Ref | 3174 | 3167.8913 | 2 | 2.776 |  |  |  |  | 2008-08-05 00:53 | 2008-08-05 00:58
2008-08-05 01:10 |  | Cand | 3074 | 3067.72 | 2 | 4.0352 | 3169.22 | 0.96797 | 0.0014073 |  | 2008-08-05 01:08 | 2008-08-05 01:13
2008-08-05 01:25 |  | Ref | 3177 | 3170.5487 | 2 | 2.776 |  |  |  |  | 2008-08-05 01:23 | 2008-08-05 01:28
2008-08-05 01:40 |  | Cand | 3074 | 3067.3774 | 2 | 4.0352 | 3169.7941 | 0.96769 | 0.0013999 |  | 2008-08-05 01:38 | 2008-08-05 01:43
2008-08-05 02:10 |  | Ref | 3175.25 | 3168.2849 | 8 | 1.388 |  |  |  |  | 2008-08-05 01:53 | 2008-08-05 02:28
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
2008-08-05 18:05 |  | Zero | 17.8695 | [7 empty] | 2008-08-05 18:03 | 2008-08-05 18:08
Bracketing Intervals
Detector Linearity | Linearity Factor at 2.7 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 5.7067 | 2008-08-04 23:38 | 2008-08-05 01:43
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 3.9259 | 2008-08-04 23:28 | 2008-08-05 02:28
| Degrees of Freedom |  |  | 4
Date/Time |  | Meas. Type | "B-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "B-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 16:03 |  | Zero | 0 | [7 empty] | 2008-08-04 16:00 | 2008-08-04 16:05
2008-08-04 23:30 |  | Ref | 3132 | 3127.3421 | 1 | 3.9954 |  |  |  |  | 2008-08-04 23:30 | 2008-08-04 23:30
2008-08-04 23:40 |  | Cand | 3039.6667 | 3034.9047 | 3 | 1.8597 | 3132.5714 | 0.96882 | 0.00099246 |  | 2008-08-04 23:35 | 2008-08-04 23:45
2008-08-04 23:55 |  | Ref | 3145.3333 | 3140.4153 | 3 | 2.3068 |  |  |  |  | 2008-08-04 23:50 | 2008-08-05
2008-08-05 00:10 |  | Cand | 3047 | 3041.9258 | 3 | 1.8597 | 3139.4258 | 0.96894 | 0.00077738 |  | 2008-08-05 00:05 | 2008-08-05 00:15
2008-08-05 00:25 |  | Ref | 3143.6667 | 3138.4363 | 3 | 2.3068 |  |  |  |  | 2008-08-05 00:20 | 2008-08-05 00:30
2008-08-05 00:40 |  | Cand | 3051.3333 | 3045.9469 | 3 | 1.8597 | 3138.7802 | 0.97042 | 0.00077804 |  | 2008-08-05 00:35 | 2008-08-05 00:45
2008-08-05 00:55 |  | Ref | 3144.6667 | 3139.1241 | 3 | 2.3068 |  |  |  |  | 2008-08-05 00:50 | 2008-08-05 01:00
2008-08-05 01:10 |  | Cand | 3048.6667 | 3042.9679 | 3 | 1.8597 | 3140.8013 | 0.96885 | 0.00077701 |  | 2008-08-05 01:05 | 2008-08-05 01:15
2008-08-05 01:25 |  | Ref | 3148.3333 | 3142.4785 | 3 | 2.3068 |  |  |  |  | 2008-08-05 01:20 | 2008-08-05 01:30
2008-08-05 01:40 |  | Cand | 3050.6667 | 3044.6557 | 3 | 1.8597 | 3141.0631 | 0.96931 | 0.00077093 |  | 2008-08-05 01:35 | 2008-08-05 01:45
2008-08-05 02:10 |  | Ref | 3144.5556 | 3138.2323 | 9 | 1.3318 |  |  |  |  | 2008-08-05 01:50 | 2008-08-05 02:30
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
2008-08-05 18:08 |  | Zero | 16.2895 | [7 empty] | 2008-08-05 18:05 | 2008-08-05 18:10
Bracketing Intervals
Detector Linearity | Linearity Factor at 2.7 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 3.221 | 2008-08-04 23:35 | 2008-08-05 01:45
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 3.9954 | 2008-08-04 23:30 | 2008-08-05 02:30
| Degrees of Freedom |  |  | 4
| Ratio Average Uncertainty Components, 1s |  |  | Pooled Detector Linearity |  |  |  | 0
|  |  |  | Measurement Stability |  |  |  | 0.00036759
|  |  |  | Ratio Repeatability |  |  |  | 0.0003325
|  | Set Average Candidate/Reference Ratio |  |  |  |  | 0.96905
|  | Number of Brackets in this Set |  |  |  |  | 10
|  | Uncertainty for This Set's Average, 1s |  |  |  |  | 0.00049566
|  | Nested "Level-2" StdErr, s2 |  |  |  |  | 0.0011254
Formulas:
First Record = MIN(IF(A_Cand="","",IF(A_Time>[First Record, row above],A_Time,"")))
First Record = MIN(IF(A_Ref="","",IF(A_Time>[First Record, row above],A_Time,"")))
Last Record = MIN(IF(A_Zero="",IF(A_Time>[First Record],A_Time1,""),""))
Last Record = MIN(IF(A_Ref="",IF(A_Time>[First Record],A_Time1,""),""))
Note on A8: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on A31: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on D3: Ed Brown: The "nominal" reference setpoint, used for identification, not for calculation
Note on D4: Ed Brown: Actual certified reference concentration
Note on D6: Ed Brown: The "nominal" candidate setpoint, used for identification, not for calculation
Note on D9: Ed Brown: Average response for measurement interval
Note on D32: Ed Brown: Average response for measurement interval
Note on E9: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E27: Ed Brown: Value from Linearity Chack reponse table
Note on E29: Ed Brown: Degrees of Freedom used for pooling below
Note on E32: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E50: Ed Brown: Value from Linearity Chack reponse table
Note on E52: Ed Brown: Degrees of Freedom used for pooling below
Note on F10: Ed Brown: Number of individual readings in measurement interval
Note on F33: Ed Brown: Number of individual readings in measurement interval
Note on G10: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on G33: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on H9: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H32: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H60: Ed Brown: The "End Result" of this bracketing procedure
Note on H62: Ed Brown: Uncertainty of the ratio result of this bracketing procedure
Note on I9: Ed Brown: Calculated measurement ratio for bracket
Note on I32: Ed Brown: Calculated measurement ratio for bracket
Note on J9: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on J32: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on K27: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K28: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on K50: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K51: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on M26: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Note on M49: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Populate Parse Table

## Point-4
Single-Point Bracketing Data
Reference Setpoint |  |  | 5.7 | mg/m3 |  | Bracketing Procedure Start |  | 2008-08-04 15:25
Reference Concentration |  |  | 6.0937 | mg/m3
[6 empty] | Bracketing Procedure End |  | 2008-08-05 18:10
Candidate Setpoint |  |  | 5.7 | mg/m3
Date/Time |  | Meas. Type | "A-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "A-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 16:00 |  | Zero | 0 | [7 empty] | 2008-08-04 15:58 | 2008-08-04 16:03
2008-08-05 02:45 |  | Ref | 6701 | 6693.6353 | 2 | 3.2562 |  |  |  |  | 2008-08-05 02:43 | 2008-08-05 02:48
2008-08-05 03:00 |  | Cand | 6553 | 6545.464 | 2 | 3.9961 | 6699.464 | 0.97701 | 0.00068451 |  | 2008-08-05 02:58 | 2008-08-05 03:03
2008-08-05 03:15 |  | Ref | 6713 | 6705.2927 | 2 | 3.2562 |  |  |  |  | 2008-08-05 03:13 | 2008-08-05 03:18
2008-08-05 03:30 |  | Cand | 6548.5 | 6540.6214 | 2 | 3.9961 | 6705.1214 | 0.97547 | 0.00068367 |  | 2008-08-05 03:28 | 2008-08-05 03:33
2008-08-05 03:45 |  | Ref | 6713 | 6704.9502 | 2 | 3.2562 |  |  |  |  | 2008-08-05 03:43 | 2008-08-05 03:48
2008-08-05 04:00 |  | Cand | 6559 | 6550.7789 | 2 | 3.9961 | 6705.2789 | 0.97696 | 0.0006839 |  | 2008-08-05 03:58 | 2008-08-05 04:03
2008-08-05 04:15 |  | Ref | 6714 | 6705.6076 | 2 | 3.2562 |  |  |  |  | 2008-08-05 04:13 | 2008-08-05 04:18
2008-08-05 04:30 |  | Cand | 6558 | 6549.4363 | 2 | 3.9961 | 6707.1863 | 0.97648 | 0.00068363 |  | 2008-08-05 04:28 | 2008-08-05 04:33
2008-08-05 04:45 |  | Ref | 6717.5 | 6708.7651 | 2 | 3.2562 |  |  |  |  | 2008-08-05 04:43 | 2008-08-05 04:48
2008-08-05 05:00 |  | Cand | 6560.5 | 6551.5938 | 2 | 3.9961 | 6706.8438 | 0.97685 | 0.00068373 |  | 2008-08-05 04:58 | 2008-08-05 05:03
2008-08-05 05:15 |  | Ref | 6714 | 6704.9225 | 2 | 3.2562 |  |  |  |  | 2008-08-05 05:13 | 2008-08-05 05:18
2008-08-05 05:30 |  | Cand | 6567 | 6557.7512 | 2 | 3.9961 | 6705.0012 | 0.97804 | 0.00068411 |  | 2008-08-05 05:28 | 2008-08-05 05:33
2008-08-05 05:45 |  | Ref | 6714.5 | 6705.08 | 2 | 3.2562 |  |  |  |  | 2008-08-05 05:43 | 2008-08-05 05:48
2008-08-05 18:05 |  | Zero | 17.8695 | [7 empty] | 2008-08-05 18:03 | 2008-08-05 18:08
Bracketing Intervals
Detector Linearity | Linearity Factor at 5.7 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 5.6514 | 2008-08-05 02:58 | 2008-08-05 05:33
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 4.605 | 2008-08-05 02:43 | 2008-08-05 05:48
| Degrees of Freedom |  |  | 5
Date/Time |  | Meas. Type | "B-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "B-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 16:03 |  | Zero | 0 | [7 empty] | 2008-08-04 16:00 | 2008-08-04 16:05
2008-08-05 02:45 |  | Ref | 6638.6667 | 6631.9791 | 3 | 3.4383 |  |  |  |  | 2008-08-05 02:40 | 2008-08-05 02:50
2008-08-05 03:00 |  | Cand | 6496.3333 | 6489.4897 | 3 | 4.5076 | 6639.4897 | 0.97741 | 0.00076748 |  | 2008-08-05 02:55 | 2008-08-05 03:05
2008-08-05 03:15 |  | Ref | 6654 | 6647.0002 | 3 | 3.4383 |  |  |  |  | 2008-08-05 03:10 | 2008-08-05 03:20
2008-08-05 03:30 |  | Cand | 6493.3333 | 6486.1774 | 3 | 4.5076 | 6647.5107 | 0.97573 | 0.00076627 |  | 2008-08-05 03:25 | 2008-08-05 03:35
2008-08-05 03:45 |  | Ref | 6655.3333 | 6648.0213 | 3 | 3.4383 |  |  |  |  | 2008-08-05 03:40 | 2008-08-05 03:50
2008-08-05 04:00 |  | Cand | 6509 | 6501.5318 | 3 | 4.5076 | 6648.6985 | 0.97787 | 0.00076649 |  | 2008-08-05 03:55 | 2008-08-05 04:05
2008-08-05 04:15 |  | Ref | 6657 | 6649.3757 | 3 | 3.4383 |  |  |  |  | 2008-08-05 04:10 | 2008-08-05 04:20
2008-08-05 04:30 |  | Cand | 6496.3333 | 6488.5529 | 3 | 4.5076 | 6649.5529 | 0.97579 | 0.00076604 |  | 2008-08-05 04:25 | 2008-08-05 04:35
2008-08-05 04:45 |  | Ref | 6657.6667 | 6649.7301 | 3 | 3.4383 |  |  |  |  | 2008-08-05 04:40 | 2008-08-05 04:50
2008-08-05 05:00 |  | Cand | 6502 | 6493.9073 | 3 | 4.5076 | 6652.074 | 0.97622 | 0.00076582 |  | 2008-08-05 04:55 | 2008-08-05 05:05
2008-08-05 05:15 |  | Ref | 6662.6667 | 6654.4178 | 3 | 3.4383 |  |  |  |  | 2008-08-05 05:10 | 2008-08-05 05:20
2008-08-05 05:30 |  | Cand | 6513 | 6504.595 | 3 | 4.5076 | 6651.2617 | 0.97795 | 0.00076621 |  | 2008-08-05 05:25 | 2008-08-05 05:35
2008-08-05 05:45 |  | Ref | 6656.6667 | 6648.1056 | 3 | 3.4383 |  |  |  |  | 2008-08-05 05:40 | 2008-08-05 05:50
2008-08-05 18:08 |  | Zero | 16.2895 | [7 empty] | 2008-08-05 18:05 | 2008-08-05 18:10
Bracketing Intervals
Detector Linearity | Linearity Factor at 5.7 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 7.8075 | 2008-08-05 02:55 | 2008-08-05 05:35
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 5.9554 | 2008-08-05 02:40 | 2008-08-05 05:50
| Degrees of Freedom |  |  | 5
| Ratio Average Uncertainty Components, 1s |  |  | Pooled Detector Linearity |  |  |  | 0
|  |  |  | Measurement Stability |  |  |  | 0.00020967
|  |  |  | Ratio Repeatability |  |  |  | 0.00024677
|  | Set Average Candidate/Reference Ratio |  |  |  |  | 0.97681
|  | Number of Brackets in this Set |  |  |  |  | 12
|  | Uncertainty for This Set's Average, 1s |  |  |  |  | 0.00032382
|  | Nested "Level-2" StdErr, s2 |  |  |  |  | 0.00089505
Formulas:
First Record = MIN(IF(A_Cand="","",IF(A_Time>[First Record, row above],A_Time,"")))
First Record = MIN(IF(A_Ref="","",IF(A_Time>[First Record, row above],A_Time,"")))
Last Record = MIN(IF(A_Zero="",IF(A_Time>[First Record],A_Time1,""),""))
Last Record = MIN(IF(A_Ref="",IF(A_Time>[First Record],A_Time1,""),""))
Note on A8: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on A31: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on D3: Ed Brown: The "nominal" reference setpoint, used for identification, not for calculation
Note on D4: Ed Brown: Actual certified reference concentration
Note on D6: Ed Brown: The "nominal" candidate setpoint, used for identification, not for calculation
Note on D9: Ed Brown: Average response for measurement interval
Note on D32: Ed Brown: Average response for measurement interval
Note on E9: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E27: Ed Brown: Value from Linearity Chack reponse table
Note on E29: Ed Brown: Degrees of Freedom used for pooling below
Note on E32: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E50: Ed Brown: Value from Linearity Chack reponse table
Note on E52: Ed Brown: Degrees of Freedom used for pooling below
Note on F10: Ed Brown: Number of individual readings in measurement interval
Note on F33: Ed Brown: Number of individual readings in measurement interval
Note on G10: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on G33: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on H9: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H32: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H60: Ed Brown: The "End Result" of this bracketing procedure
Note on H62: Ed Brown: Uncertainty of the ratio result of this bracketing procedure
Note on I9: Ed Brown: Calculated measurement ratio for bracket
Note on I32: Ed Brown: Calculated measurement ratio for bracket
Note on J9: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on J32: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on K27: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K28: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on K50: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K51: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on M26: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Note on M49: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Populate Parse Table

## Point-5
Single-Point Bracketing Data
Reference Setpoint |  |  | 9.5 | mg/m3 |  | Bracketing Procedure Start |  | 2008-08-04 15:25
Reference Concentration |  |  | 10.4051 | mg/m3
[6 empty] | Bracketing Procedure End |  | 2008-08-05 18:10
Candidate Setpoint |  |  | 9.5 | mg/m3
Date/Time |  | Meas. Type | "A-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "A-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 16:00 |  | Zero | 0 | [7 empty] | 2008-08-04 15:58 | 2008-08-04 16:03
2008-08-05 06:00 |  | Ref | 11076.5 | 11066.9087 | 2 | 30.4039 |  |  |  |  | 2008-08-05 05:58 | 2008-08-05 06:03
2008-08-05 06:15 |  | Cand | 10789.5 | 10779.7374 | 2 | 26.1947 | 11012.7374 | 0.97884 | 0.0030511 |  | 2008-08-05 06:13 | 2008-08-05 06:18
2008-08-05 06:30 |  | Ref | 10968.5 | 10958.5662 | 2 | 30.4039 |  |  |  |  | 2008-08-05 06:28 | 2008-08-05 06:33
2008-08-05 06:45 |  | Cand | 10670.5 | 10660.3949 | 2 | 26.1947 | 10924.3949 | 0.97583 | 0.0030721 |  | 2008-08-05 06:43 | 2008-08-05 06:48
2008-08-05 07:00 |  | Ref | 10900.5 | 10890.2236 | 2 | 30.4039 |  |  |  |  | 2008-08-05 06:58 | 2008-08-05 07:03
2008-08-05 07:15 |  | Cand | 10662 | 10651.5523 | 2 | 26.1947 | 10888.8023 | 0.97821 | 0.003085 |  | 2008-08-05 07:13 | 2008-08-05 07:18
2008-08-05 07:30 |  | Ref | 10898 | 10887.3811 | 2 | 30.4039 |  |  |  |  | 2008-08-05 07:28 | 2008-08-05 07:33
2008-08-05 07:45 |  | Cand | 10661.5 | 10650.7098 | 2 | 26.1947 | 10895.2098 | 0.97756 | 0.0030824 |  | 2008-08-05 07:43 | 2008-08-05 07:48
2008-08-05 08:00 |  | Ref | 10914 | 10903.0385 | 2 | 30.4039 |  |  |  |  | 2008-08-05 07:58 | 2008-08-05 08:03
2008-08-05 08:15 |  | Cand | 10647 | 10635.8672 | 2 | 26.1947 | 10893.1172 | 0.97638 | 0.0030815 |  | 2008-08-05 08:13 | 2008-08-05 08:18
2008-08-05 08:30 |  | Ref | 10894.5 | 10883.196 | 2 | 30.4039 |  |  |  |  | 2008-08-05 08:28 | 2008-08-05 08:33
2008-08-05 08:45 |  | Cand | 10632 | 10620.5247 | 2 | 26.1947 | 10877.7747 | 0.97635 | 0.0030859 |  | 2008-08-05 08:43 | 2008-08-05 08:48
2008-08-05 09:00 |  | Ref | 10884 | 10872.3534 | 2 | 30.4039 |  |  |  |  | 2008-08-05 08:58 | 2008-08-05 09:03
2008-08-05 18:05 |  | Zero | 17.8695 | [7 empty] | 2008-08-05 18:03 | 2008-08-05 18:08
Bracketing Intervals
Detector Linearity | Linearity Factor at 9.5 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 37.0449 | 2008-08-05 06:13 | 2008-08-05 08:48
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 42.9976 | 2008-08-05 05:58 | 2008-08-05 09:03
| Degrees of Freedom |  |  | 5
Date/Time |  | Meas. Type | "B-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "B-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 16:03 |  | Zero | 0 | [7 empty] | 2008-08-04 16:00 | 2008-08-04 16:05
2008-08-05 06:00 |  | Ref | 10962.3333 | 10953.6161 | 3 | 23.0891 |  |  |  |  | 2008-08-05 05:55 | 2008-08-05 06:05
2008-08-05 06:15 |  | Cand | 10702.6667 | 10693.7933 | 3 | 22.3805 | 10909.96 | 0.98019 | 0.0025219 |  | 2008-08-05 06:10 | 2008-08-05 06:20
2008-08-05 06:30 |  | Ref | 10875.3333 | 10866.3039 | 3 | 23.0891 |  |  |  |  | 2008-08-05 06:25 | 2008-08-05 06:35
2008-08-05 06:45 |  | Cand | 10587 | 10577.8144 | 3 | 22.3805 | 10827.4811 | 0.97694 | 0.0025382 |  | 2008-08-05 06:40 | 2008-08-05 06:50
2008-08-05 07:00 |  | Ref | 10798 | 10788.6583 | 3 | 23.0891 |  |  |  |  | 2008-08-05 06:55 | 2008-08-05 07:05
2008-08-05 07:15 |  | Cand | 10563 | 10553.5021 | 3 | 22.3805 | 10796.1688 | 0.97752 | 0.0025461 |  | 2008-08-05 07:10 | 2008-08-05 07:20
2008-08-05 07:30 |  | Ref | 10813.3333 | 10803.6793 | 3 | 23.0891 |  |  |  |  | 2008-08-05 07:25 | 2008-08-05 07:35
2008-08-05 07:45 |  | Cand | 10582 | 10572.1899 | 3 | 22.3805 | 10802.5232 | 0.97868 | 0.0025456 |  | 2008-08-05 07:40 | 2008-08-05 07:50
2008-08-05 08:00 |  | Ref | 10811.3333 | 10801.3671 | 3 | 23.0891 |  |  |  |  | 2008-08-05 07:55 | 2008-08-05 08:05
2008-08-05 08:15 |  | Cand | 10561.6667 | 10551.5443 | 3 | 22.3805 | 10795.5443 | 0.9774 | 0.0025461 |  | 2008-08-05 08:10 | 2008-08-05 08:20
2008-08-05 08:30 |  | Ref | 10800 | 10789.7215 | 3 | 23.0891 |  |  |  |  | 2008-08-05 08:25 | 2008-08-05 08:35
2008-08-05 08:45 |  | Cand | 10558 | 10547.5654 | 3 | 22.3805 | 10786.732 | 0.97783 | 0.0025486 |  | 2008-08-05 08:40 | 2008-08-05 08:50
2008-08-05 09:00 |  | Ref | 10794.3333 | 10783.7426 | 3 | 23.0891 |  |  |  |  | 2008-08-05 08:55 | 2008-08-05 09:05
2008-08-05 18:08 |  | Zero | 16.2895 | [7 empty] | 2008-08-05 18:05 | 2008-08-05 18:10
Bracketing Intervals
Detector Linearity | Linearity Factor at 9.5 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 38.7642 | 2008-08-05 06:10 | 2008-08-05 08:50
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 39.9916 | 2008-08-05 05:55 | 2008-08-05 09:05
| Degrees of Freedom |  |  | 5
| Ratio Average Uncertainty Components, 1s |  |  | Pooled Detector Linearity |  |  |  | 0
|  |  |  | Measurement Stability |  |  |  | 0.00081448
|  |  |  | Ratio Repeatability |  |  |  | 0.00018992
|  | Set Average Candidate/Reference Ratio |  |  |  |  | 0.97764
|  | Number of Brackets in this Set |  |  |  |  | 12
|  | Uncertainty for This Set's Average, 1s |  |  |  |  | 0.00083633
|  | Nested "Level-2" StdErr, s2 |  |  |  |  | 0.0012224
Formulas:
First Record = MIN(IF(A_Cand="","",IF(A_Time>[First Record, row above],A_Time,"")))
First Record = MIN(IF(A_Ref="","",IF(A_Time>[First Record, row above],A_Time,"")))
Last Record = MIN(IF(A_Zero="",IF(A_Time>[First Record],A_Time1,""),""))
Last Record = MIN(IF(A_Ref="",IF(A_Time>[First Record],A_Time1,""),""))
Note on A8: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on A31: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on D3: Ed Brown: The "nominal" reference setpoint, used for identification, not for calculation
Note on D4: Ed Brown: Actual certified reference concentration
Note on D6: Ed Brown: The "nominal" candidate setpoint, used for identification, not for calculation
Note on D9: Ed Brown: Average response for measurement interval
Note on D32: Ed Brown: Average response for measurement interval
Note on E9: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E27: Ed Brown: Value from Linearity Chack reponse table
Note on E29: Ed Brown: Degrees of Freedom used for pooling below
Note on E32: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E50: Ed Brown: Value from Linearity Chack reponse table
Note on E52: Ed Brown: Degrees of Freedom used for pooling below
Note on F10: Ed Brown: Number of individual readings in measurement interval
Note on F33: Ed Brown: Number of individual readings in measurement interval
Note on G10: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on G33: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on H9: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H32: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H60: Ed Brown: The "End Result" of this bracketing procedure
Note on H62: Ed Brown: Uncertainty of the ratio result of this bracketing procedure
Note on I9: Ed Brown: Calculated measurement ratio for bracket
Note on I32: Ed Brown: Calculated measurement ratio for bracket
Note on J9: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on J32: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on K27: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K28: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on K50: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K51: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on M26: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Note on M49: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Populate Parse Table

## Point-6
Single-Point Bracketing Data
Reference Setpoint |  |  | 19 | mg/m3 |  | Bracketing Procedure Start |  | 2008-08-04 15:25
Reference Concentration |  |  | 20.4651 | mg/m3
[6 empty] | Bracketing Procedure End |  | 2008-08-05 18:10
Candidate Setpoint |  |  | 19 | mg/m3
Date/Time |  | Meas. Type | "A-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "A-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 16:00 |  | Zero | 0 | [7 empty] | 2008-08-04 15:58 | 2008-08-04 16:03
|  | Ref |  |  | 0
2008-08-05 10:00 |  | Cand | 20663 | 20650.6683 | 2 | 31.3513 |  | 0 |  |  | 2008-08-05 09:58 | 2008-08-05 10:03
2008-08-05 10:15 |  | Ref | 21137 | 21124.4971 | 2 | 27.9639 |  |  |  |  | 2008-08-05 10:13 | 2008-08-05 10:18
2008-08-05 10:30 |  | Cand | 20746 | 20733.3258 | 2 | 31.3513 | 21143.3258 | 0.98061 | 0.0017435 |  | 2008-08-05 10:28 | 2008-08-05 10:33
2008-08-05 10:45 |  | Ref | 21175 | 21162.1545 | 2 | 27.9639 |  |  |  |  | 2008-08-05 10:43 | 2008-08-05 10:48
2008-08-05 11:00 |  | Cand | 20782 | 20768.9832 | 2 | 31.3513 | 21192.9832 | 0.97999 | 0.0017391 |  | 2008-08-05 10:58 | 2008-08-05 11:03
2008-08-05 11:15 |  | Ref | 21237 | 21223.812 | 2 | 27.9639 |  |  |  |  | 2008-08-05 11:13 | 2008-08-05 11:18
2008-08-05 11:30 |  | Cand | 20794.5 | 20781.1407 | 2 | 31.3513 | 21213.6407 | 0.97961 | 0.0017372 |  | 2008-08-05 11:28 | 2008-08-05 11:33
2008-08-05 11:45 |  | Ref | 21217 | 21203.4694 | 2 | 27.9639 |  |  |  |  | 2008-08-05 11:43 | 2008-08-05 11:48
2008-08-05 12:00 |  | Cand | 20742.5 | 20728.7981 | 2 | 31.3513 | 21182.493 | 0.97858 | 0.0017508 |  | 2008-08-05 11:58 | 2008-08-05 12:03
2008-08-05 12:38 |  | Ref | 21144.1818 | 21130.0518 | 11 | 11.9238 |  |  |  |  | 2008-08-05 12:13 | 2008-08-05 13:03
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
2008-08-05 18:05 |  | Zero | 17.8695 | [7 empty] | 2008-08-05 18:03 | 2008-08-05 18:08
Bracketing Intervals
Detector Linearity | Linearity Factor at 19 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 44.3375 | 2008-08-05 09:58 | 2008-08-05 12:03
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 39.5469 | 2008-08-05 09:58 | 2008-08-05 13:03
| Degrees of Freedom |  |  | 3
Date/Time |  | Meas. Type | "B-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "B-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 16:03 |  | Zero | 0 | [7 empty] | 2008-08-04 16:00 | 2008-08-04 16:05
2008-08-05 10:14 |  | Ref | 20967.3333 | 20955.9749 | 3 | 24.9021 |  |  |  |  | 2008-08-05 10:08 | 2008-08-05 10:20
2008-08-05 10:30 |  | Cand | 20532.3333 | 20520.8058 | 3 | 22.2892 | 20970.8858 | 0.97854 | 0.0013438 |  | 2008-08-05 10:25 | 2008-08-05 10:35
2008-08-05 10:45 |  | Ref | 20996.3333 | 20984.6497 | 3 | 24.9021 |  |  |  |  | 2008-08-05 10:40 | 2008-08-05 10:50
2008-08-05 11:00 |  | Cand | 20598 | 20586.1602 | 3 | 22.2892 | 21017.6602 | 0.97947 | 0.0013409 |  | 2008-08-05 10:55 | 2008-08-05 11:05
2008-08-05 11:15 |  | Ref | 21062.6667 | 21050.6707 | 3 | 24.9021 |  |  |  |  | 2008-08-05 11:10 | 2008-08-05 11:20
2008-08-05 11:30 |  | Cand | 20586.3333 | 20574.1813 | 3 | 22.2892 | 21030.8479 | 0.97829 | 0.0013395 |  | 2008-08-05 11:25 | 2008-08-05 11:35
2008-08-05 11:45 |  | Ref | 21023.3333 | 21011.0251 | 3 | 24.9021 |  |  |  |  | 2008-08-05 11:40 | 2008-08-05 11:50
2008-08-05 12:00 |  | Cand | 20544 | 20531.5357 | 3 | 22.2892 | 20988.1071 | 0.97825 | 0.0013569 |  | 2008-08-05 11:55 | 2008-08-05 12:05
2008-08-05 12:38 |  | Ref | 20943.6667 | 20930.812 | 12 | 12.4511 |  |  |  |  | 2008-08-05 12:10 | 2008-08-05 13:05
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
|  | Cand |  |  | 0 |  |  | 0
|  | Ref |  |  | 0
2008-08-05 18:08 |  | Zero | 16.2895 | [7 empty] | 2008-08-05 18:05 | 2008-08-05 18:10
Bracketing Intervals
Detector Linearity | Linearity Factor at 19 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 38.606 | 2008-08-05 10:25 | 2008-08-05 12:05
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 43.1318 | 2008-08-05 10:08 | 2008-08-05 13:05
| Degrees of Freedom |  |  | 3
| Ratio Average Uncertainty Components, 1s |  |  | Pooled Detector Linearity |  |  |  | 0
|  |  |  | Measurement Stability |  |  |  | 0.00055038
|  |  |  | Ratio Repeatability |  |  |  | 0.00025613
|  | Set Average Candidate/Reference Ratio |  |  |  |  | 0.97917
|  | Number of Brackets in this Set |  |  |  |  | 8
|  | Uncertainty for This Set's Average, 1s |  |  |  |  | 0.00060706
|  | Nested "Level-2" StdErr, s2 |  |  |  |  | 0.00087941
Formulas:
First Record = MIN(IF(A_Ref="","",IF(A_Time>[First Record, row above],A_Time,"")))
First Record = MIN(IF(A_Cand="","",IF(A_Time>[First Record, row above],A_Time,"")))
Last Record = MIN(IF(A_Zero="",IF(A_Time>[First Record],A_Time1,""),""))
Last Record = MIN(IF(A_Cand="",IF(A_Time>[First Record],A_Time1,""),""))
Note on A8: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on A31: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on D3: Ed Brown: The "nominal" reference setpoint, used for identification, not for calculation
Note on D4: Ed Brown: Actual certified reference concentration
Note on D6: Ed Brown: The "nominal" candidate setpoint, used for identification, not for calculation
Note on D9: Ed Brown: Average response for measurement interval
Note on D32: Ed Brown: Average response for measurement interval
Note on E9: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E27: Ed Brown: Value from Linearity Chack reponse table
Note on E29: Ed Brown: Degrees of Freedom used for pooling below
Note on E32: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E50: Ed Brown: Value from Linearity Chack reponse table
Note on E52: Ed Brown: Degrees of Freedom used for pooling below
Note on F10: Ed Brown: Number of individual readings in measurement interval
Note on F33: Ed Brown: Number of individual readings in measurement interval
Note on G10: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on G33: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on H9: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H32: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H60: Ed Brown: The "End Result" of this bracketing procedure
Note on H62: Ed Brown: Uncertainty of the ratio result of this bracketing procedure
Note on I9: Ed Brown: Calculated measurement ratio for bracket
Note on I32: Ed Brown: Calculated measurement ratio for bracket
Note on J9: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on J32: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on K27: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K28: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on K50: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K51: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on M26: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Note on M49: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Populate Parse Table

## Point-7
Single-Point Bracketing Data
Reference Setpoint |  |  | 38 | mg/m3 |  | Bracketing Procedure Start |  | 2008-08-04 15:25
Reference Concentration |  |  | 39.6771 | mg/m3
[6 empty] | Bracketing Procedure End |  | 2008-08-05 18:10
Candidate Setpoint |  |  | 38 | mg/m3
Date/Time |  | Meas. Type | "A-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "A-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 16:00 |  | Zero | 0 | [7 empty] | 2008-08-04 15:58 | 2008-08-04 16:03
2008-08-05 13:58 |  | Ref | 40253 | 40237.9565 | 1 | 70.2705 |  |  |  |  | 2008-08-05 13:58 | 2008-08-05 13:58
2008-08-05 14:10 |  | Cand | 39780.5 | 39765.3138 | 2 | 34.9179 | 40330.3138 | 0.98599 | 0.001391 |  | 2008-08-05 14:08 | 2008-08-05 14:13
2008-08-05 14:25 |  | Ref | 40456.5 | 40441.1425 | 2 | 49.6888 |  |  |  |  | 2008-08-05 14:23 | 2008-08-05 14:28
2008-08-05 14:40 |  | Cand | 39819.5 | 39803.9712 | 2 | 34.9179 | 40463.4712 | 0.9837 | 0.0012142 |  | 2008-08-05 14:38 | 2008-08-05 14:43
2008-08-05 14:55 |  | Ref | 40501.5 | 40485.8 | 2 | 49.6888 |  |  |  |  | 2008-08-05 14:53 | 2008-08-05 14:58
2008-08-05 15:10 |  | Cand | 39868.5 | 39852.6287 | 2 | 34.9179 | 40497.3787 | 0.98408 | 0.0012134 |  | 2008-08-05 15:08 | 2008-08-05 15:13
2008-08-05 15:25 |  | Ref | 40525 | 40508.9574 | 2 | 49.6888 |  |  |  |  | 2008-08-05 15:23 | 2008-08-05 15:28
2008-08-05 15:40 |  | Cand | 39813.5 | 39797.2861 | 2 | 34.9179 | 40467.5361 | 0.98344 | 0.0012139 |  | 2008-08-05 15:38 | 2008-08-05 15:43
2008-08-05 15:55 |  | Ref | 40442.5 | 40426.1149 | 2 | 49.6888 |  |  |  |  | 2008-08-05 15:53 | 2008-08-05 15:58
2008-08-05 16:10 |  | Cand | 39758 | 39741.4436 | 2 | 34.9179 | 40421.2636 | 0.98318 | 0.0011685 |  | 2008-08-05 16:08 | 2008-08-05 16:13
2008-08-05 16:33 |  | Ref | 40430.8 | 40413.9867 | 5 | 31.4259 |  |  |  |  | 2008-08-05 16:23 | 2008-08-05 16:43
2008-08-05 16:55 |  | Cand | 39755.5 | 39738.4298 | 2 | 34.9179 | 40445.6004 | 0.98252 | 0.0011324 |  | 2008-08-05 16:53 | 2008-08-05 16:58
2008-08-05 17:15 |  | Ref | 40491 | 40473.7014 | 2 | 49.6888 |  |  |  |  | 2008-08-05 17:13 | 2008-08-05 17:18
2008-08-05 18:05 |  | Zero | 17.8695 | [7 empty] | 2008-08-05 18:03 | 2008-08-05 18:08
Bracketing Intervals
Detector Linearity | Linearity Factor at 38 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 49.3814 | 2008-08-05 14:08 | 2008-08-05 16:58
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 70.2705 | 2008-08-05 13:58 | 2008-08-05 17:18
| Degrees of Freedom |  |  | 5
Date/Time |  | Meas. Type | "B-Channel" Responses |  |  |  |  | Ref/Cand Ratios |  |  | "B-Channel" Meas. Intervals
|  |  | Average | mavg, Zero Adjusted | SE(mavg) |  | Ref. Interp. | Ratio | Stability Uncertainty |  | First Record | Last Record
[5 empty] | n | SE/Ön |  |  |  |  | Date/Time | Date/Time
2008-08-04 16:03 |  | Zero | 0 | [7 empty] | 2008-08-04 16:00 | 2008-08-04 16:05
|  | Ref |  |  | 0
2008-08-05 14:10 |  | Cand | 39347.6667 | 39333.8492 | 3 | 32.5829 |  | 0 |  |  | 2008-08-05 14:05 | 2008-08-05 14:15
2008-08-05 14:25 |  | Ref | 40050.6667 | 40036.6931 | 3 | 34.5531 |  |  |  |  | 2008-08-05 14:20 | 2008-08-05 14:30
2008-08-05 14:40 |  | Cand | 39411 | 39396.8703 | 3 | 32.5829 | 40067.3703 | 0.98327 | 0.0010103 |  | 2008-08-05 14:35 | 2008-08-05 14:45
2008-08-05 14:55 |  | Ref | 40112.3333 | 40098.0475 | 3 | 34.5531 |  |  |  |  | 2008-08-05 14:50 | 2008-08-05 15:00
2008-08-05 15:10 |  | Cand | 39465.6667 | 39451.2247 | 3 | 32.5829 | 40078.8914 | 0.98434 | 0.0010104 |  | 2008-08-05 15:05 | 2008-08-05 15:15
2008-08-05 15:25 |  | Ref | 40074.3333 | 40059.7352 | 3 | 34.5531 |  |  |  |  | 2008-08-05 15:20 | 2008-08-05 15:30
2008-08-05 15:40 |  | Cand | 39423.6667 | 39408.9124 | 3 | 32.5829 | 40033.2458 | 0.9844 | 0.0010116 |  | 2008-08-05 15:35 | 2008-08-05 15:45
2008-08-05 15:55 |  | Ref | 40021.6667 | 40006.7563 | 3 | 34.5531 |  |  |  |  | 2008-08-05 15:50 | 2008-08-05 16:00
2008-08-05 16:10 |  | Cand | 39401.6667 | 39386.6002 | 3 | 32.5829 | 40016.2073 | 0.98427 | 0.00099437 |  | 2008-08-05 16:05 | 2008-08-05 16:15
2008-08-05 16:35 |  | Ref | 40047.2857 | 40031.959 | 7 | 22.6203 |  |  |  |  | 2008-08-05 16:20 | 2008-08-05 16:50
2008-08-05 17:00 |  | Cand | 39371 | 39355.4131 | 3 | 32.5829 | 40073.175 | 0.98209 | 0.001083 |  | 2008-08-05 16:55 | 2008-08-05 17:05
2008-08-05 17:13 |  | Ref | 40109.5 | 40093.783 | 2 | 42.3187 |  |  |  |  | 2008-08-05 17:10 | 2008-08-05 17:15
2008-08-05 18:08 |  | Zero | 16.2895 | [7 empty] | 2008-08-05 18:05 | 2008-08-05 18:10
Bracketing Intervals
Detector Linearity | Linearity Factor at 38 |  |  | 0 |  | Drift Linearity & Stability Uncertainties, 1s |  |  | Cand | 56.4352 | 2008-08-05 14:05 | 2008-08-05 17:05
| Uncertainty, 3s |  |  | 0 |  |  |  |  | Ref | 59.8477 | 2008-08-05 14:05 | 2008-08-05 17:15
| Degrees of Freedom |  |  | 4
| Ratio Average Uncertainty Components, 1s |  |  | Pooled Detector Linearity |  |  |  | 0
|  |  |  | Measurement Stability |  |  |  | 0.00034293
|  |  |  | Ratio Repeatability |  |  |  | 0.00029515
|  | Set Average Candidate/Reference Ratio |  |  |  |  | 0.98375
|  | Number of Brackets in this Set |  |  |  |  | 11
|  | Uncertainty for This Set's Average, 1s |  |  |  |  | 0.00045246
|  | Nested "Level-2" StdErr, s2 |  |  |  |  | 0.0010542
Formulas:
First Record = MIN(IF(A_Cand="","",IF(A_Time>[First Record, row above],A_Time,"")))
First Record = MIN(IF(A_Ref="","",IF(A_Time>[First Record, row above],A_Time,"")))
Last Record = MIN(IF(A_Zero="",IF(A_Time>[First Record],A_Time1,""),""))
Last Record = MIN(IF(A_Ref="",IF(A_Time>[First Record],A_Time1,""),""))
Note on A8: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on A31: Ed Brown: RepresentativeTimestamp for measurement interval. May be entered manually, or computed as an average from the interval table to the right.
Note on D3: Ed Brown: The "nominal" reference setpoint, used for identification, not for calculation
Note on D4: Ed Brown: Actual certified reference concentration
Note on D6: Ed Brown: The "nominal" candidate setpoint, used for identification, not for calculation
Note on D9: Ed Brown: Average response for measurement interval
Note on D32: Ed Brown: Average response for measurement interval
Note on E9: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E27: Ed Brown: Value from Linearity Chack reponse table
Note on E29: Ed Brown: Degrees of Freedom used for pooling below
Note on E32: Ed Brown: Average interval measurement, adjusted for time-interpolated zero offset
Note on E50: Ed Brown: Value from Linearity Chack reponse table
Note on E52: Ed Brown: Degrees of Freedom used for pooling below
Note on F10: Ed Brown: Number of individual readings in measurement interval
Note on F33: Ed Brown: Number of individual readings in measurement interval
Note on G10: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on G33: Ed Brown: Standard error of interval average due to measurement stability and drift linearity effects
Note on H9: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H32: Ed Brown: Time-interpolated reference measurement corresponding to candidate measurement
Note on H60: Ed Brown: The "End Result" of this bracketing procedure
Note on H62: Ed Brown: Uncertainty of the ratio result of this bracketing procedure
Note on I9: Ed Brown: Calculated measurement ratio for bracket
Note on I32: Ed Brown: Calculated measurement ratio for bracket
Note on J9: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on J32: Ed Brown: Ratio uncertainty from measurement stability and drift linearity effects
Note on K27: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K28: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on K50: Ed Brown: Standard Error of Time Regression for all Candidate Measurements at ths Setting
Note on K51: Ed Brown: Standard Error of Time Regression for all Reference Measurements at ths Setting
Note on M26: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Note on M49: Ed Brown: Inclusive intervals. These may be entered manually or calculated from the table above.
Populate Parse Table

## Bracketing Results
Bracketing Procedure Results Summary
Operator: |  | Chris Winterrowd |  |  | Candidate ID: |  | Tekran 6
Company/Lab: |  | EPA/RTP Laboratories |  |  | Candidate S/N: |  | 3009
Start Date/Time: |  | 2008-08-04 15:25 |  |  | Reference ID: |  | Tekran 3
End Date/Time: |  | 2008-08-05 18:10 |  |  | Reference S/N: |  | 3039
Reference Generator Certification Parameters |  |  | Bracketing Parameters |  |  |  | Candidate Generator Parameters
|  |  | Uncertainty Terms |  | Ratio Measurements
|  |  | StdErr Among Brackets S2 | Uncertainty of Mean Ratio u(Rmean)
Setpoint ug/scm | Certified ug/scm | Uncertainty 1s ug/scm |  |  | No. of Brackets K | Average Ratio Rmean | Setpoint ug/scm | Bracketing (Ccand)j ug/scm | Spreadsheet Layer
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | Point-1
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | Point-2
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | Point-3
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | Point-4
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | Point-5
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | Point-6
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | Point-7
Note on A11: Ed Brown: The "nominal" setpoint, used for identification, not for calculation
Note on H11: Ed Brown: The "nominal" setpoint, used for identification, not for calculation

## Uncertainty
Uncertainty Calculation
|  |  | Candidate Location: | EPA/RTP Laboratories
|  |  | Candidate Device: | Tekran 6
|  |  | Candidate S/N: | 3009
|  |  | Reference ID: | Tekran 3
|  |  | Reference S/N: | 3039
|  |  | Certification Dates: | 4-Aug / 5-Aug-2008
| From Bracketing |  | Reference | Reproducibility | Combined | Expanded Uncertainty
| Average | Combined | Gas | (Type-B) | Overall
Setpoint | Ratio | uComparison | uReference | uReproducibility | uCandidate
ug/scm | Rmean | ug/scm | ug/scm | ug/scm | ug/scm | ug/scm | %
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Note on A13: Ed Brown: The "nominal" setpoint, used for identification, not for calculation

## Certification
Certification Report
|  |  | Candidate Location: | EPA/RTP Laboratories
|  |  | Candidate Device: | Tekran 6
|  |  | Candidate S/N: | 3009
|  |  | Reference ID: | Tekran 3
|  |  | Reference S/N: | 3039
|  |  | Certification Date(s): | 4-Aug / 5-Aug-2008
Reference Parameters |  |  |  | Candidate Generator Certification
Concentrations |  | Uncertainties |  | Certification Values |  | Certification Uncertainties
Setpoint | Certified | 2s | Relative | Setpoint | Certified | 1s | 2s | Relative
ug/scm | ug/scm | ug/scm | 2s, % | ug/scm | ug/scm | ug/scm | ug/scm | 2s, %
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Note on A12: Ed Brown: The "nominal" setpoint, used for identification, not for calculation
Note on A13: Ed Brown: The "nominal" setpoint, used for identification, not for calculation
Note on C12: Ed Brown: The "nominal" setpoint, used for identification, not for calculation
Note on E13: Ed Brown: The "nominal" setpoint, used for identification, not for calculation