## TSPB1A (2)
Total Suspended Particulate Emission Calculator
NoSouth Sway
Project No. | BB019012.0000
Scrubber (Boiler 7)
Run |  |  | 1 | 2 | 3
Date |  |  | 2014-08-04 | 2014-08-04 | 2014-08-04
Time |  |  | 1035-1240 | 1330-1535 | 1610-1815
Stack Flowrate (DSCFM) (A) |  |  | 26901.595 | 27302.3637 | 27439.1008
Stack Moisture (%) (B) |  |  | 9.3697 | 9.8413 | 9.5858
Meter Volume (DSCF) (C) |  |  | 108.4527 | 112.8736 | 114.7426
Stack Oxygen (%) (D) |  |  | 13.2167 | 12.2625 | 12.689
Fuel Usage (BTU/Hour) (E)
Fuel F-Factor -EPA M19 (dscf/mmBTU) (E1) |  |  | 9780 | 9780 | 9780
Correct to What % O2? |  | 7 | (D1)
Front-Half Analytical Results - P&N Wash + Filter (Grams)
Acetone Evaporated (ml):
Blank (G) | 95 | Sample (H) | 86 | 81 | 84
Net Particulate (grams):
|  | Filter (J) | 0.191 | 0.189 | 0.201
Acetone:
Blank (K) | 0.0007 | Sample (M) | 0.0047 | 0.0052 | 0.0033
|  | Blank corrected Sample (N) | 0.0040663 | 0.0046032 | 0.0026811
[6 empty] | AVERAGE | ALLOWABLE
Concentrations and Emissions:
Grains/scf (R) |  |  | 0.025152 | 0.023861 | 0.024764
Grains/dscf (S) |  |  | 0.027753 | 0.026466 | 0.02739
Pounds/hour (T) |  |  | 6.4015 | 6.1956 | 6.444 | 6.4227
Pounds/Million BTU by F-Factor(U1) |  |  | 0.10549 | 0.089486 | 0.097421 | 0.097466 | 0.309
Rinse Volume required |  |  | 248 | 248 | 248
Fraction of required volume- actual volume (V) |  |  | 2.8837 | 3.0617 | 2.9524
ADJUSTED Pounds/Million BTU by F-Factor(U1) |  |  | 0.30421 | 0.27398 | 0.28762 | 0.2886 | 0.309
Equations:
(N) = (M) - ((H)(K / G))
(R) = (S)(100 - B) / 100
(S) = ((J + N) / C) * 15.43 grains/gram
(T) = (A)((J + N) / C) * pound/453.59 grams* 60 min/hr
(U1) = (20.9 / (20.9-D))(E1 * (J + N)) / (C * 453.59)
(V) = 248 / H

## TSPB1A
Total Suspended Particulate Emission Calculator
NoSouth Sway
Project No. | BB019012.0000
Scrubber (Boiler 7)
Run |  |  | 1 | 2 | 3
Date |  |  | 2014-08-04 | 2014-08-04 | 2014-08-04
Time |  |  | 1035-1240 | 1330-1535 | 1610-1815
Stack Flowrate (DSCFM) (A) |  |  | 26901.595 | 27302.3637 | 27439.1008
Stack Moisture (%) (B) |  |  | 9.3697 | 9.8413 | 9.5858
Meter Volume (DSCF) (C) |  |  | 108.4527 | 112.8736 | 114.7426
Stack Oxygen (%) (D) |  |  | 13.2167 | 12.2625 | 12.689
Fuel Usage (BTU/Hour) (E)
Fuel F-Factor -EPA M19 (dscf/mmBTU) (E1) |  |  | 9780 | 9780 | 9780
Correct to What % O2? |  | 7 | (D1)
Front-Half Analytical Results - P&N Wash + Filter (Grams)
Acetone Evaporated (ml):
Blank (G) | 95 | Sample (H) | 86 | 81 | 84
Net Particulate (grams):
|  | Filter (J) | 0.191 | 0.189 | 0.201
Acetone:
Blank (K) | 0.0007 | Sample (M) | 0.0047 | 0.0052 | 0.0033
|  | Blank corrected Sample (N) | 0.0040663 | 0.0046032 | 0.0026811
[6 empty] | AVERAGE | ALLOWABLE
Concentrations and Emissions:
Grains/scf (R) |  |  | 0.025152 | 0.023861 | 0.024764
Grains/dscf (S) |  |  | 0.027753 | 0.026466 | 0.02739
Pounds/hour (T) |  |  | 6.4015 | 6.1956 | 6.444 | 6.4227
Pounds/Million BTU by F-Factor(U1) |  |  | 0.10549 | 0.089486 | 0.097421 | 0.097466 | 0.309
Equations:
(N) = (M) - ((H)(K / G)) | [5 empty] | revised- average wrong in original!
(R) = (S)(100 - B) / 100
(S) = ((J + N) / C) * 15.43 grains/gram
(T) = (A)((J + N) / C) * pound/453.59 grams* 60 min/hr
(U1) = (20.9 / (20.9-D))(E1 * (J + N)) / (C * 453.59)

## LBSHG
Mercury Emission Calculator
NoSouth Sway
Project No. | BB019012.0000
Scrubber (Boiler 7)
Mercury (EPA M29/Multi-metals)
Run |  | 1 | 2 | 3
Date |  | 2014-08-04 | 2014-08-04 | 2014-08-04
Time |  | 1035-1240 | 1330-1535 | 1610-1815
Meter Volume (DSCF)(A) |  | 108.4527 | 112.8736 | 114.7426
Stack Flow (DSCFM)(B) |  | 26901.595 | 27302.3637 | 27439.1008
Oxygen content (%) (I) |  | 13.2167 | 12.2625 | 12.689
Fuel F-Factor -EPA M19 (dscf/mmBTU) (C) |  | 9780 | 9780 | 9780
Front-Half Filter & Probe Rinse (Containers 1 & 3) (Analytical Fraction - 1B)
Sample (ug) (C1) |  | 0 | 0.022 | 0.161
Blank (ug) (D1) |  | 0 | 0 | 0
Net (micrograms) (E1) |  | 0 | 0.022 | 0.161
Grams/second (F1) |  | 0 | 8.8691e-08 | 6.4168e-07
Pounds/hour (G1) |  | 0 | 7.0403e-07 | 5.0937e-06
Impingers 1-3 Contents and Rinse (HNO3/H2O2) (Container 4) (Analytical Fraction - 2B)
Sample (ug) (C2) |  | 1.71 | 3.17 | 2.58
Blank (ug) (D2) |  | 0 | 0 | 0
Net (micrograms) (E2) |  | 1.71 | 3.17 | 2.58
Grams/second (F2) |  | 7.0694e-06 | 1.278e-05 | 1.0283e-05
Pounds/hour (G2) |  | 5.6117e-05 | 0.00010144 | 8.1625e-05
Impinger 4 Contents and Rinse (HNO3) (Container 5A) (Analytical Fraction - 3A)
Sample (ug) (C3) |  | 0.0053 | 0.007 | 0.007
Blank (ug) (D3) |  | 0 | 0 | 0
Net (micrograms) (E3) |  | 0.0053 | 0.007 | 0.007
Grams/second (F3) |  | 2.1911e-08 | 2.822e-08 | 2.7899e-08
Pounds/hour (G3) |  | 1.7393e-07 | 2.2401e-07 | 2.2146e-07
Impinger Content and Rinse (KMNO4) (Container 5B) (Analytical Fraction - 3B)
Sample (ug) (C4) |  | 0 | 0 | 0
Blank (ug) (D4) |  | 0 | 0 | 0
Net (micrograms) (E4) |  | 0 | 0 | 0
Grams/second (F4) |  | 0 | 0 | 0
Pounds/hour (G4) |  | 0 | 0 | 0
Impinger Rinse (HCl) (Container 5C) (Analytical Fraction - 3C)
Sample (ug) (C5) |  | 0.592 | 0.241 | 0.186
Blank (ug) (D5) |  | 0 | 0 | 0
Net (micrograms) (E5) |  | 0.592 | 0.241 | 0.186
Grams/second (F5) |  | 2.4474e-06 | 9.7157e-07 | 7.4132e-07
Pounds/hour (G5) |  | 1.9428e-05 | 7.7123e-06 | 5.8846e-06
Emissions |  |  |  |  | Averages | Allowable
Grams/second |  | 9.5387e-06 | 1.3868e-05 | 1.1694e-05 | 1.17e-05
Pounds/hour |  | 7.5718e-05 | 0.00011008 | 9.2825e-05 | 9.2876e-05
ug/dscm (H) |  | 0.75121 | 1.0761 | 0.90289 | 0.91007
ug/dscm @ 7%O2 (J) |  | 1.359 | 1.7318 | 1.5285 | 1.5398
Pounds/MMBtu (K) |  | 1.2478e-06 | 1.59e-06 | 1.4033e-06 | 1.4137e-06 | 2.2e-05
Equations:
E(n) = (C(n) - D(n))
F(n) = (E(n) / A) (B) (0.016667 minutes/second) (1E-6 grams/microgram)
G(n) = (F(n)) (3600 seconds/hour) (0.002205 pounds/gram)
H = ((E1 + E2 + E3 + E4 + E5)/(A))(35.31 cu.ft./dscm)
J = (H)((13.9/(20.9-I))
(K) = (20.9 / (20.9-I))((E1 + E2 + E3 + E4 + E5)/1000000) * C / (A * 453.59)

## CEMS
EMISSION CALCULATIONS
NoSouth Sway
Project No. | BB019012.0000
Scrubber (Boiler 7)
|  | Run | 1 | 2 | 3
|  | Date | 2013-12-04 | 2013-12-04 | 2013-12-04
|  | Time | 1035-1135 | 1330-1430 | 1610-1710
Stack Flowrate (DSCFM)(A) |  |  | 26901.595 | 27302.3637 | 27439.1008
Stack Moisture (%) (B) |  |  | 9.3697 | 9.8413 | 9.5858
Correct to | 3 | % O2 (C)
Fuel F-Factor -EPA M19 (dscf/mmBTU) (D) |  |  | 9780 | 9780 | 9780
[6 empty] | Averages | Allowables
Oxygen
Average Zero Response (%)(F) |  |  | 0 | 0 | 0
Average Upscale Response (%)(G) |  |  | 11.55 | 11.6 | 11.55
Upscale Standard Value (%)(H) |  |  | 11.5 | 11.5 | 11.5
%, dry (I), uncorrected |  |  | 13.2742 | 12.3692 | 12.7442
%, dry (J), corrected |  |  | 13.2167 | 12.2625 | 12.689 | 12.7227
Carbon Dioxide
Average Zero Response (%)(F) |  |  | 0.1 | 0.1 | 0.1
Average Upscale Response (%)(G) |  |  | 9 | 8.95 | 8.85
Upscale Standard Value (%)(H) |  |  | 9.06 | 9.06 | 9.06
%, dry (I), uncorrected |  |  | 7.4592 | 7.4292 | 7.2433
%, dry (J), corrected |  |  | 7.4915 | 7.5031 | 7.3964 | 7.4637
| Carbon Monoxide (CO) Molar Mass = (E) |  |  | 28.01
Average Zero Response (ppmv)(F) |  |  | 1.05 | 1.45 | 1.5
Average Upscale Response (ppmv)(G) |  |  | 47.15 | 47.2 | 47.4
Upscale Standard Value (ppmv)(H) |  |  | 48.1 | 48.1 | 48.1
ppmv, dry (I), uncorrected |  |  | 24.5792 | 24.4583 | 44.5325
ppmv, dry (K), corrected |  |  | 24.55 | 24.1902 | 45.0951 | 31.2784
pounds/hour (L) |  |  | 2.8695 | 2.8696 | 5.3762 | 3.7051
ppmv, dry at | 3 | % O2 (M) | 57.1947 | 50.1309 | 98.3073 | 68.5443 | 420
pounds/MMBtu (N) |  |  | 0.047289 | 0.041449 | 0.081281 | 0.056673
Equations: |  |  |  | Definitions:
(J or K)=(I-F)(H)/(G-F) |  |  |  | (A) = Instack Flow Rate as Measured By USEPA Methods 1 through 4 (See Stack Flow Calculations Sheets)
(L) = (A)(K, P1, or P3)(E)(60 min/hr)(1lb-mol/386.8 ft3)(10^-6) |  |  |  | (B) = Instack Moisture as Measured By USEPA Method 4. (See Stack Flow Calculations Sheets)
(M) = (20.9 - C) / (20.9 - JO2) (K, P1, or P3) |  |  |  | (D) = USEPA Method 19 F-Factor.
(N) = ((20.9/(20.9-JO2))(10E-06)(E)(D)(K, P1, or P3)/(386.8 ft3/lb-mol) |  |  |  | (E) = Molar mass of compound in lb/lb-mole.
(F) = Analyzer response to zero gas.
(G) = Analyzer response to span gas.
(H) = Span gas value.
(I or P) = Analyzer response to Stack gas.

## CEMS Run 001
DASYLab - V 7.00.00
WORKSHEET : O2,CO2,CO,voltmetr
Recording Date : 12/4/2013, 10:29:15 AM
Block Length : 1
Delta : 30 sec.
Number of Channels : 3
Date | Time | O2 [%] | CO2 [%] | CO [ppm]
2013-12-04 | 10:35 | 14.1 | 6.8 | 20.2
2013-12-04 | 10:36 | 14.1 | 6.8 | 20.3
2013-12-04 | 10:36 | 14.2 | 6.7 | 20.7
2013-12-04 | 10:37 | 14.2 | 6.7 | 21.2
2013-12-04 | 10:37 | 14.4 | 6.6 | 21.2
2013-12-04 | 10:38 | 14.5 | 6.5 | 21.7
2013-12-04 | 10:38 | 14.3 | 6.6 | 22.4
2013-12-04 | 10:39 | 14.2 | 6.7 | 22.6
2013-12-04 | 10:39 | 14.2 | 6.7 | 22.2
2013-12-04 | 10:40 | 14.3 | 6.7 | 21.7
2013-12-04 | 10:40 | 14.2 | 6.7 | 22.1
2013-12-04 | 10:41 | 14.2 | 6.7 | 22.2
2013-12-04 | 10:41 | 14 | 6.8 | 22.1
2013-12-04 | 10:42 | 13.9 | 6.9 | 21.7
2013-12-04 | 10:42 | 13.9 | 6.9 | 21.7
2013-12-04 | 10:43 | 14 | 6.9 | 21.7
2013-12-04 | 10:43 | 14 | 6.9 | 21.6
2013-12-04 | 10:44 | 13.8 | 7.1 | 20.6
2013-12-04 | 10:44 | 13.7 | 7.1 | 20.1
2013-12-04 | 10:45 | 13.6 | 7.2 | 19.7
2013-12-04 | 10:45 | 13.6 | 7.2 | 19.6
2013-12-04 | 10:46 | 13.6 | 7.2 | 19.2
2013-12-04 | 10:46 | 13.7 | 7.2 | 19.2
2013-12-04 | 10:47 | 13.7 | 7.2 | 19.2
2013-12-04 | 10:47 | 13.6 | 7.2 | 19.3
2013-12-04 | 10:48 | 13.5 | 7.3 | 19.2
2013-12-04 | 10:48 | 13.3 | 7.4 | 19.2
2013-12-04 | 10:49 | 13.3 | 7.4 | 19.2
2013-12-04 | 10:49 | 13.2 | 7.5 | 19.2
2013-12-04 | 10:50 | 13.2 | 7.5 | 19.2
2013-12-04 | 10:50 | 13.1 | 7.6 | 19.3
2013-12-04 | 10:51 | 13 | 7.7 | 19.8
2013-12-04 | 10:51 | 12.9 | 7.8 | 20.2
2013-12-04 | 10:52 | 12.8 | 7.8 | 20.3
2013-12-04 | 10:52 | 12.7 | 7.9 | 21
2013-12-04 | 10:53 | 12.7 | 7.9 | 21.6
2013-12-04 | 10:53 | 12.6 | 8 | 22.1
2013-12-04 | 10:54 | 12.5 | 8 | 22.9
2013-12-04 | 10:54 | 12.5 | 8 | 23.3
2013-12-04 | 10:55 | 12.4 | 8.2 | 24.6
2013-12-04 | 10:55 | 12.3 | 8.2 | 27.5
2013-12-04 | 10:56 | 12.4 | 8.1 | 29.9
2013-12-04 | 10:56 | 12.4 | 8.1 | 30.6
2013-12-04 | 10:57 | 12.4 | 8.1 | 30.6
2013-12-04 | 10:57 | 12.6 | 8 | 29
2013-12-04 | 10:58 | 12.7 | 7.9 | 26.8
2013-12-04 | 10:58 | 12.9 | 7.8 | 24.5
2013-12-04 | 10:59 | 13.1 | 7.7 | 23.1
2013-12-04 | 10:59 | 13.2 | 7.6 | 22.2
2013-12-04 | 11:00 | 13.2 | 7.6 | 22.2
2013-12-04 | 11:00 | 13.2 | 7.5 | 22.2
2013-12-04 | 11:01 | 13.1 | 7.6 | 22.2
2013-12-04 | 11:01 | 13 | 7.7 | 22.3
2013-12-04 | 11:02 | 12.8 | 7.8 | 23.1
2013-12-04 | 11:02 | 12.8 | 7.9 | 24.6
2013-12-04 | 11:03 | 12.8 | 7.8 | 26.4
2013-12-04 | 11:03 | 12.8 | 7.8 | 27.1
2013-12-04 | 11:04 | 12.8 | 7.8 | 27.7
2013-12-04 | 11:04 | 12.8 | 7.8 | 27.2
2013-12-04 | 11:05 | 12.8 | 7.8 | 27.2
2013-12-04 | 11:05 | 12.8 | 7.9 | 27.3
2013-12-04 | 11:06 | 12.8 | 7.9 | 28
2013-12-04 | 11:06 | 12.6 | 8 | 29.1
2013-12-04 | 11:07 | 12.5 | 8 | 31.5
2013-12-04 | 11:07 | 12.5 | 8 | 34.6
2013-12-04 | 11:08 | 12.7 | 7.9 | 36
2013-12-04 | 11:08 | 12.8 | 7.9 | 35.7
2013-12-04 | 11:09 | 12.8 | 7.8 | 34.9
2013-12-04 | 11:09 | 12.9 | 7.8 | 33.8
2013-12-04 | 11:10 | 12.9 | 7.7 | 31.7
2013-12-04 | 11:10 | 13 | 7.7 | 29.2
2013-12-04 | 11:11 | 13 | 7.7 | 26.8
2013-12-04 | 11:11 | 13.1 | 7.7 | 25.6
2013-12-04 | 11:12 | 13.1 | 7.6 | 25.2
2013-12-04 | 11:12 | 13.1 | 7.6 | 25.2
2013-12-04 | 11:13 | 13.1 | 7.6 | 25.2
2013-12-04 | 11:13 | 13.1 | 7.6 | 24.9
2013-12-04 | 11:14 | 13.2 | 7.6 | 24.1
2013-12-04 | 11:14 | 13.2 | 7.5 | 23.4
2013-12-04 | 11:15 | 13.3 | 7.5 | 22.6
2013-12-04 | 11:15 | 13.4 | 7.4 | 22.2
2013-12-04 | 11:16 | 13.4 | 7.4 | 22.3
2013-12-04 | 11:16 | 13.5 | 7.3 | 22.7
2013-12-04 | 11:17 | 13.6 | 7.3 | 22.7
2013-12-04 | 11:17 | 13.7 | 7.2 | 22.7
2013-12-04 | 11:18 | 13.8 | 7.1 | 22.7
2013-12-04 | 11:18 | 13.9 | 7 | 22.7
2013-12-04 | 11:19 | 13.9 | 7 | 23.1
2013-12-04 | 11:19 | 13.9 | 7 | 23.6
2013-12-04 | 11:20 | 13.8 | 7.1 | 23.8
2013-12-04 | 11:20 | 13.6 | 7.1 | 24.6
2013-12-04 | 11:21 | 13.6 | 7.2 | 25.8
2013-12-04 | 11:21 | 13.7 | 7.1 | 25.7
2013-12-04 | 11:22 | 13.7 | 7.1 | 25.5
2013-12-04 | 11:22 | 13.6 | 7.2 | 25.7
2013-12-04 | 11:23 | 13.7 | 7.2 | 26.2
2013-12-04 | 11:23 | 13.8 | 7.1 | 26.4
2013-12-04 | 11:24 | 13.8 | 7.1 | 25.7
2013-12-04 | 11:24 | 13.6 | 7.2 | 25.7
2013-12-04 | 11:25 | 13.5 | 7.2 | 25.7
2013-12-04 | 11:25 | 13.5 | 7.3 | 25.7
2013-12-04 | 11:26 | 13.3 | 7.4 | 26.2
2013-12-04 | 11:26 | 13.1 | 7.6 | 26.5
2013-12-04 | 11:27 | 12.8 | 7.8 | 27.9
2013-12-04 | 11:27 | 12.7 | 7.9 | 31.2
2013-12-04 | 11:28 | 12.7 | 7.9 | 33.2
2013-12-04 | 11:28 | 12.7 | 7.9 | 32.2
2013-12-04 | 11:29 | 12.8 | 7.8 | 30.7
2013-12-04 | 11:29 | 12.7 | 7.9 | 29.7
2013-12-04 | 11:30 | 12.8 | 7.8 | 28.7
2013-12-04 | 11:30 | 12.8 | 7.8 | 27.2
2013-12-04 | 11:31 | 12.9 | 7.7 | 25.7
2013-12-04 | 11:31 | 13 | 7.6 | 24.9
2013-12-04 | 11:32 | 13 | 7.6 | 24.4
2013-12-04 | 11:32 | 13.1 | 7.6 | 24.2
2013-12-04 | 11:33 | 13.1 | 7.6 | 24.2
2013-12-04 | 11:33 | 13.3 | 7.4 | 24.2
2013-12-04 | 11:34 | 13.5 | 7.3 | 23.9
2013-12-04 | 11:34 | 13.6 | 7.2 | 23.6
2013-12-04 | 11:35 | 13.7 | 7.1 | 23.2
AVERAGES |  | 13.2742 | 7.4592 | 24.5792

## CEMS Run 002
DASYLab - V 7.00.00
WORKSHEET : O2,CO2,CO,voltmetr
Recording Date : 12/4/2013, 1:25:47 PM
Block Length : 1
Delta : 30 sec.
Number of Channels : 3
Date | Time | O2 [%] | CO2 [%] | CO [ppm]
2013-12-04 | 13:30 | 12.1 | 7.7 | 25.5
2013-12-04 | 13:31 | 12.2 | 7.7 | 24.5
2013-12-04 | 13:31 | 12.4 | 7.6 | 23.5
2013-12-04 | 13:32 | 12.6 | 7.4 | 22.4
2013-12-04 | 13:32 | 12.7 | 7.3 | 21.1
2013-12-04 | 13:33 | 12.9 | 7.1 | 20.7
2013-12-04 | 13:33 | 13.1 | 7 | 20.7
2013-12-04 | 13:34 | 13.3 | 6.9 | 20.8
2013-12-04 | 13:34 | 13.4 | 6.8 | 21.2
2013-12-04 | 13:35 | 13.4 | 6.7 | 21.4
2013-12-04 | 13:35 | 13.3 | 6.8 | 21.2
2013-12-04 | 13:36 | 13 | 7 | 21.6
2013-12-04 | 13:36 | 12.8 | 7.2 | 22.1
2013-12-04 | 13:37 | 12.6 | 7.3 | 22.3
2013-12-04 | 13:37 | 12.5 | 7.4 | 22.2
2013-12-04 | 13:38 | 12.2 | 7.6 | 22.3
2013-12-04 | 13:38 | 12.1 | 7.7 | 23
2013-12-04 | 13:39 | 12.1 | 7.7 | 23.5
2013-12-04 | 13:39 | 11.9 | 7.9 | 24.6
2013-12-04 | 13:40 | 11.8 | 8 | 26.4
2013-12-04 | 13:40 | 11.7 | 8 | 27.8
2013-12-04 | 13:41 | 11.8 | 8 | 28.1
2013-12-04 | 13:41 | 11.9 | 7.9 | 26.9
2013-12-04 | 13:42 | 12 | 7.8 | 25.4
2013-12-04 | 13:42 | 12 | 7.8 | 24.7
2013-12-04 | 13:43 | 12.1 | 7.7 | 24.7
2013-12-04 | 13:43 | 12.2 | 7.6 | 24.3
2013-12-04 | 13:44 | 12.4 | 7.5 | 23.4
2013-12-04 | 13:44 | 12.5 | 7.5 | 22.4
2013-12-04 | 13:45 | 12.6 | 7.4 | 22.2
2013-12-04 | 13:45 | 12.8 | 7.2 | 22.2
2013-12-04 | 13:46 | 13.3 | 6.9 | 21.7
2013-12-04 | 13:46 | 13.3 | 6.8 | 21.7
2013-12-04 | 13:47 | 13.3 | 6.8 | 21.6
2013-12-04 | 13:47 | 13.2 | 6.8 | 20.9
2013-12-04 | 13:48 | 13.1 | 7 | 20.2
2013-12-04 | 13:48 | 12.8 | 7.2 | 19.7
2013-12-04 | 13:49 | 12.6 | 7.3 | 20.3
2013-12-04 | 13:49 | 12.6 | 7.3 | 20.8
2013-12-04 | 13:50 | 12.6 | 7.2 | 20.7
2013-12-04 | 13:50 | 12.4 | 7.3 | 20.3
2013-12-04 | 13:51 | 12.2 | 7.4 | 20.7
2013-12-04 | 13:51 | 12.1 | 7.5 | 22.5
2013-12-04 | 13:52 | 12.1 | 7.5 | 23.2
2013-12-04 | 13:52 | 12 | 7.6 | 23.2
2013-12-04 | 13:53 | 12 | 7.6 | 24.1
2013-12-04 | 13:53 | 12.1 | 7.5 | 25
2013-12-04 | 13:54 | 12.1 | 7.6 | 25.3
2013-12-04 | 13:54 | 11.9 | 7.7 | 25.8
2013-12-04 | 13:55 | 12 | 7.6 | 26.5
2013-12-04 | 13:55 | 12 | 7.6 | 25.7
2013-12-04 | 13:56 | 11.9 | 7.7 | 25
2013-12-04 | 13:56 | 11.9 | 7.7 | 24.5
2013-12-04 | 13:57 | 12 | 7.7 | 24.2
2013-12-04 | 13:57 | 12.2 | 7.5 | 23.4
2013-12-04 | 13:58 | 12.3 | 7.4 | 22.3
2013-12-04 | 13:58 | 12.3 | 7.4 | 21.4
2013-12-04 | 13:59 | 12.3 | 7.4 | 21.3
2013-12-04 | 13:59 | 12.4 | 7.3 | 21.7
2013-12-04 | 14:00 | 12.4 | 7.3 | 21.4
2013-12-04 | 14:00 | 12.5 | 7.3 | 21.2
2013-12-04 | 14:01 | 12.3 | 7.4 | 21.6
2013-12-04 | 14:01 | 12.2 | 7.5 | 22.6
2013-12-04 | 14:02 | 12.2 | 7.5 | 23.8
2013-12-04 | 14:02 | 12 | 7.6 | 24.8
2013-12-04 | 14:03 | 12.1 | 7.6 | 25.2
2013-12-04 | 14:03 | 12.2 | 7.5 | 25.2
2013-12-04 | 14:04 | 12.3 | 7.4 | 24.7
2013-12-04 | 14:04 | 12.4 | 7.3 | 23.3
2013-12-04 | 14:05 | 12.6 | 7.2 | 22.1
2013-12-04 | 14:05 | 12.8 | 7.1 | 21.4
2013-12-04 | 14:06 | 12.8 | 7.1 | 20.9
2013-12-04 | 14:06 | 13 | 7 | 20.7
2013-12-04 | 14:07 | 12.7 | 7.1 | 21
2013-12-04 | 14:07 | 12.7 | 7.2 | 21.8
2013-12-04 | 14:08 | 12.7 | 7.2 | 23
2013-12-04 | 14:08 | 12.6 | 7.2 | 23.5
2013-12-04 | 14:09 | 12.6 | 7.2 | 24
2013-12-04 | 14:09 | 12.5 | 7.3 | 25.1
2013-12-04 | 14:10 | 12.3 | 7.4 | 27.1
2013-12-04 | 14:10 | 12.1 | 7.5 | 29.3
2013-12-04 | 14:11 | 12.1 | 7.6 | 30.5
2013-12-04 | 14:11 | 12 | 7.7 | 30.7
2013-12-04 | 14:12 | 11.9 | 7.8 | 30.8
2013-12-04 | 14:12 | 11.9 | 7.8 | 31.8
2013-12-04 | 14:13 | 11.9 | 7.8 | 32.1
2013-12-04 | 14:13 | 12 | 7.7 | 30.3
2013-12-04 | 14:14 | 12.1 | 7.6 | 28.1
2013-12-04 | 14:14 | 12.2 | 7.5 | 26.5
2013-12-04 | 14:15 | 12.4 | 7.4 | 25.5
2013-12-04 | 14:15 | 12.6 | 7.2 | 24.4
2013-12-04 | 14:16 | 12.7 | 7.1 | 23.6
2013-12-04 | 14:16 | 12.8 | 7 | 23.8
2013-12-04 | 14:17 | 12.9 | 7 | 24.3
2013-12-04 | 14:17 | 12.7 | 7.1 | 24.4
2013-12-04 | 14:18 | 12.5 | 7.2 | 24.7
2013-12-04 | 14:18 | 12.5 | 7.3 | 26.2
2013-12-04 | 14:19 | 12.3 | 7.4 | 27.6
2013-12-04 | 14:19 | 12.2 | 7.5 | 29.1
2013-12-04 | 14:20 | 12 | 7.7 | 30.6
2013-12-04 | 14:20 | 12 | 7.7 | 31.2
2013-12-04 | 14:21 | 12.1 | 7.6 | 29.3
2013-12-04 | 14:21 | 12.1 | 7.6 | 26.4
2013-12-04 | 14:22 | 12.1 | 7.6 | 25.8
2013-12-04 | 14:22 | 12 | 7.7 | 26.6
2013-12-04 | 14:23 | 12 | 7.7 | 27.6
2013-12-04 | 14:23 | 12 | 7.7 | 28.9
2013-12-04 | 14:24 | 12.1 | 7.7 | 29
2013-12-04 | 14:24 | 12 | 7.7 | 28.4
2013-12-04 | 14:25 | 12.1 | 7.6 | 27.7
2013-12-04 | 14:25 | 12.2 | 7.6 | 27
2013-12-04 | 14:26 | 12.2 | 7.5 | 25.7
2013-12-04 | 14:26 | 12.3 | 7.5 | 24.9
2013-12-04 | 14:27 | 12.2 | 7.6 | 24.7
2013-12-04 | 14:27 | 12.1 | 7.6 | 25.9
2013-12-04 | 14:28 | 12.1 | 7.7 | 27.3
2013-12-04 | 14:28 | 12.1 | 7.6 | 27.5
2013-12-04 | 14:29 | 12.3 | 7.5 | 26.7
2013-12-04 | 14:29 | 12.5 | 7.3 | 25.2
2013-12-04 | 14:30 | 12.7 | 7.2 | 23.7
AVERAGES |  | 12.3692 | 7.4292 | 24.4583

## CEMS Run 003
DASYLab - V 7.00.00
WORKSHEET : O2,CO2,CO,voltmetr
Recording Date : 12/4/2013, 4:05:31 PM
Block Length : 1
Delta : 30 sec.
Number of Channels : 3
Date | Time | O2 [%] | CO2 [%] | CO [ppm]
2013-12-04 | 16:10 | 12.3 | 7.7 | 72.2
2013-12-04 | 16:11 | 12.4 | 7.6 | 72.8
2013-12-04 | 16:11 | 12.4 | 7.6 | 69
2013-12-04 | 16:12 | 12.4 | 7.6 | 67.8
2013-12-04 | 16:12 | 12.4 | 7.6 | 69.2
2013-12-04 | 16:13 | 12.4 | 7.6 | 71
2013-12-04 | 16:13 | 12.5 | 7.5 | 73.4
2013-12-04 | 16:14 | 12.5 | 7.5 | 72.2
2013-12-04 | 16:14 | 12.5 | 7.5 | 67.8
2013-12-04 | 16:15 | 12.7 | 7.4 | 62.7
2013-12-04 | 16:15 | 12.8 | 7.3 | 54.9
2013-12-04 | 16:16 | 12.8 | 7.3 | 47.5
2013-12-04 | 16:16 | 12.9 | 7.2 | 42.2
2013-12-04 | 16:17 | 12.9 | 7.2 | 39.8
2013-12-04 | 16:17 | 12.9 | 7.2 | 40.4
2013-12-04 | 16:18 | 13.1 | 7.1 | 40.5
2013-12-04 | 16:18 | 13.1 | 7 | 38
2013-12-04 | 16:19 | 13.1 | 7 | 36.3
2013-12-04 | 16:19 | 13.2 | 6.9 | 35.4
2013-12-04 | 16:20 | 13.2 | 6.9 | 34.4
2013-12-04 | 16:20 | 13.2 | 6.9 | 34
2013-12-04 | 16:21 | 13.1 | 7 | 35.1
2013-12-04 | 16:21 | 12.9 | 7.1 | 38.5
2013-12-04 | 16:22 | 12.8 | 7.2 | 42.3
2013-12-04 | 16:22 | 12.9 | 7.2 | 41.1
2013-12-04 | 16:23 | 12.7 | 7.3 | 37.6
2013-12-04 | 16:23 | 12.6 | 7.4 | 36.3
2013-12-04 | 16:24 | 12.4 | 7.6 | 38.7
2013-12-04 | 16:24 | 12.2 | 7.7 | 45
2013-12-04 | 16:25 | 12.2 | 7.7 | 50.1
2013-12-04 | 16:25 | 12.2 | 7.7 | 52.1
2013-12-04 | 16:26 | 12.1 | 7.7 | 54
2013-12-04 | 16:26 | 12 | 7.8 | 58
2013-12-04 | 16:27 | 12 | 7.8 | 62.4
2013-12-04 | 16:27 | 12 | 7.8 | 64
2013-12-04 | 16:28 | 11.9 | 8 | 64.4
2013-12-04 | 16:28 | 11.9 | 8 | 69.3
2013-12-04 | 16:29 | 11.9 | 7.9 | 71.6
2013-12-04 | 16:29 | 12.1 | 7.8 | 64.7
2013-12-04 | 16:30 | 12.3 | 7.7 | 54.5
2013-12-04 | 16:30 | 12.5 | 7.5 | 46.9
2013-12-04 | 16:31 | 12.7 | 7.4 | 41.8
2013-12-04 | 16:31 | 12.9 | 7.2 | 37.5
2013-12-04 | 16:32 | 13.2 | 7 | 34.7
2013-12-04 | 16:32 | 13.3 | 6.8 | 34.2
2013-12-04 | 16:33 | 13.3 | 6.8 | 34.2
2013-12-04 | 16:33 | 13.3 | 6.8 | 34.2
2013-12-04 | 16:34 | 13.2 | 6.9 | 34.5
2013-12-04 | 16:34 | 13.1 | 7 | 36.2
2013-12-04 | 16:35 | 13 | 7 | 41.3
2013-12-04 | 16:35 | 13 | 7 | 46.6
2013-12-04 | 16:36 | 12.9 | 7.1 | 47.8
2013-12-04 | 16:36 | 12.9 | 7.1 | 45.1
2013-12-04 | 16:37 | 12.9 | 7.2 | 40.9
2013-12-04 | 16:37 | 12.9 | 7.1 | 38.8
2013-12-04 | 16:38 | 12.9 | 7.1 | 40.2
2013-12-04 | 16:38 | 12.9 | 7.2 | 42.3
2013-12-04 | 16:39 | 12.8 | 7.2 | 41.7
2013-12-04 | 16:39 | 12.8 | 7.2 | 39.7
2013-12-04 | 16:40 | 12.7 | 7.2 | 39.2
2013-12-04 | 16:40 | 12.6 | 7.4 | 42
2013-12-04 | 16:41 | 12.5 | 7.4 | 44
2013-12-04 | 16:41 | 12.3 | 7.6 | 45.4
2013-12-04 | 16:42 | 12.2 | 7.7 | 51.1
2013-12-04 | 16:42 | 12.2 | 7.6 | 57.5
2013-12-04 | 16:43 | 12.3 | 7.6 | 57.5
2013-12-04 | 16:43 | 12.2 | 7.7 | 52.1
2013-12-04 | 16:44 | 12.2 | 7.7 | 47.3
2013-12-04 | 16:44 | 12.3 | 7.6 | 44.6
2013-12-04 | 16:45 | 12.2 | 7.7 | 45.8
2013-12-04 | 16:45 | 12.3 | 7.6 | 50.6
2013-12-04 | 16:46 | 12.5 | 7.5 | 52.5
2013-12-04 | 16:46 | 12.5 | 7.4 | 49.8
2013-12-04 | 16:47 | 12.4 | 7.5 | 47.3
2013-12-04 | 16:47 | 12.3 | 7.6 | 50.1
2013-12-04 | 16:48 | 12.3 | 7.6 | 55
2013-12-04 | 16:48 | 12.4 | 7.6 | 54.5
2013-12-04 | 16:49 | 12.5 | 7.5 | 49.6
2013-12-04 | 16:49 | 12.7 | 7.3 | 44.6
2013-12-04 | 16:50 | 13 | 7.1 | 40.8
2013-12-04 | 16:50 | 13 | 7.1 | 37.6
2013-12-04 | 16:51 | 13.1 | 7 | 35
2013-12-04 | 16:51 | 13.2 | 6.9 | 33.1
2013-12-04 | 16:52 | 13.3 | 6.8 | 32.6
2013-12-04 | 16:52 | 13.4 | 6.7 | 33.1
2013-12-04 | 16:53 | 13.6 | 6.6 | 33
2013-12-04 | 16:53 | 13.7 | 6.5 | 32.6
2013-12-04 | 16:54 | 13.7 | 6.5 | 32.6
2013-12-04 | 16:54 | 13.6 | 6.5 | 32.5
2013-12-04 | 16:55 | 13.6 | 6.5 | 32.8
2013-12-04 | 16:55 | 13.6 | 6.6 | 33.6
2013-12-04 | 16:56 | 13.5 | 6.6 | 34.5
2013-12-04 | 16:56 | 13.4 | 6.7 | 34.3
2013-12-04 | 16:57 | 13.3 | 6.8 | 33.6
2013-12-04 | 16:57 | 13 | 7 | 33.3
2013-12-04 | 16:58 | 12.8 | 7.1 | 35.8
2013-12-04 | 16:58 | 12.6 | 7.3 | 40.1
2013-12-04 | 16:59 | 12.5 | 7.4 | 44.3
2013-12-04 | 16:59 | 12.3 | 7.6 | 47
2013-12-04 | 17:00 | 12.2 | 7.6 | 48.5
2013-12-04 | 17:00 | 12.2 | 7.7 | 48.6
2013-12-04 | 17:01 | 12.2 | 7.7 | 48
2013-12-04 | 17:01 | 12.2 | 7.6 | 46.6
2013-12-04 | 17:02 | 12.2 | 7.6 | 45.1
2013-12-04 | 17:02 | 12.3 | 7.5 | 43.6
2013-12-04 | 17:03 | 12.6 | 7.3 | 42.6
2013-12-04 | 17:03 | 12.7 | 7.1 | 40.5
2013-12-04 | 17:04 | 12.9 | 7 | 37.5
2013-12-04 | 17:04 | 13 | 6.9 | 34.7
2013-12-04 | 17:05 | 13 | 6.9 | 33.4
2013-12-04 | 17:05 | 13.1 | 6.8 | 32.9
2013-12-04 | 17:06 | 13 | 6.9 | 32.6
2013-12-04 | 17:06 | 12.9 | 6.9 | 32.8
2013-12-04 | 17:07 | 13 | 6.9 | 32.4
2013-12-04 | 17:07 | 13.1 | 6.8 | 30.4
2013-12-04 | 17:08 | 13.1 | 6.8 | 28.8
2013-12-04 | 17:08 | 13.2 | 6.8 | 28.1
2013-12-04 | 17:09 | 13.4 | 6.6 | 28.1
2013-12-04 | 17:09 | 13.4 | 6.6 | 28.1
2013-12-04 | 17:10 | 13.4 | 6.5 | 28.1
AVERAGES |  | 12.7442 | 7.2433 | 44.5325

## TSP HG-ISO Run1
ISOKINETIC SAMPLE CALCULATIONS
| Client Name: | NoSouth Sway |  |  | Traverse | Delta-P | Delta-H | Gas Meter Temps. |  | Stack | Square
| Jobsite: | Atlas, MA |  |  | Point No. | ("H2O) | ("H2O) | Inlet | Outlet | Temp. | Root of
| Emission Source(s): | Scrubber (Boiler 7) | [5 empty] | (deg. F) | (deg. F) | (deg. F) | Delta-P
| ARCADIS Project No.: | BB019012.0000 |  |  | 1 | 0.05 | 2.25 | 69 | 69 | 105 | 0.22361
| Date: | 2014-08-04 |  |  | 2 | 0.05 | 2.25 | 68 | 68 | 105 | 0.22361
| Sampling Location: | Outlet Stack |  |  | 3 | 0.055 | 2.475 | 68 | 68 | 106 | 0.23452
| Run No.: | 1 |  |  | 4 | 0.055 | 2.475 | 68 | 68 | 106 | 0.23452
| Run Times: | 1035-1240 |  | PM/Metals | 5 | 0.055 | 2.475 | 70 | 70 | 108 | 0.23452
[5 empty] | 6 | 0.06 | 2.7 | 71 | 71 | 107 | 0.24495
| Input Data: |  |  |  | 7 | 0.06 | 2.7 | 72 | 72 | 107 | 0.24495
Beginning Meter Setting (Cubic feet): |  |  |  | 795.235 | 8 | 0.06 | 2.7 | 74 | 74 | 107 | 0.24495
Ending Meter Setting (Cubic feet): |  |  |  | 905.012 | 9 | 0.06 | 2.7 | 75 | 75 | 106 | 0.24495
Total Metered Volume (Cubic Feet) [Vm]: |  |  |  | 109.777 | 10 | 0.055 | 2.475 | 75 | 75 | 105 | 0.23452
Total Water Caught (grams) [Wc]: |  |  |  | 237.7 | 11 | 0.045 | 2.025 | 75 | 75 | 105 | 0.21213
Stack Static Pressure ("H2O) [Ps]: |  |  |  | -0.085 | 12 | 0.045 | 2.025 | 76 | 76 | 106 | 0.21213
Barometric Pressure ("Hg) [Pb]: |  |  |  | 30.05 | 13 | 0.05 | 2.25 | 78 | 78 | 106 | 0.22361
Carbon Dioxide (%) [CO2]: |  |  |  | 7.4915 | 14 | 0.055 | 2.475 | 77 | 77 | 106 | 0.23452
Oxygen (%) [O2]: |  |  |  | 13.2167 | 15 | 0.055 | 2.475 | 77 | 77 | 106 | 0.23452
Nozzle Diameter (inches) [Dn]: |  |  |  | 0.502 | 16 | 0.055 | 2.475 | 78 | 78 | 107 | 0.23452
Pitot Tube Factor [Cp]: |  |  |  | 0.83 | 17 | 0.055 | 2.475 | 77 | 77 | 107 | 0.23452
Meter Correction Factor [Y]: |  |  |  | 0.99 | 18 | 0.06 | 2.7 | 78 | 78 | 106 | 0.24495
Stack Inside Diameter (in) [ID]: |  |  |  | 85 | 19 | 0.06 | 2.7 | 77 | 77 | 106 | 0.24495
Stack Cross-Section Area (Square Feet) [CSA]: |  |  |  | 39.4063 | 20 | 0.065 | 2.925 | 77 | 77 | 107 | 0.25495
Total Sample Time (minutes) [min]: |  |  |  | 120 | 21 | 0.06 | 2.7 | 78 | 78 | 106 | 0.24495
[5 empty] | 22 | 0.05 | 2.25 | 78 | 78 | 106 | 0.22361
[5 empty] | 23 | 0.045 | 2.025 | 77 | 77 | 106 | 0.21213
[5 empty] | 24 | 0.045 | 2.025 | 77 | 77 | 105 | 0.21213
| Output Data: |  |  |  | 25
Metered Volume (dry cubic feet @stp)[Vms]: |  |  |  | 108.4527 |  |  | [H] |  | [Tm] | [Ts] | [dP]
Water Volume (cubic feet @ stp) [Vws]: |  |  |  | 11.2123
Moisture (%) [Bws]: |  |  |  | 9.3697 | Avgs. | 0.054375 | 2.4469 |  | 74.5833 | 106.125 | 0.23286
Dry Molecular Weight (lb/lb-mol) [MWd]: |  |  |  | 29.7273
Wet Molecular Weight (lb/lb-mol) [MWw]: |  |  |  | 28.6285 |  |  | K factor | 45
Absolute Stack Pressure ("Hg) [Pa]: |  |  |  | 30.0437
Gas Velocities (feet/second) | [6 empty] | Impinger Weights
| Stack - Actual [Vsa]: |  |  | 13.4052 |  |  | Final Wt | Initial Wt | Gain
| Stack - @ stp [Vss]: |  |  | 12.5542 |  |  | 581.2 | 541.3 | 39.9
| Nozzle - Actual [Vn]: |  |  | 12.0982 |  |  | 882.4 | 767.8 | 114.6
Stack Flowrate (cubic feet/ minute) | [6 empty] | 843.7 | 797.9 | 45.8
| Actual [ACFM]: |  |  | 31694.9836 |  |  | 660.4 | 650.8 | 9.6
| @ stp [SCFM]: |  |  | 29682.7867 |  |  | 776.3 | 771.1 | 5.2
| Dry @ stp [DSCFM]: |  |  | 26901.595 |  |  | 781.9 | 779.9 | 2
Sample Isokinetics (%) [I]: |  |  |  | 96.3678 |  |  | 933.9 | 913.3 | 20.6
[7 empty] | Water Caught (gms) |  | 237.7
This Spreadsheet Uses The Following Equations/Calculations:
Vms = (Vm)(Y)((A+460)/B)((H / 13.6)+(Pb))/(Tm + 460) | [7 empty] | STP = Standard Conditions = | 68 | Deg F (A)
Vws = (Wc) / 21.2 | [8 empty] | 29.92 | "Hg (B)
Bws = (100)(Vws)/(Vws + Vms)
MWd = ((0.44)(CO2))+((0.32)(O2))+((0.28)(100-(CO2 + O2)))
MWw = (0.18)(Bws)+((100 - Bws)(MWd) / 100)
Pa = Pb + ((Ps) / 13.6) | [6 empty] | ARCADIS US Inc., Branchburg NJ
Vsa = (85.49)(Cp)(dP)((Ts+460)/((MWw)(Pa)))^0.5
Vss = ((A+460)/B)(Vsa)(Pa)/((Ts) + 460)
ACFM = (Vsa)(CSA)(60)
SCFM = (Vss)(CSA)(60)
DSCFM = (SCFM)(100 - (Bws) / 100)
Vn = (Vms + Vws)/(((Dn / 2)^2)(3.14 / 144)(min)(60))
CSA = ID2/183.3465
I = (Vn / Vss)(100)

## TSP HG-ISO Run2
ISOKINETIC SAMPLE CALCULATIONS
| Client Name: | NoSouth Sway |  |  | Traverse | Delta-P | Delta-H | Gas Meter Temps. |  | Stack | Square
| Jobsite: | Atlas, MA |  |  | Point No. | ("H2O) | ("H2O) | Inlet | Outlet | Temp. | Root of
| Emission Source(s): | Scrubber (Boiler 7) | [5 empty] | (deg. F) | (deg. F) | (deg. F) | Delta-P
| ARCADIS Project No.: | BB019012.0000 |  |  | 1 | 0.05 | 2.3 | 74 | 74 | 107 | 0.22361
| Date: | 2014-08-04 |  |  | 2 | 0.05 | 2.3 | 75 | 75 | 107 | 0.22361
| Sampling Location: | Outlet Stack |  |  | 3 | 0.065 | 2.99 | 75 | 75 | 109 | 0.25495
| Run No.: | 2 |  |  | 4 | 0.06 | 2.76 | 75 | 75 | 108 | 0.24495
| Run Times: | 1330-1535 |  | PM/Metals | 5 | 0.06 | 2.76 | 75 | 75 | 108 | 0.24495
[5 empty] | 6 | 0.06 | 2.76 | 76 | 76 | 109 | 0.24495
| Input Data: |  |  |  | 7 | 0.055 | 2.53 | 76 | 76 | 108 | 0.23452
Beginning Meter Setting (Cubic feet): |  |  |  | 905.465 | 8 | 0.055 | 2.53 | 76 | 76 | 108 | 0.23452
Ending Meter Setting (Cubic feet): |  |  |  | 1020.374 | 9 | 0.06 | 2.76 | 77 | 77 | 108 | 0.24495
Total Metered Volume (Cubic Feet) [Vm]: |  |  |  | 114.909 | 10 | 0.055 | 2.53 | 77 | 77 | 108 | 0.23452
Total Water Caught (grams) [Wc]: |  |  |  | 261.2 | 11 | 0.045 | 2.07 | 77 | 77 | 108 | 0.21213
Stack Static Pressure ("H2O) [Ps]: |  |  |  | -0.085 | 12 | 0.045 | 2.07 | 77 | 77 | 108 | 0.21213
Barometric Pressure ("Hg) [Pb]: |  |  |  | 30.02 | 13 | 0.05 | 2.3 | 78 | 78 | 108 | 0.22361
Carbon Dioxide (%) [CO2]: |  |  |  | 7.5031 | 14 | 0.055 | 2.53 | 78 | 78 | 108 | 0.23452
Oxygen (%) [O2]: |  |  |  | 12.2625 | 15 | 0.055 | 2.53 | 78 | 78 | 107 | 0.23452
Nozzle Diameter (inches) [Dn]: |  |  |  | 0.502 | 16 | 0.06 | 2.76 | 78 | 78 | 108 | 0.24495
Pitot Tube Factor [Cp]: |  |  |  | 0.83 | 17 | 0.06 | 2.76 | 79 | 79 | 108 | 0.24495
Meter Correction Factor [Y]: |  |  |  | 0.99 | 18 | 0.065 | 2.99 | 79 | 79 | 108 | 0.25495
Stack Inside Diameter (in) [ID]: |  |  |  | 85 | 19 | 0.065 | 2.99 | 79 | 79 | 107 | 0.25495
Stack Cross-Section Area (Square Feet) [CSA]: |  |  |  | 39.4063 | 20 | 0.065 | 2.99 | 79 | 79 | 108 | 0.25495
Total Sample Time (minutes) [min]: |  |  |  | 120 | 21 | 0.065 | 2.99 | 79 | 79 | 108 | 0.25495
[5 empty] | 22 | 0.06 | 2.76 | 79 | 79 | 108 | 0.24495
[5 empty] | 23 | 0.05 | 2.3 | 80 | 80 | 107 | 0.22361
[5 empty] | 24 | 0.05 | 2.3 | 80 | 80 | 107 | 0.22361
| Output Data: |  |  |  | 25 | [5 empty] | 0
Metered Volume (dry cubic feet @stp)[Vms]: |  |  |  | 112.8736 |  |  | [H] |  | [Tm] | [Ts] | [dP]
Water Volume (cubic feet @ stp) [Vws]: |  |  |  | 12.3208
Moisture (%) [Bws]: |  |  |  | 9.8413 | Avgs. | 0.056667 | 2.6067 |  | 77.3333 | 107.8333 | 0.23768
Dry Molecular Weight (lb/lb-mol) [MWd]: |  |  |  | 29.691
Wet Molecular Weight (lb/lb-mol) [MWw]: |  |  |  | 28.5404 |  |  | K factor | 46
Absolute Stack Pressure ("Hg) [Pa]: |  |  |  | 30.0137
Gas Velocities (feet/second) | [6 empty] | Impinger Weights
| Stack - Actual [Vsa]: |  |  | 13.7311 |  |  | Final Wt | Initial Wt | Gain
| Stack - @ stp [Vss]: |  |  | 12.8078 |  |  | 593.1 | 541.4 | 51.7
| Nozzle - Actual [Vn]: |  |  | 12.6572 |  |  | 820.1 | 688.8 | 131.3
Stack Flowrate (cubic feet/ minute) | [6 empty] | 847.9 | 804.5 | 43.4
| Actual [ACFM]: |  |  | 32465.4103 |  |  | 659.8 | 651.4 | 8.4
| @ stp [SCFM]: |  |  | 30282.5615 |  |  | 734.2 | 728.5 | 5.7
| Dry @ stp [DSCFM]: |  |  | 27302.3637 |  |  | 771.6 | 769.4 | 2.2
Sample Isokinetics (%) [I]: |  |  |  | 98.8238 |  |  | 927.5 | 909 | 18.5
[7 empty] | Water Caught (gms) |  | 261.2
This Spreadsheet Uses The Following Equations/Calculations:
Vms = (Vm)(Y)((A+460)/B)((H / 13.6)+(Pb))/(Tm + 460) | [7 empty] | STP = Standard Conditions = | 68 | Deg F (A)
Vws = (Wc) / 21.2 | [8 empty] | 29.92 | "Hg (B)
Bws = (100)(Vws)/(Vws + Vms)
MWd = ((0.44)(CO2))+((0.32)(O2))+((0.28)(100-(CO2 + O2)))
MWw = (0.18)(Bws)+((100 - Bws)(MWd) / 100) | [6 empty] | ARCADIS US Inc., Branchburg NJ
Pa = Pb + ((Ps) / 13.6)
Vsa = (85.49)(Cp)(dP)((Ts+460)/((MWw)(Pa)))^0.5
Vss = ((A+460)/B)(Vsa)(Pa)/((Ts) + 460)
ACFM = (Vsa)(CSA)(60)
SCFM = (Vss)(CSA)(60)
DSCFM = (SCFM)(100 - (Bws) / 100)
Vn = (Vms + Vws)/(((Dn / 2)^2)(3.14 / 144)(min)(60))
CSA = ID2/183.3465
I = (Vn / Vss)(100)

## TSP HG-ISO Run3
ISOKINETIC SAMPLE CALCULATIONS
| Client Name: | NoSouth Sway |  |  | Traverse | Delta-P | Delta-H | Gas Meter Temps. |  | Stack | Square
| Jobsite: | Atlas, MA |  |  | Point No. | ("H2O) | ("H2O) | Inlet | Outlet | Temp. | Root of
| Emission Source(s): | Scrubber (Boiler 7) | [5 empty] | (deg. F) | (deg. F) | (deg. F) | Delta-P
| ARCADIS Project No.: | BB019012.0000 |  |  | 1 | 0.05 | 2.3 | 74 | 74 | 108 | 0.22361
| Date: | 2014-08-04 |  |  | 2 | 0.05 | 2.3 | 74 | 74 | 108 | 0.22361
| Sampling Location: | Outlet Stack |  |  | 3 | 0.055 | 2.53 | 75 | 75 | 107 | 0.23452
| Run No.: | 3 |  |  | 4 | 0.06 | 2.76 | 75 | 75 | 107 | 0.24495
| Run Times: | 1610-1815 |  | PM/Metals | 5 | 0.06 | 2.76 | 76 | 76 | 108 | 0.24495
[5 empty] | 6 | 0.06 | 2.76 | 77 | 77 | 107 | 0.24495
| Input Data: |  |  |  | 7 | 0.065 | 2.99 | 77 | 77 | 107 | 0.25495
Beginning Meter Setting (Cubic feet): |  |  |  | 20.633 | 8 | 0.065 | 2.99 | 78 | 78 | 108 | 0.25495
Ending Meter Setting (Cubic feet): |  |  |  | 137.768 | 9 | 0.06 | 2.76 | 79 | 79 | 108 | 0.24495
Total Metered Volume (Cubic Feet) [Vm]: |  |  |  | 117.135 | 10 | 0.055 | 2.53 | 80 | 80 | 106 | 0.23452
Total Water Caught (grams) [Wc]: |  |  |  | 257.9 | 11 | 0.05 | 2.3 | 81 | 81 | 107 | 0.22361
Stack Static Pressure ("H2O) [Ps]: |  |  |  | -0.08 | 12 | 0.05 | 2.3 | 81 | 81 | 107 | 0.22361
Barometric Pressure ("Hg) [Pb]: |  |  |  | 30.02 | 13 | 0.055 | 2.53 | 81 | 81 | 107 | 0.23452
Carbon Dioxide (%) [CO2]: |  |  |  | 7.3964 | 14 | 0.055 | 2.53 | 81 | 81 | 108 | 0.23452
Oxygen (%) [O2]: |  |  |  | 12.689 | 15 | 0.06 | 2.76 | 81 | 81 | 109 | 0.24495
Nozzle Diameter (inches) [Dn]: |  |  |  | 0.502 | 16 | 0.06 | 2.76 | 81 | 81 | 107 | 0.24495
Pitot Tube Factor [Cp]: |  |  |  | 0.83 | 17 | 0.055 | 2.53 | 81 | 81 | 108 | 0.23452
Meter Correction Factor [Y]: |  |  |  | 0.99 | 18 | 0.055 | 2.53 | 80 | 80 | 109 | 0.23452
Stack Inside Diameter (in) [ID]: |  |  |  | 85 | 19 | 0.06 | 2.76 | 80 | 80 | 108 | 0.24495
Stack Cross-Section Area (Square Feet) [CSA]: |  |  |  | 39.4063 | 20 | 0.065 | 2.99 | 80 | 80 | 107 | 0.25495
Total Sample Time (minutes) [min]: |  |  |  | 120 | 21 | 0.065 | 2.99 | 80 | 80 | 107 | 0.25495
[5 empty] | 22 | 0.055 | 2.53 | 80 | 80 | 107 | 0.23452
[5 empty] | 23 | 0.05 | 2.3 | 80 | 80 | 107 | 0.22361
[5 empty] | 24 | 0.05 | 2.3 | 80 | 80 | 107 | 0.22361
| Output Data: |  |  |  | 25 | [5 empty] | 0
Metered Volume (dry cubic feet @stp)[Vms]: |  |  |  | 114.7426 |  |  | [H] |  | [Tm] | [Ts] | [dP]
Water Volume (cubic feet @ stp) [Vws]: |  |  |  | 12.1651
Moisture (%) [Bws]: |  |  |  | 9.5858 | Avgs. | 0.056875 | 2.6162 |  | 78.8333 | 107.4583 | 0.23824
Dry Molecular Weight (lb/lb-mol) [MWd]: |  |  |  | 29.691
Wet Molecular Weight (lb/lb-mol) [MWw]: |  |  |  | 28.5703 |  |  | K factor | 46
Absolute Stack Pressure ("Hg) [Pa]: |  |  |  | 30.0141
Gas Velocities (feet/second) | [6 empty] | Impinger Weights
| Stack - Actual [Vsa]: |  |  | 13.7516 |  |  | Final Wt | Initial Wt | Gain
| Stack - @ stp [Vss]: |  |  | 12.8356 |  |  | 607.1 | 542.6 | 64.5
| Nozzle - Actual [Vn]: |  |  | 12.8304 |  |  | 896.3 | 770.9 | 125.4
Stack Flowrate (cubic feet/ minute) | [6 empty] | 756.1 | 717.8 | 38.3
| Actual [ACFM]: |  |  | 32513.9114 |  |  | 648 | 641 | 7
| @ stp [SCFM]: |  |  | 30348.2151 |  |  | 774.4 | 770.4 | 4
| Dry @ stp [DSCFM]: |  |  | 27439.1008 |  |  | 773 | 772.7 | 0.3
Sample Isokinetics (%) [I]: |  |  |  | 99.9595 |  |  | 942.9 | 924.5 | 18.4
[7 empty] | Water Caught (gms) |  | 257.9
This Spreadsheet Uses The Following Equations/Calculations:
Vms = (Vm)(Y)((A+460)/B)((H / 13.6)+(Pb))/(Tm + 460) | [7 empty] | STP = Standard Conditions = | 68 | Deg F (A)
Vws = (Wc) / 21.2 | [8 empty] | 29.92 | "Hg (B)
Bws = (100)(Vws)/(Vws + Vms)
MWd = ((0.44)(CO2))+((0.32)(O2))+((0.28)(100-(CO2 + O2)))
MWw = (0.18)(Bws)+((100 - Bws)(MWd) / 100) | [6 empty] | ARCADIS US Inc., Branchburg NJ
Pa = Pb + ((Ps) / 13.6)
Vsa = (85.49)(Cp)(dP)((Ts+460)/((MWw)(Pa)))^0.5
Vss = ((A+460)/B)(Vsa)(Pa)/((Ts) + 460)
ACFM = (Vsa)(CSA)(60)
SCFM = (Vss)(CSA)(60)
DSCFM = (SCFM)(100 - (Bws) / 100)
Vn = (Vms + Vws)/(((Dn / 2)^2)(3.14 / 144)(min)(60))
CSA = ID2/183.3465
I = (Vn / Vss)(100)

## TSP HG- stack cond
RUN | 1 | 2 | 3
DATE | 2014-08-04 | 2014-08-04 | 2014-08-04
TIME | 1035-1240 | 1330-1535 | 1610-1815
STACK INSIDE DIAMETER (in) | 85 | 85 | 85
STACK CROSS SECTION (sq ft) | 39.4063 | 39.4063 | 39.4063
BAROMETRIC PRESSURE | 30.05 | 30.02 | 30.02
AVG. STACK TEMP (oF) | 106.125 | 107.8333 | 107.4583
STACK PRESSURE ("H2O-gage) | -0.085 | -0.085 | -0.08
MOISTURE (% vol) | 9.3697 | 9.8413 | 9.5858
O2 (% vol) | 13.2167 | 12.2625 | 12.689
CO2 (%vol) | 7.4915 | 7.5031 | 7.3964
N2 (% vol by difference) | 79.2918 | 80.2344 | 79.9146
AVG. ACTUAL VELOCITY (ft/sec) | 13.4052 | 13.7311 | 13.7516
ACTUAL FLOW RATE (acfm) | 31690 | 32470 | 32510
STD FLOW RATE (scfm) | 29680 | 30280 | 30350
DRY STD FLOW RATE (dscfm) | 26900 | 27300 | 27440
Standard Conditions are 68 °F, 29.92" Hg

## lbmmbtu Hg Coal
Laboratory Coal Analysis- hg
(Confirmation of their calculations)
Concentration of Hg (ppbw) (A) | 64
HHV Coal (Btu/Lb Coal) (B) | 13793
Emission Rate Hg (Lb Hg/MMBtu) (C ) | 4.64e-06
(A) = reported concentration of mercury in coal (lb Hg/1,000,000,000 lb Coal)
(B) = reported HHV
(C ) = (A)/(B)/1000
(lb Hg/100000000 lb-coal) / (Btu/lb-coal) x 1000000/1000000000

## Data Sheet 1
Plant: | NoSouth Sway | [12 empty] | Optional Fields
Data required in ¶ 8.5 of EPA Method 5 and presented in Figure 5-3 of Method 5.
Sampling Location: |  | Scrubber (Boiler 7) |  |  |  | Method: | PM/Metals | Checks |  |  | Pre | Mid | Post
Run No.: |  | 1 |  | Testing Operator: | TWB |  |  | Vacuum (in Hg): |  |  | 15 |  | 7
Run Date: |  | 2014-08-04 |  | Job Number: | BB019012.0000 |  |  | Leak Check Total Volume (ft3):
Equipment Identification | [7 empty] | Leak Rate (cfm): |  |  | 0.001 |  | 0.001
Dry Gas Meter ID: |  | ES245 |  | Calibration |  |  |  | Pitot (Yes/No): |  |  | Yes |  | Yes
Reagent Box ID: |  | ES246 |  | Meter Box Correction Factor, Y: |  |  | 0.99 | Nozzle (Yes/No): |  |  | Yes |  | Yes
Unbilical Line ID: |  | 1 |  | Meter Box pressure differential, ΔH@ (in H2O): |  |  | 1.63 | Stack Thermocouple (Yes/No): |  |  | Yes |  | Yes
Stack Thermocouple ID: |  | TypeK-456-9 |  | Pitot tube coefficient, Cp: |  |  | 0.83
Tedlar Bag IDs: |  |  |  | Nozzle diameter, Dn (in): |  |  | 0.502
Orsat Pump ID: | [7 empty] | Moisture in: |  | Initial | Final | Micro-manometer Used (Y/N)? |  |  | No
Probe / Pitot ID: |  | 128 |  | Ambient Conditions |  |  |  | Impinger 1: |  | 541.3 | 581.2 | Sensitivity (in. H2O)
Nozzle ID: |  | SetES2987-.5 |  | Pbar (in. Hg): |  | 30.05 |  | Impinger 2: |  | 767.8 | 882.4
|  |  |  | Temperature (°F): |  |  |  | Impinger 3: |  | 797.9 | 843.7
Filters |  |  |  | Pg (in. H2O): |  | -0.085 |  | Impinger 4: |  | 650.8 | 660.4
Filter No. 1 ID: |  | QtzP1245-7 |  | Concentrations |  |  |  | Impinger 5: |  | 771.1 | 776.3
Filter No. 2 ID: |  |  |  | O2 Concentration (%): |  | 13.2167 |  | Impinger 6: |  | 779.9 | 781.9
Filter No. 3 ID: |  |  |  | CO2 Concentration (%): |  | 7.4915 |  | Silica Gel. (gm): |  | 913.3 | 933.9 | Impinger Water Weight (g), Vlc: |  |  | 237.7
Traverse Point No. | Sampling Time (min) |  | Clock Time 24 hr | Dry Gas Meter Reading (Vm,ft3) | Velocity Head (ΔP in H2O) | Orfice Pressure Differential (ΔH, in H20) |  | Stack Temperature °F | Probe Temperature °F | Filter Temperature °F |  | Final exit Temperature °F | Dry Gas Meter Temperature °F |  | Pump Vacuum (in. Hg) | Sample Rate
| Begin | End |  |  |  | Desired | Actual |  |  | First | Second |  | Inlet | Outlet
1 | 0 | 5 | 10:35 | 795.235 | 0.05 | 2.25 | 2.25 | 105 | 245 | 248 |  | 33 | 69 | 69 | 2
2 | 5 | 10 | 10:40 | 799.809 | 0.05 | 2.25 | 2.25 | 105 | 245 | 248 |  | 33 | 68 | 68 | 2
3 | 10 | 15 | 10:45 | 804.383 | 0.055 | 2.475 | 2.475 | 106 | 243 | 250 |  | 34 | 68 | 68 | 2
4 | 15 | 20 | 10:50 | 808.957 | 0.055 | 2.475 | 2.475 | 106 | 239 | 250 |  | 36 | 68 | 68 | 2
5 | 20 | 25 | 10:55 | 813.531 | 0.055 | 2.475 | 2.475 | 108 | 248 | 258 |  | 38 | 70 | 70 | 2
6 | 25 | 30 | 11:00 | 818.105 | 0.06 | 2.7 | 2.7 | 107 | 247 | 252 |  | 40 | 71 | 71 | 2
7 | 30 | 35 | 11:05 | 822.679 | 0.06 | 2.7 | 2.7 | 107 | 248 | 253 |  | 40 | 72 | 72 | 3
8 | 35 | 40 | 11:10 | 827.253 | 0.06 | 2.7 | 2.7 | 107 | 252 | 255 |  | 40 | 74 | 74 | 3
9 | 40 | 45 | 11:15 | 831.827 | 0.06 | 2.7 | 2.7 | 106 | 251 | 249 |  | 41 | 75 | 75 | 3
10 | 45 | 50 | 11:20 | 836.401 | 0.055 | 2.475 | 2.475 | 105 | 253 | 252 |  | 41 | 75 | 75 | 3
11 | 50 | 55 | 11:25 | 840.975 | 0.045 | 2.025 | 2.025 | 105 | 255 | 254 |  | 41 | 75 | 75 | 3
12 | 55 | 60 | 11:30 | 845.549 | 0.045 | 2.025 | 2.025 | 106 | 258 | 255 |  | 40 | 76 | 76 | 3
13 | 60 | 65 | 11:35 | 850.123 | 0.05 | 2.25 | 2.25 | 106 | 258 | 251 |  | 40 | 78 | 78 | 3
14 | 65 | 70 | 11:40 | 854.697 | 0.055 | 2.475 | 2.475 | 106 | 255 | 255 |  | 41 | 77 | 77 | 3
15 | 70 | 75 | 11:45 | 859.271 | 0.055 | 2.475 | 2.475 | 106 | 250 | 253 |  | 36 | 77 | 77 | 3
16 | 75 | 80 | 11:50 | 863.845 | 0.055 | 2.475 | 2.475 | 107 | 252 | 252 |  | 37 | 78 | 78 | 3
17 | 80 | 85 | 11:55 | 868.419 | 0.055 | 2.475 | 2.475 | 107 | 252 | 247 |  | 38 | 77 | 77 | 3
18 | 85 | 90 | 12:00 | 872.993 | 0.06 | 2.7 | 2.7 | 106 | 257 | 254 |  | 39 | 78 | 78 | 3
19 | 90 | 95 | 12:05 | 877.567 | 0.06 | 2.7 | 2.7 | 106 | 256 | 252 |  | 39 | 77 | 77 | 4
20 | 95 | 100 | 12:10 | 882.141 | 0.065 | 2.925 | 2.925 | 107 | 252 | 251 |  | 39 | 77 | 77 | 4
21 | 100 | 105 | 12:15 | 886.715 | 0.06 | 2.7 | 2.7 | 106 | 256 | 251 |  | 39 | 78 | 78 | 4
22 | 105 | 110 | 12:20 | 891.289 | 0.05 | 2.25 | 2.25 | 106 | 251 | 253 |  | 39 | 78 | 78 | 4
23 | 110 | 115 | 12:25 | 895.863 | 0.045 | 2.025 | 2.025 | 106 | 251 | 254 |  | 40 | 77 | 77 | 4
24 | 115 | 120 | 12:30 | 900.437 | 0.045 | 2.025 | 2.025 | 105 | 254 | 254 |  | 41 | 77 | 77 | 4
| 120 |  | 12:35 | 905.012
Note on B2: Enter the name of the facility.
Note on B4: Enter the sampling location name, such as inlet, stack, ESP inlet, scrubber outlet, etc.
Note on H4: The method to be used to determine the analyte emissions.
Note on B5: Enter the run number.
Note on F5: Enter the name of the operator performing the run.
Note on J5: Record the vacuum at which the Pre and Post sampling train leak checks were performed.
Note on B6: Enter the date that the run was performed.
Note on J6: Enter the value representing the total volume of all leak checks.
Note on B7: Enter an appropriate identification value for each piece of equipment.
Note on J7: Record the Pre and Post test sampling train leak check rates. For Method 5, the Post-test leak rate must be less than or equal to 0.02 acfm.
Note on B8: The Dry Gas Meter is a specialized flow meter, used to measure the volume of the extracted stack gas. THe Gas Meter measures a defined volume, regardless of the pressurized quantity or quality of the gas flowing through the meter. Temperature and pressure compensation must be made to measure the standard volume of gas moving through the meter.The ID is needed
Note on J8: Record the Pre and Post-test leak check results, as applicable.
Note on B9: Reagent Box, optional. A reagent box is a container that is sealable (for custody purposes) and is used to transport multiple recovered samples.
Note on F9: This is the dry gas meter correction coefficient, gamma, of an isokinetic sampling train meterbox (such as used for Method 5 sampling). The value should be between 0.95 and 1.05.
Note on J9: Nozzle inspections for dents, nicks, etc.
Note on B10: Sample gas transport line from the sample box to the meterbox, optional. The umbilical usually consists of bundled tubing, thermocouple, electrical lines, etc., used to control the probe and sample box filter temperatures.
Note on F10: This is the orifice coefficient of an isokinetic sampling train meterbox (such as used for Method 5 sampling).
Note on J10: Record the Pre and Post-test results of the thermocouple check, as applicable.
Note on B11: The ID of the thermocouple device used for monitoring the stack gas temperature. The ID is necessary for calibration documentation purposes.
Note on F11: This is the pitot tube coefficient and is usually 0.84 for S-type pitots and 0.99 for standard pitots.
Note on B12: The ID of a Tedlar bag, if used. Optional.
Note on F12: This is the nozzle diameter, inches.
Note on B13: The ID of the pump used for filling a Tedlar bag, for example. Optional.
Note on K13: Select the units of measure for the impinger moisture readings.
Note on N13: Was a micro-manometer used for flow rate measurement?
Note on J14: Enter the initial and final moisture readings for impinger 1 OR enter a value in the "Impinger Water Weight" cell.
Note on N14: Enter sensitivity of micro-manometer if used.
Note on B15: The nozzle ID.
Note on F15: Input the barometric pressure of the sampling location. If the pressure is at sea level, remember to adjust the pressure for the elevation of the location above sea level. A 0.1 inches Hg decrease for every 100 feet of elevation is typically used.
Note on J15: Enter the initial and final moisture readings for impinger 2.
Note on J16: Enter the initial and final moisture readings for impinger 3.
Note on F17: Record the static pressure, inches of water, of the sampling location.
Note on J17: Enter the initial and final moisture readings for impinger 4.
Note on B18: For particulate sampling, the filters used have a unique identification. Record the number here.
Note on J18: Enter the initial and final moisture readings for impinger 5.
Note on B19: If two filters are used, record the second ID here.
Note on F19: Record the oxygen percent of the gas stream tested.
Note on J19: Enter the initial and final moisture readings for impinger 6.
Note on B20: If three filters are used, record the third filter ID here.
Note on F20: Record the carbon dioxide percent of the gas stream tested.
Note on J20: Enter the initial and final moisture readings for silica gel.
Note on N20: The volume or mass of the reagents increase in the impingers of an isokinetic sampling train is input here. This value is calculated from the impinger and silica gel weights entered above or can be entered directly.
Note on B22: The sampling point label, such as A1, A-1, D-2, etc.
Note on E22: This is the actual clock time at the start of sampling at a point. Necessary for non-isokinetic methods, see instructions.
Note on F22: This is the Dry Gas Meter volume reading at the beginning of the sampling at a point. This means that the final volume reading is recorded in a row without a Point label and no other recorded point data. Sometimes the sampling data is recorded at the end of sampling at a point which would require that the first volume reading is recorded without any other sampling data. Necessary for non-isokinetic methods, see instructions.
Note on G22: The velocity pressure (delta p) is expressed in inches of water.
Note on J22: Stack temperature is the temperature of the effluent gas at the sampling point and is expressed as degrees Fahrenheit.
Note on K22: Temperature of the sampling probe, degrees F. Necessary for non-isokinetic methods, see instructions.
Note on L22: Temperature of the filter box or compartment.
Note on N22: Enter the final impinger or condenser exit temperature. Necessary for non-isokinetic methods, see instructions.
Note on Q22: Vacuum of the sampling pump, inches mercury. Necessary for non-isokinetic methods, see instructions.
Note on R22: Sample flow rate.
Note on C23: The cumulative sampling time that sampling at the sample point was started, in minutes. Port changes DO NOT reset the time to zero (0).
Note on D23: The cumulative sampling time that sampling at the sample point was ended and is the begin time plus the sampling time per point.
Note on H23: Orifice pressure setting required for sampling isokinetically.
Note on I23: Orifice pressure actually sampled or reached, inches water
Note on L23: Filter inlet temperature or box temperature.
Note on M23: Filter outlet temperature only.
Note on O23: Dry gas meter inlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.
Note on P23: Dry gas meter outlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.

## Data Sheet 2
Plant: | NoSouth Sway | [12 empty] | Optional Fields
Data required in ¶ 8.5 of EPA Method 5 and presented in Figure 5-3 of Method 5.
Sampling Location: |  | Scrubber (Boiler 7) |  |  |  | Method: | PM/Metals | Checks |  |  | Pre | Mid | Post
Run No.: |  | 2 |  | Testing Operator: | TWB |  |  | Vacuum (in Hg): |  |  | 15 |  | 7
Run Date: |  | 2014-08-04 |  | Job Number: | BB019012.0000 |  |  | Leak Check Total Volume (ft3):
Equipment Identification | [7 empty] | Leak Rate (cfm): |  |  | 0.001 |  | 0.001
Dry Gas Meter ID: |  | ES245 |  | Calibration |  |  |  | Pitot (Yes/No): |  |  | Yes |  | Yes
Reagent Box ID: |  | ES246 |  | Meter Box Correction Factor, Y: |  |  | 0.99 | Nozzle (Yes/No): |  |  | Yes |  | Yes
Unbilical Line ID: |  | 1 |  | Meter Box pressure differential, ΔH@ (in H2O): |  |  | 1.63 | Stack Thermocouple (Yes/No): |  |  | Yes |  | Yes
Stack Thermocouple ID: |  | TypeK-456-9 |  | Pitot tube coefficient, Cp: |  |  | 0.83
Tedlar Bag IDs: |  |  |  | Nozzle diameter, Dn (in): |  |  | 0.502
Orsat Pump ID: | [7 empty] | Moisture in: |  | Initial | Final | Micro-manometer Used (Y/N)? |  |  | No
Probe / Pitot ID: |  | 128 |  | Ambient Conditions |  |  |  | Impinger 1: |  | 541.4 | 593.1 | Sensitivity (in. H2O)
Nozzle ID: |  | SetES2987-.5 |  | Pbar (in. Hg): |  | 30.02 |  | Impinger 2: |  | 688.8 | 820.1
|  |  |  | Temperature (°F): |  |  |  | Impinger 3: |  | 804.5 | 847.9
Filters |  |  |  | Pg (in. H2O): |  | -0.085 |  | Impinger 4: |  | 651.4 | 659.8
Filter No. 1 ID: |  | QtzP1245-8 |  | Concentrations |  |  |  | Impinger 5: |  | 728.5 | 734.2
Filter No. 2 ID: |  |  |  | O2 Concentration (%): |  | 12.2625 |  | Impinger 6: |  | 769.4 | 771.6
Filter No. 3 ID: |  |  |  | CO2 Concentration (%): |  | 7.5031 |  | Silica Gel. (gm): |  | 909 | 927.5 | Impinger Water Weight (g), Vlc: |  |  | 261.2
Traverse Point No. | Sampling Time (min) |  | Clock Time 24 hr | Dry Gas Meter Reading (Vm,ft3) | Velocity Head (ΔP in H2O) | Orfice Pressure Differential (ΔH, in H20) |  | Stack Temperature °F | Probe Temperature °F | Filter Temperature °F |  | Final exit Temperature °F | Dry Gas Meter Temperature °F |  | Pump Vacuum (in. Hg) | Sample Rate
| Begin | End |  |  |  | Desired | Actual |  |  | First | Second |  | Inlet | Outlet
1 | 0 | 5 | 13:30:00 AM | 795.235 | 0.05 | 2.3 | 2.3 | 107 | 245 | 248 |  | 40 | 74 | 74 | 2
2 | 5 | 10 | 13:35:00 AM | 804.616 | 0.05 | 2.3 | 2.3 | 107 | 245 | 248 |  | 41 | 75 | 75 | 2
3 | 10 | 15 | 13:40:00 AM | 813.997 | 0.065 | 2.99 | 2.99 | 109 | 245 | 248 |  | 41 | 75 | 75 | 2
4 | 15 | 20 | 13:45:00 AM | 823.378 | 0.06 | 2.76 | 2.76 | 108 | 245 | 248 |  | 41 | 75 | 75 | 2
5 | 20 | 25 | 13:50:00 AM | 832.759 | 0.06 | 2.76 | 2.76 | 108 | 245 | 248 |  | 40 | 75 | 75 | 2
6 | 25 | 30 | 13:55:00 AM | 842.14 | 0.06 | 2.76 | 2.76 | 109 | 245 | 248 |  | 40 | 76 | 76 | 2
7 | 30 | 35 | 14:00:00 AM | 851.521 | 0.055 | 2.53 | 2.53 | 108 | 245 | 248 |  | 41 | 76 | 76 | 2
8 | 35 | 40 | 14:05:00 AM | 860.902 | 0.055 | 2.53 | 2.53 | 108 | 245 | 248 |  | 36 | 76 | 76 | 2
9 | 40 | 45 | 14:10:00 AM | 870.283 | 0.06 | 2.76 | 2.76 | 108 | 245 | 248 |  | 37 | 77 | 77 | 2
10 | 45 | 50 | 14:15:00 AM | 879.664 | 0.055 | 2.53 | 2.53 | 108 | 245 | 248 |  | 38 | 77 | 77 | 2
11 | 50 | 55 | 14:20:00 AM | 889.045 | 0.045 | 2.07 | 2.07 | 108 | 245 | 248 |  | 39 | 77 | 77 | 2
12 | 55 | 60 | 14:25:00 AM | 898.426 | 0.045 | 2.07 | 2.07 | 108 | 245 | 248 |  | 39 | 77 | 77 | 2
13 | 60 | 65 | 14:30:00 AM | 907.807 | 0.05 | 2.3 | 2.3 | 108 | 245 | 248 |  | 39 | 78 | 78 | 2
14 | 65 | 70 | 14:35:00 AM | 917.188 | 0.055 | 2.53 | 2.53 | 108 | 245 | 248 |  | 39 | 78 | 78 | 2
15 | 70 | 75 | 14:40:00 AM | 926.569 | 0.055 | 2.53 | 2.53 | 107 | 245 | 248 |  | 39 | 78 | 78 | 2
16 | 75 | 80 | 14:45:00 AM | 935.95 | 0.06 | 2.76 | 2.76 | 108 | 245 | 248 |  | 40 | 78 | 78 | 2
17 | 80 | 85 | 14:50:00 AM | 945.331 | 0.06 | 2.76 | 2.76 | 108 | 245 | 248 |  | 41 | 79 | 79 | 2
18 | 85 | 90 | 14:55:00 AM | 954.712 | 0.065 | 2.99 | 2.99 | 108 | 245 | 248 |  | 33 | 79 | 79 | 2
19 | 90 | 95 | 15:00:00 AM | 964.093 | 0.065 | 2.99 | 2.99 | 107 | 245 | 248 |  | 33 | 79 | 79 | 2
20 | 95 | 100 | 15:05:00 AM | 973.474 | 0.065 | 2.99 | 2.99 | 108 | 245 | 248 |  | 34 | 79 | 79 | 2
21 | 100 | 105 | 15:10:00 AM | 982.855 | 0.065 | 2.99 | 2.99 | 108 | 245 | 248 |  | 36 | 79 | 79 | 2
22 | 105 | 110 | 15:15:00 AM | 992.236 | 0.06 | 2.76 | 2.76 | 108 | 245 | 248 |  | 38 | 79 | 79 | 2
23 | 110 | 115 | 15:20:00 AM | 1001.617 | 0.05 | 2.3 | 2.3 | 107 | 245 | 248 |  | 40 | 80 | 80 | 2
24 | 115 | 120 | 15:25:00 AM | 1010.998 | 0.05 | 2.3 | 2.3 | 107 | 245 | 248 |  | 40 | 80 | 80 | 2
| 120 |  |  | 1020.374
Note on B2: Enter the name of the facility.
Note on B4: Enter the sampling location name, such as inlet, stack, ESP inlet, scrubber outlet, etc.
Note on H4: The method to be used to determine the analyte emissions is input here.
Note on B5: Enter the run number.
Note on F5: Enter the name of the operator performing the run.
Note on J5: Record the vacuum at which the Pre and Post sampling train leak checks were performed.
Note on B6: Enter the date that the run was performed.
Note on J6: Enter the value representing the total volume of all leak checks.
Note on B7: Enter an appropriate identification value for each piece of equipment.
Note on J7: Record the Pre and Post test sampling train leak check rates. For Method 5, the Post-test leak rate must be less than or equal to 0.02 acfm.
Note on B8: The Dry Gas Meter is a specialized flow meter, used to measure the volume of the extracted stack gas. THe Gas Meter measures a defined volume, regardless of the pressurized quantity or quality of the gas flowing through the meter. Temperature and pressure compensation must be made to measure the standard volume of gas moving through the meter.The ID is needed
Note on J8: Record the Pre and Post-test leak check results, as applicable.
Note on B9: Reagent Box, optional. A reagent box is a container that is sealable (for custody purposes) and is used to transport multiple recovered samples.
Note on F9: This is the dry gas meter correction coefficient, gamma, of an isokinetic sampling train meterbox (such as used for Method 5 sampling). The value should be between 0.95 and 1.05.
Note on J9: Nozzle inspections for dents, nicks, etc.
Note on B10: Sample gas transport line from the sample box to the meterbox, optional. The umbilical usually consists of bundled tubing, thermocouple, electrical lines, etc., used to control the probe and sample box filter temperatures.
Note on F10: This is the orifice coefficient of an isokinetic sampling train meterbox (such as used for Method 5 sampling).
Note on J10: Record the Pre and Post-test results of the thermocouple check, as applicable.
Note on B11: The ID of the thermocouple device used for monitoring the stack gas temperature. The ID is necessary for calibration documentation purposes
Note on F11: This is the pitot tube coefficient and is usually 0.84 for S-type pitots and 0.99 for standard pitots.
Note on B12: The ID of a Tedlar bag, if used. Optional.
Note on F12: This is the nozzle diameter, inches.
Note on B13: The ID of the pump used for filling a Tedlar bag, for example. Optional.
Note on K13: Select the units of measure for the impinger moisture readings.
Note on N13: Was a micro-manometer used for flow rate measurement?
Note on J14: Enter the initial and final moisture readings for impinger 1.
Note on N14: Enter sensitivity of micro-manometer if used.
Note on B15: The nozzle ID.
Note on F15: Input the barometric pressure of the sampling location. If the pressure is at sea level, remember to adjust the pressure for the elevation of the location above sea level. A 0.1 inches Hg decrease for every 100 feet of elevation is typically used.
Note on J15: Enter the initial and final moisture readings for impinger 2.
Note on J16: Enter the initial and final moisture readings for impinger 3.
Note on F17: Record the static pressure, inches of water, of the sampling location.
Note on J17: Enter the initial and final moisture readings for impinger 4.
Note on B18: For particulate sampling, the filters used have a unique identification. Record the number here.
Note on J18: Enter the initial and final moisture readings for impinger 5.
Note on B19: If two filters are used, record the second ID here.
Note on F19: Record the oxygen percent of the gas stream tested.
Note on J19: Enter the initial and final moisture readings for impinger 6.
Note on B20: If three filters are used, record the third filter ID here.
Note on F20: Record the carbon dioxide percent of the gas stream tested.
Note on J20: Enter the initial and final moisture readings for silica gel.
Note on N20: The volume or mass of the reagents increase in the impingers of an isokinetic sampling train is input here. This value is calculated from the impinger and silica gel weights entered above or can be entered directly.
Note on B22: The sampling point label, such as A1, A-1, D-2, etc.
Note on E22: This is the actual clock time at the start of sampling at a point. Necessary for non-isokinetic methods, see instructions.
Note on F22: This is the Dry Gas Meter volume reading at the beginning of the sampling at a point. This means that the final volume reading is recorded in a row without a Point label and no other recorded point data. Sometimes the sampling data is recorded at the end of sampling at a point which would require that the first volume reading is recorded without any other sampling data. Necessary for non-isokinetic methods, see instructions.
Note on G22: The velocity pressure (delta p) is expressed in inches of water.
Note on J22: Stack temperature is the temperature of the effluent gas at the sampling point and is expressed as degrees Fahrenheit.
Note on K22: Temperature of the sampling probe, degrees F. Necessary for non-isokinetic methods, see instructions.
Note on L22: Temperature of the filter box or compartment.
Note on N22: Enter the final impinger or condenser exit temperature. Necessary for non-isokinetic methods, see instructions.
Note on Q22: Vacuum of the sampling pump, inches mercury. Necessary for non-isokinetic methods, see instructions.
Note on R22: Sample flow rate.
Note on C23: The cumulative sampling time that sampling at the sample point was started, in minutes. Port changes DO NOT reset the time to zero (0).
Note on D23: The cumulative sampling time that sampling at the sample point was ended and is the begin time plus the sampling time per point.
Note on H23: Orifice pressure setting required for sampling isokinetically.
Note on I23: Orifice pressure actually sampled or reached, inches water
Note on L23: Filter inlet temperature or box temperature.
Note on M23: Filter outlet temperature only.
Note on O23: Dry gas meter inlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.
Note on P23: Dry gas meter outlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.

## Data Sheet 3
Plant: | NoSouth Sway | [12 empty] | Optional Fields
Data required in ¶ 8.5 of EPA Method 5 and presented in Figure 5-3 of Method 5.
Sampling Location: |  | Scrubber (Boiler 7) |  |  |  | Method: | PM/Metals | Checks |  |  | Pre | Mid | Post
Run No.: |  | 3 |  | Testing Operator: | TWB |  |  | Vacuum (in Hg): |  |  | 15 |  | 7
Run Date: |  | 2014-08-04 |  | Job Number: | BB019012.0000 |  |  | Leak Check Total Volume (ft3):
Equipment Identification | [7 empty] | Leak Rate (cfm): |  |  | 0.001 |  | 0.001
Dry Gas Meter ID: |  | ES245 |  | Calibration |  |  |  | Pitot (Yes/No): |  |  | Yes |  | Yes
Reagent Box ID: |  | ES246 |  | Meter Box Correction Factor, Y: |  |  | 0.99 | Nozzle (Yes/No): |  |  | Yes |  | Yes
Unbilical Line ID: |  | 1 |  | Meter Box pressure differential, ΔH@ (in H2O): |  |  | 1.68 | Stack Thermocouple (Yes/No): |  |  | Yes |  | Yes
Stack Thermocouple ID: |  | TypeK-456-9 |  | Pitot tube coefficient, Cp: |  |  | 0.83
Tedlar Bag IDs: |  |  |  | Nozzle diameter, Dn (in): |  |  | 0.502
Orsat Pump ID: | [7 empty] | Moisture in: |  | Initial | Final | Micro-manometer Used (Y/N)? |  |  | No
Probe / Pitot ID: |  | 128 |  | Ambient Conditions |  |  |  | Impinger 1: |  | 542.6 | 607.1 | Sensitivity (in. H2O)
Nozzle ID: |  | SetES2987-.5 |  | Pbar (in. Hg): |  | 30.02 |  | Impinger 2: |  | 770.9 | 896.3
|  |  |  | Temperature (°F): |  |  |  | Impinger 3: |  | 717.8 | 756.1
Filters |  |  |  | Pg (in. H2O): |  | -0.08 |  | Impinger 4: |  | 641 | 648
Filter No. 1 ID: |  | QtzP1245-7 |  | Concentrations |  |  |  | Impinger 5: |  | 770.4 | 774.4
Filter No. 2 ID: |  |  |  | O2 Concentration (%): |  | 12.689 |  | Impinger 6: |  | 772.7 | 773
Filter No. 3 ID: |  |  |  | CO2 Concentration (%): |  | 7.3964 |  | Silica Gel. (gm): |  | 924.5 | 942.9 | Impinger Water Weight (g), Vlc: |  |  | 257.9
Traverse Point No. | Sampling Time (min) |  | Clock Time 24 hr | Dry Gas Meter Reading (Vm,ft3) | Velocity Head (ΔP in H2O) | Orfice Pressure Differential (ΔH, in H20) |  | Stack Temperature °F | Probe Temperature °F | Filter Temperature °F |  | Final exit Temperature °F | Dry Gas Meter Temperature °F |  | Pump Vacuum (in. Hg) | Sample Rate
| Begin | End |  |  |  | Desired | Actual |  |  | First | Second |  | Inlet | Outlet
1 | 0 | 5 | 16:10 | 20.633 | 0.05 | 2.3 | 2.3 | 108 | 245 | 248 |  | 33 | 74 | 74 | 2
2 | 5 | 10 | 16:15 | 25.514 | 0.05 | 2.3 | 2.3 | 108 | 245 | 248 |  | 33 | 74 | 74 | 3
3 | 10 | 15 | 16:20 | 30.395 | 0.055 | 2.53 | 2.53 | 107 | 245 | 248 |  | 33 | 75 | 75 | 4
4 | 15 | 20 | 16:25 | 35.276 | 0.06 | 2.76 | 2.76 | 107 | 245 | 248 |  | 33 | 75 | 75 | 5
5 | 20 | 25 | 16:30 | 40.157 | 0.06 | 2.76 | 2.76 | 108 | 245 | 248 |  | 33 | 76 | 76 | 6
6 | 25 | 30 | 16:35 | 45.038 | 0.06 | 2.76 | 2.76 | 107 | 245 | 248 |  | 33 | 77 | 77 | 7
7 | 30 | 35 | 16:40 | 49.919 | 0.065 | 2.99 | 2.99 | 107 | 245 | 248 |  | 33 | 77 | 77 | 8
8 | 35 | 40 | 16:45 | 54.8 | 0.065 | 2.99 | 2.99 | 108 | 245 | 248 |  | 33 | 78 | 78 | 9
9 | 40 | 45 | 16:50 | 59.681 | 0.06 | 2.76 | 2.76 | 108 | 245 | 248 |  | 33 | 79 | 79 | 10
10 | 45 | 50 | 16:55 | 64.562 | 0.055 | 2.53 | 2.53 | 106 | 245 | 248 |  | 33 | 80 | 80 | 11
11 | 50 | 55 | 17:00 | 69.443 | 0.05 | 2.3 | 2.3 | 107 | 245 | 248 |  | 33 | 81 | 81 | 12
12 | 55 | 60 | 17:05 | 74.324 | 0.05 | 2.3 | 2.3 | 107 | 245 | 248 |  | 33 | 81 | 81 | 13
13 | 60 | 65 | 17:10 | 79.205 | 0.055 | 2.53 | 2.53 | 107 | 245 | 248 |  | 33 | 81 | 81 | 14
14 | 65 | 70 | 17:15 | 84.086 | 0.055 | 2.53 | 2.53 | 108 | 245 | 248 |  | 33 | 81 | 81 | 15
15 | 70 | 75 | 17:20 | 88.967 | 0.06 | 2.76 | 2.76 | 109 | 245 | 248 |  | 33 | 81 | 81 | 16
16 | 75 | 80 | 17:25 | 93.848 | 0.06 | 2.76 | 2.76 | 107 | 245 | 248 |  | 33 | 81 | 81 | 17
17 | 80 | 85 | 17:30 | 98.729 | 0.055 | 2.53 | 2.53 | 108 | 245 | 248 |  | 33 | 81 | 81 | 18
18 | 85 | 90 | 17:35 | 103.61 | 0.055 | 2.53 | 2.53 | 109 | 245 | 248 |  | 33 | 80 | 80 | 19
19 | 90 | 95 | 17:40 | 108.491 | 0.06 | 2.76 | 2.76 | 108 | 245 | 248 |  | 33 | 80 | 80 | 20
20 | 95 | 100 | 17:45 | 113.372 | 0.065 | 2.99 | 2.99 | 107 | 245 | 248 |  | 33 | 80 | 80 | 21
21 | 100 | 105 | 17:50 | 118.253 | 0.065 | 2.99 | 2.99 | 107 | 245 | 248 |  | 33 | 80 | 80 | 22
22 | 105 | 110 | 17:55 | 123.134 | 0.055 | 2.53 | 2.53 | 107 | 245 | 248 |  | 33 | 80 | 80 | 23
23 | 110 | 115 | 18:00 | 128.015 | 0.05 | 2.3 | 2.3 | 107 | 245 | 248 |  | 33 | 80 | 80 | 24
24 | 115 | 120 | 18:05 | 132.896 | 0.05 | 2.3 | 2.3 | 107 | 245 | 248 |  | 33 | 80 | 80 | 25
| 120 |  |  | 137.768
Note on B2: Enter the name of the facility.
Note on B4: Enter the sampling location name, such as inlet, stack, ESP inlet, scrubber outlet, etc.
Note on H4: The method to be used to determine the analyte emissions is input here.
Note on B5: Enter the run number.
Note on F5: Enter the name of the operator performing the run.
Note on J5: Record the vacuum at which the Pre and Post sampling train leak checks were performed.
Note on B6: Enter the date that the run was performed.
Note on J6: Enter the value representing the total volume of all leak checks.
Note on B7: Enter an appropriate identification value for each piece of equipment.
Note on J7: Record the Pre and Post test sampling train leak check rates. For Method 5, the Post-test leak rate must be less than or equal to 0.02 acfm.
Note on B8: The Dry Gas Meter is a specialized flow meter, used to measure the volume of the extracted stack gas. THe Gas Meter measures a defined volume, regardless of the pressurized quantity or quality of the gas flowing through the meter. Temperature and pressure compensation must be made to measure the standard volume of gas moving through the meter.The ID is needed
Note on J8: Record the Pre and Post-test leak check results, as applicable.
Note on B9: Reagent Box, optional. A reagent box is a container that is sealable (for custody purposes) and is used to transport multiple recovered samples.
Note on F9: This is the dry gas meter correction coefficient, gamma, of an isokinetic sampling train meterbox (such as used for Method 5 sampling). The value should be between 0.95 and 1.05.
Note on J9: Nozzle inspections for dents, nicks, etc.
Note on B10: Sample gas transport line from the sample box to the meterbox, optional. The umbilical usually consists of bundled tubing, thermocouple, electrical lines, etc., used to control the probe and sample box filter temperatures.
Note on F10: This is the orifice coefficient of an isokinetic sampling train meterbox (such as used for Method 5 sampling).
Note on J10: Record the Pre and Post-test results of the thermocouple check, as applicable.
Note on B11: The ID of the thermocouple device used for monitoring the stack gas temperature. The ID is necessary for calibration documentation purposes
Note on F11: This is the pitot tube coefficient and is usually 0.84 for S-type pitots and 0.99 for standard pitots.
Note on B12: The ID of a Tedlar bag, if used. Optional.
Note on F12: This is the nozzle diameter, inches.
Note on B13: The ID of the pump used for filling a Tedlar bag, for example. Optional.
Note on K13: Select the units of measure for the impinger moisture readings.
Note on N13: Was a micro-manometer used for flow rate measurement?
Note on J14: Enter the initial and final moisture readings for impinger 1.
Note on N14: Enter sensitivity of micro-manometer if used.
Note on B15: The nozzle ID.
Note on F15: Input the barometric pressure of the sampling location. If the pressure is at sea level, remember to adjust the pressure for the elevation of the location above sea level. A 0.1 inches Hg decrease for every 100 feet of elevation is typically used.
Note on J15: Enter the initial and final moisture readings for impinger 2.
Note on J16: Enter the initial and final moisture readings for impinger 3.
Note on F17: Record the static pressure, inches of water, of the sampling location.
Note on J17: Enter the initial and final moisture readings for impinger 4.
Note on B18: For particulate sampling, the filters used have a unique identification. Record the number here.
Note on J18: Enter the initial and final moisture readings for impinger 5.
Note on B19: If two filters are used, record the second ID here.
Note on F19: Record the oxygen percent of the gas stream tested.
Note on J19: Enter the initial and final moisture readings for impinger 6.
Note on B20: If three filters are used, record the third filter ID here.
Note on F20: Record the carbon dioxide percent of the gas stream tested.
Note on J20: Enter the initial and final moisture readings for silica gel.
Note on N20: The volume or mass of the reagents increase in the impingers of an isokinetic sampling train is input here. This value is calculated from the impinger and silica gel weights entered above or can be entered directly.
Note on B22: The sampling point label, such as A1, A-1, D-2, etc.
Note on E22: This is the actual clock time at the start of sampling at a point. Necessary for non-isokinetic methods, see instructions.
Note on F22: This is the Dry Gas Meter volume reading at the beginning of the sampling at a point. This means that the final volume reading is recorded in a row without a Point label and no other recorded point data. Sometimes the sampling data is recorded at the end of sampling at a point which would require that the first volume reading is recorded without any other sampling data. Necessary for non-isokinetic methods, see instructions.
Note on G22: The velocity pressure (delta p) is expressed in inches of water.
Note on J22: Stack temperature is the temperature of the effluent gas at the sampling point and is expressed as degrees Fahrenheit.
Note on K22: Temperature of the sampling probe, degrees F. Necessary for non-isokinetic methods, see instructions.
Note on L22: Temperature of the filter box or compartment.
Note on N22: Enter the final impinger or condenser exit temperature. Necessary for non-isokinetic methods, see instructions.
Note on Q22: Vacuum of the sampling pump, inches mercury. Necessary for non-isokinetic methods, see instructions.
Note on R22: Sample flow rate.
Note on C23: The cumulative sampling time that sampling at the sample point was started, in minutes. Port changes DO NOT reset the time to zero (0).
Note on D23: The cumulative sampling time that sampling at the sample point was ended and is the begin time plus the sampling time per point.
Note on H23: Orifice pressure setting required for sampling isokinetically.
Note on I23: Orifice pressure actually sampled or reached, inches water
Note on L23: Filter inlet temperature or box temperature.
Note on M23: Filter outlet temperature only.
Note on O23: Dry gas meter inlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.
Note on P23: Dry gas meter outlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.

## Data Sheet 4
Plant: | [13 empty] | Optional Fields
Data required in ¶ 8.5 of EPA Method 5 and presented in Figure 5-3 of Method 5.
Sampling Location: | [5 empty] | Method: |  | Checks |  |  | Pre | Mid | Post
Run No.: |  |  |  | Testing Operator: |  |  |  | Vacuum (in Hg):
Run Date: |  |  |  | Job Number: |  |  |  | Leak Check Total Volume (ft3):
Equipment Identification | [7 empty] | Leak Rate (cfm):
Dry Gas Meter ID: |  |  |  | Calibration |  |  |  | Pitot (Yes/No):
Reagent Box ID: |  |  |  | Meter Box Correction Factor, Y: |  |  |  | Nozzle (Yes/No):
Unbilical Line ID: |  |  |  | Meter Box pressure differential, ΔH@ (in H2O): |  |  |  | Stack Thermocouple (Yes/No):
Stack Thermocouple ID: |  |  |  | Pitot tube coefficient, Cp:
Tedlar Bag IDs: |  |  |  | Nozzle diameter, Dn (in):
Orsat Pump ID: | [7 empty] | Moisture in: |  | Initial | Final | Micro-manometer Used (Y/N)?
Probe / Pitot ID: |  |  |  | Ambient Conditions |  |  |  | Impinger 1: |  |  |  | Sensitivity (in. H2O)
Nozzle ID: |  |  |  | Pbar (in. Hg): |  |  |  | Impinger 2:
|  |  |  | Temperature (°F): |  |  |  | Impinger 3:
Filters |  |  |  | Pg (in. H2O): |  |  |  | Impinger 4:
Filter No. 1 ID: |  |  |  | Concentrations |  |  |  | Impinger 5:
Filter No. 2 ID: |  |  |  | O2 Concentration (%): |  |  |  | Impinger 6:
Filter No. 3 ID: |  |  |  | CO2 Concentration (%): |  |  |  | Silica Gel. (gm): |  |  |  | Impinger Water Weight (g), Vlc: |  |  | 0
Traverse Point No. | Sampling Time (min) |  | Clock Time 24 hr | Dry Gas Meter Reading (Vm,ft3) | Velocity Head (ΔP in H2O) | Orfice Pressure Differential (ΔH, in H20) |  | Stack Temperature °F | Probe Temperature °F | Filter Temperature °F |  | Final exit Temperature °F | Dry Gas Meter Temperature °F |  | Pump Vacuum (in. Hg) | Sample Rate
| Begin | End |  |  |  | Desired | Actual |  |  | First | Second |  | Inlet | Outlet
Note on B2: Enter the name of the facility.
Note on B4: Enter the sampling location name, such as inlet, stack, ESP inlet, scrubber outlet, etc.
Note on H4: The method to be used to determine the analyte emissions is input here.
Note on B5: Enter the run number.
Note on F5: Enter the name of the operator performing the run.
Note on J5: Record the vacuum at which the Pre and Post sampling train leak checks were performed.
Note on B6: Enter the date that the run was performed.
Note on J6: Enter the value representing the total volume of all leak checks.
Note on B7: Enter an appropriate identification value for each piece of equipment.
Note on J7: Record the Pre and Post test sampling train leak check rates. For Method 5, the Post-test leak rate must be less than or equal to 0.02 acfm.
Note on B8: The Dry Gas Meter is a specialized flow meter, used to measure the volume of the extracted stack gas. THe Gas Meter measures a defined volume, regardless of the pressurized quantity or quality of the gas flowing through the meter. Temperature and pressure compensation must be made to measure the standard volume of gas moving through the meter.The ID is needed
Note on J8: Record the Pre and Post-test leak check results, as applicable.
Note on B9: Reagent Box, optional. A reagent box is a container that is sealable (for custody purposes) and is used to transport multiple recovered samples.
Note on F9: This is the dry gas meter correction coefficient, gamma, of an isokinetic sampling train meterbox (such as used for Method 5 sampling). The value should be between 0.95 and 1.05.
Note on J9: Nozzle inspections for dents, nicks, etc.
Note on B10: Sample gas transport line from the sample box to the meterbox, optional. The umbilical usually consists of bundled tubing, thermocouple, electrical lines, etc., used to control the probe and sample box filter temperatures.
Note on F10: This is the orifice coefficient of an isokinetic sampling train meterbox (such as used for Method 5 sampling).
Note on J10: Record the Pre and Post-test results of the thermocouple check, as applicable.
Note on B11: The ID of the thermocouple device used for monitoring the stack gas temperature. The ID is necessary for calibration documentation purposes
Note on F11: This is the pitot tube coefficient and is usually 0.84 for S-type pitots and 0.99 for standard pitots.
Note on B12: The ID of a Tedlar bag, if used. Optional.
Note on F12: This is the nozzle diameter, inches.
Note on B13: The ID of the pump used for filling a Tedlar bag, for example. Optional.
Note on K13: Select the units of measure for the impinger moisture readings.
Note on N13: Was a micro-manometer used for flow rate measurement?
Note on J14: Enter the initial and final moisture readings for impinger 1.
Note on N14: Enter sensitivity of micro-manometer if used.
Note on B15: The nozzle ID.
Note on F15: Input the barometric pressure of the sampling location. If the pressure is at sea level, remember to adjust the pressure for the elevation of the location above sea level. A 0.1 inches Hg decrease for every 100 feet of elevation is typically used.
Note on J15: Enter the initial and final moisture readings for impinger 2.
Note on J16: Enter the initial and final moisture readings for impinger 3.
Note on F17: Record the static pressure, inches of water, of the sampling location.
Note on J17: Enter the initial and final moisture readings for impinger 4.
Note on B18: For particulate sampling, the filters used have a unique identification. Record the number here.
Note on J18: Enter the initial and final moisture readings for impinger 5.
Note on B19: If two filters are used, record the second ID here.
Note on F19: Record the oxygen percent of the gas stream tested.
Note on J19: Enter the initial and final moisture readings for impinger 6.
Note on B20: If three filters are used, record the third filter ID here.
Note on F20: Record the carbon dioxide percent of the gas stream tested.
Note on J20: Enter the initial and final moisture readings for silica gel.
Note on N20: The volume or mass of the reagents increase in the impingers of an isokinetic sampling train is input here. This value is calculated from the impinger and silica gel weights entered above or can be entered directly.
Note on B22: The sampling point label, such as A1, A-1, D-2, etc.
Note on E22: This is the actual clock time at the start of sampling at a point. Necessary for non-isokinetic methods, see instructions.
Note on F22: This is the Dry Gas Meter volume reading at the beginning of the sampling at a point. This means that the final volume reading is recorded in a row without a Point label and no other recorded point data. Sometimes the sampling data is recorded at the end of sampling at a point which would require that the first volume reading is recorded without any other sampling data. Necessary for non-isokinetic methods, see instructions.
Note on G22: The velocity pressure (delta p) is expressed in inches of water.
Note on J22: Stack temperature is the temperature of the effluent gas at the sampling point and is expressed as degrees Fahrenheit.
Note on K22: Temperature of the sampling probe, degrees F. Necessary for non-isokinetic methods, see instructions.
Note on L22: Temperature of the filter box or compartment, degrees F.
Note on N22: Enter the final impinger or condenser exit temperature in degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.
Note on Q22: Vacuum of the sampling pump, inches mercury. Necessary for non-isokinetic methods, see instructions.
Note on R22: Sample flow rate.
Note on C23: The cumulative sampling time that sampling at the sample point was started, in minutes. Port changes DO NOT reset the time to zero (0).
Note on D23: The cumulative sampling time that sampling at the sample point was ended and is the begin time plus the sampling time per point.
Note on H23: Orifice pressure setting required for sampling isokinetically, inches water
Note on I23: Orifice pressure actually sampled or reached, inches water
Note on L23: Filter inlet temperature or box temperature.
Note on M23: Filter outlet temperature only.
Note on O23: Dry gas meter inlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.
Note on P23: Dry gas meter outlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.

## Data Sheet 5
Plant: | [13 empty] | Optional Fields
Data required in ¶ 8.5 of EPA Method 5 and presented in Figure 5-3 of Method 5.
Sampling Location: | [5 empty] | Method: |  | Checks |  |  | Pre | Mid | Post
Run No.: |  |  |  | Testing Operator: |  |  |  | Vacuum (in Hg):
Run Date: |  |  |  | Job Number: |  |  |  | Leak Check Total Volume (ft3):
Equipment Identification | [7 empty] | Leak Rate (cfm):
Dry Gas Meter ID: |  |  |  | Calibration |  |  |  | Pitot (Yes/No):
Reagent Box ID: |  |  |  | Meter Box Correction Factor, Y: |  |  |  | Nozzle (Yes/No):
Unbilical Line ID: |  |  |  | Meter Box pressure differential, ΔH@ (in H2O): |  |  |  | Stack Thermocouple (Yes/No):
Stack Thermocouple ID: |  |  |  | Pitot tube coefficient, Cp:
Tedlar Bag IDs: |  |  |  | Nozzle diameter, Dn (in):
Orsat Pump ID: | [7 empty] | Moisture in: |  | Initial | Final | Micro-manometer Used (Y/N)?
Probe / Pitot ID: |  |  |  | Ambient Conditions |  |  |  | Impinger 1: |  |  |  | Sensitivity (in. H2O)
Nozzle ID: |  |  |  | Pbar (in. Hg): |  |  |  | Impinger 2:
|  |  |  | Temperature (°F): |  |  |  | Impinger 3:
Filters |  |  |  | Pg (in. H2O): |  |  |  | Impinger 4:
Filter No. 1 ID: |  |  |  | Concentrations |  |  |  | Impinger 5:
Filter No. 2 ID: |  |  |  | O2 Concentration (%): |  |  |  | Impinger 6:
Filter No. 3 ID: |  |  |  | CO2 Concentration (%): |  |  |  | Silica Gel. (gm): |  |  |  | Impinger Water Weight (g), Vlc: |  |  | 0
Traverse Point No. | Sampling Time (min) |  | Clock Time 24 hr | Dry Gas Meter Reading (Vm,ft3) | Velocity Head (ΔP in H2O) | Orfice Pressure Differential (ΔH, in H20) |  | Stack Temperature °F | Probe Temperature °F | Filter Temperature °F |  | Final exit Temperature °F | Dry Gas Meter Temperature °F |  | Pump Vacuum (in. Hg) | Sample Rate
| Begin | End |  |  |  | Desired | Actual |  |  | First | Second |  | Inlet | Outlet
Note on B2: Enter the name of the facility.
Note on B4: Enter the sampling location name, such as inlet, stack, ESP inlet, scrubber outlet, etc.
Note on H4: The method to be used to determine the analyte emissions is input here.
Note on B5: Enter the run number.
Note on F5: Enter the name of the operator performing the run.
Note on J5: Record the vacuum at which the Pre and Post sampling train leak checks were performed.
Note on B6: Enter the date that the run was performed.
Note on J6: Enter the value representing the total volume of all leak checks.
Note on B7: Enter an appropriate identification value for each piece of equipment.
Note on J7: Record the Pre and Post test sampling train leak check rates. For Method 5, the Post-test leak rate must be less than or equal to 0.02 acfm.
Note on B8: The Dry Gas Meter is a specialized flow meter, used to measure the volume of the extracted stack gas. THe Gas Meter measures a defined volume, regardless of the pressurized quantity or quality of the gas flowing through the meter. Temperature and pressure compensation must be made to measure the standard volume of gas moving through the meter.The ID is needed
Note on J8: Record the Pre and Post-test leak check results, as applicable.
Note on B9: Reagent Box, optional. A reagent box is a container that is sealable (for custody purposes) and is used to transport multiple recovered samples.
Note on F9: This is the dry gas meter correction coefficient, gamma, of an isokinetic sampling train meterbox (such as used for Method 5 sampling). The value should be between 0.95 and 1.05.
Note on J9: Nozzle inspections for dents, nicks, etc.
Note on B10: Sample gas transport line from the sample box to the meterbox, optional. The umbilical usually consists of bundled tubing, thermocouple, electrical lines, etc., used to control the probe and sample box filter temperatures.
Note on F10: This is the orifice coefficient of an isokinetic sampling train meterbox (such as used for Method 5 sampling).
Note on J10: Record the Pre and Post-test results of the thermocouple check, as applicable.
Note on B11: The ID of the thermocouple device used for monitoring the stack gas temperature. The ID is necessary for calibration documentation purposes
Note on F11: This is the pitot tube coefficient and is usually 0.84 for S-type pitots and 0.99 for standard pitots.
Note on B12: The ID of a Tedlar bag, if used. Optional.
Note on F12: This is the nozzle diameter, inches.
Note on B13: The ID of the pump used for filling a Tedlar bag, for example. Optional.
Note on K13: Select the units of measure for the impinger moisture readings.
Note on N13: Was a micro-manometer used for flow rate measurement?
Note on J14: Enter the initial and final moisture readings for impinger 1.
Note on N14: Enter sensitivity of micro-manometer if used.
Note on B15: The nozzle ID.
Note on F15: Input the barometric pressure of the sampling location. If the pressure is at sea level, remember to adjust the pressure for the elevation of the location above sea level. A 0.1 inches Hg decrease for every 100 feet of elevation is typically used.
Note on J15: Enter the initial and final moisture readings for impinger 2.
Note on J16: Enter the initial and final moisture readings for impinger 3.
Note on F17: Record the static pressure, inches of water, of the sampling location.
Note on J17: Enter the initial and final moisture readings for impinger 4.
Note on B18: For particulate sampling, the filters used have a unique identification. Record the number here.
Note on J18: Enter the initial and final moisture readings for impinger 5.
Note on B19: If two filters are used, record the second ID here.
Note on F19: Record the oxygen percent of the gas stream tested.
Note on J19: Enter the initial and final moisture readings for impinger 6.
Note on B20: If three filters are used, record the third filter ID here.
Note on F20: Record the carbon dioxide percent of the gas stream tested.
Note on J20: Enter the initial and final moisture readings for silica gel.
Note on N20: The volume or mass of the reagents increase in the impingers of an isokinetic sampling train is input here. This value is calculated from the impinger and silica gel weights entered above or can be entered directly.
Note on B22: The sampling point label, such as A1, A-1, D-2, etc.
Note on E22: This is the actual clock time at the start of sampling at a point. Necessary for non-isokinetic methods, see instructions.
Note on F22: This is the Dry Gas Meter volume reading at the beginning of the sampling at a point. This means that the final volume reading is recorded in a row without a Point label and no other recorded point data. Sometimes the sampling data is recorded at the end of sampling at a point which would require that the first volume reading is recorded without any other sampling data. Necessary for non-isokinetic methods, see instructions.
Note on G22: The velocity pressure (delta p) is expressed in inches of water.
Note on J22: Stack temperature is the temperature of the effluent gas at the sampling point and is expressed as degrees Fahrenheit.
Note on K22: Temperature of the sampling probe, degrees F. Necessary for non-isokinetic methods, see instructions.
Note on L22: Temperature of the filter box or compartment.
Note on N22: Enter the final impinger or condenser exit temperature. Necessary for non-isokinetic methods, see instructions.
Note on Q22: Vacuum of the sampling pump, inches mercury. Necessary for non-isokinetic methods, see instructions.
Note on R22: Sample flow rate.
Note on C23: The cumulative sampling time that sampling at the sample point was started, in minutes. Port changes DO NOT reset the time to zero (0).
Note on D23: The cumulative sampling time that sampling at the sample point was ended and is the begin time plus the sampling time per point.
Note on H23: Orifice pressure setting required for sampling isokinetically.
Note on I23: Orifice pressure actually sampled or reached, inches water
Note on L23: Filter inlet temperature or box temperature.
Note on M23: Filter outlet temperature only.
Note on O23: Dry gas meter inlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.
Note on P23: Dry gas meter outlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.

## Data Sheet 6
Plant: | [13 empty] | Optional Fields
Data required in ¶ 8.5 of EPA Method 5 and presented in Figure 5-3 of Method 5.
Sampling Location: | [5 empty] | Method: |  | Checks |  |  | Pre | Mid | Post
Run No.: |  |  |  | Testing Operator: |  |  |  | Vacuum (in Hg):
Run Date: |  |  |  | Job Number: |  |  |  | Leak Check Total Volume (ft3):
Equipment Identification | [7 empty] | Leak Rate (cfm):
Dry Gas Meter ID: |  |  |  | Calibration |  |  |  | Pitot (Yes/No):
Reagent Box ID: |  |  |  | Meter Box Correction Factor, Y: |  |  |  | Nozzle (Yes/No):
Unbilical Line ID: |  |  |  | Meter Box pressure differential, ΔH@ (in H2O): |  |  |  | Stack Thermocouple (Yes/No):
Stack Thermocouple ID: |  |  |  | Pitot tube coefficient, Cp:
Tedlar Bag IDs: |  |  |  | Nozzle diameter, Dn (in):
Orsat Pump ID: | [7 empty] | Moisture in: |  | Initial | Final | Micro-manometer Used (Y/N)?
Probe / Pitot ID: |  |  |  | Ambient Conditions |  |  |  | Impinger 1: |  |  |  | Sensitivity (in. H2O)
Nozzle ID: |  |  |  | Pbar (in. Hg): |  |  |  | Impinger 2:
|  |  |  | Temperature (°F): |  |  |  | Impinger 3:
Filters |  |  |  | Pg (in. H2O): |  |  |  | Impinger 4:
Filter No. 1 ID: |  |  |  | Concentrations |  |  |  | Impinger 5:
Filter No. 2 ID: |  |  |  | O2 Concentration (%): |  |  |  | Impinger 6:
Filter No. 3 ID: |  |  |  | CO2 Concentration (%): |  |  |  | Silica Gel. (gm): |  |  |  | Impinger Water Weight (g), Vlc: |  |  | 0
Traverse Point No. | Sampling Time (min) |  | Clock Time 24 hr | Dry Gas Meter Reading (Vm,ft3) | Velocity Head (ΔP in H2O) | Orfice Pressure Differential (ΔH, in H20) |  | Stack Temperature °F | Probe Temperature °F | Filter Temperature °F |  | Final exit Temperature °F | Dry Gas Meter Temperature °F |  | Pump Vacuum (in. Hg) | Sample Rate
| Begin | End |  |  |  | Desired | Actual |  |  | First | Second |  | Inlet | Outlet
Note on B2: Enter the name of the facility.
Note on B4: Enter the sampling location name, such as inlet, stack, ESP inlet, scrubber outlet, etc.
Note on H4: The method to be used to determine the analyte emissions is input here.
Note on B5: Enter the run number.
Note on F5: Enter the name of the operator performing the run.
Note on J5: Record the vacuum at which the Pre and Post sampling train leak checks were performed.
Note on B6: Enter the date that the run was performed.
Note on J6: Enter the value representing the total volume of all leak checks.
Note on B7: Enter an appropriate identification value for each piece of equipment.
Note on J7: Record the Pre and Post test sampling train leak check rates. For Method 5, the Post-test leak rate must be less than or equal to 0.02 acfm.
Note on B8: The Dry Gas Meter is a specialized flow meter, used to measure the volume of the extracted stack gas. THe Gas Meter measures a defined volume, regardless of the pressurized quantity or quality of the gas flowing through the meter. Temperature and pressure compensation must be made to measure the standard volume of gas moving through the meter.The ID is needed
Note on J8: Record the Pre and Post-test leak check results, as applicable.
Note on B9: Reagent Box, optional. A reagent box is a container that is sealable (for custody purposes) and is used to transport multiple recovered samples.
Note on F9: This is the dry gas meter correction coefficient, gamma, of an isokinetic sampling train meterbox (such as used for Method 5 sampling). The value should be between 0.95 and 1.05.
Note on J9: Nozzle inspections for dents, nicks, etc.
Note on B10: Sample gas transport line from the sample box to the meterbox, optional. The umbilical usually consists of bundled tubing, thermocouple, electrical lines, etc., used to control the probe and sample box filter temperatures.
Note on F10: This is the orifice coefficient of an isokinetic sampling train meterbox (such as used for Method 5 sampling).
Note on J10: Record the Pre and Post-test results of the thermocouple check, as applicable.
Note on B11: The ID of the thermocouple device used for monitoring the stack gas temperature. The ID is necessary for calibration documentation purposes
Note on F11: This is the pitot tube coefficient and is usually 0.84 for S-type pitots and 0.99 for standard pitots.
Note on B12: The ID of a Tedlar bag, if used. Optional.
Note on F12: This is the nozzle diameter, inches.
Note on B13: The ID of the pump used for filling a Tedlar bag, for example. Optional.
Note on K13: Select the units of measure for the impinger moisture readings.
Note on N13: Was a micro-manometer used for flow rate measurement?
Note on J14: Enter the initial and final moisture readings for impinger 1.
Note on N14: Enter sensitivity of micro-manometer if used.
Note on B15: The nozzle ID.
Note on F15: Input the barometric pressure of the sampling location. If the pressure is at sea level, remember to adjust the pressure for the elevation of the location above sea level. A 0.1 inches Hg decrease for every 100 feet of elevation is typically used.
Note on J15: Enter the initial and final moisture readings for impinger 2.
Note on J16: Enter the initial and final moisture readings for impinger 3.
Note on F17: Record the static pressure, inches of water, of the sampling location.
Note on J17: Enter the initial and final moisture readings for impinger 4.
Note on B18: For particulate sampling, the filters used have a unique identification. Record the number here.
Note on J18: Enter the initial and final moisture readings for impinger 5.
Note on B19: If two filters are used, record the second ID here.
Note on F19: Record the oxygen percent of the gas stream tested.
Note on J19: Enter the initial and final moisture readings for impinger 6.
Note on B20: If three filters are used, record the third filter ID here.
Note on F20: Record the carbon dioxide percent of the gas stream tested.
Note on J20: Enter the initial and final moisture readings for silica gel.
Note on N20: The volume or mass of the reagents increase in the impingers of an isokinetic sampling train is input here. This value is calculated from the impinger and silica gel weights entered above or can be entered directly.
Note on B22: The sampling point label, such as A1, A-1, D-2, etc.
Note on E22: This is the actual clock time at the start of sampling at a point. Necessary for non-isokinetic methods, see instructions.
Note on F22: This is the Dry Gas Meter volume reading at the beginning of the sampling at a point. This means that the final volume reading is recorded in a row without a Point label and no other recorded point data. Sometimes the sampling data is recorded at the end of sampling at a point which would require that the first volume reading is recorded without any other sampling data. Necessary for non-isokinetic methods, see instructions.
Note on G22: The velocity pressure (delta p) is expressed in inches of water.
Note on J22: Stack temperature is the temperature of the effluent gas at the sampling point and is expressed as degrees Fahrenheit.
Note on K22: Temperature of the sampling probe, degrees F. Necessary for non-isokinetic methods, see instructions.
Note on L22: Temperature of the filter box or compartment.
Note on N22: Enter the final impinger or condenser exit temperature. Necessary for non-isokinetic methods, see instructions.
Note on Q22: Vacuum of the sampling pump, inches mercury. Necessary for non-isokinetic methods, see instructions.
Note on R22: Sample flow rate.
Note on C23: The cumulative sampling time that sampling at the sample point was started, in minutes. Port changes DO NOT reset the time to zero (0).
Note on D23: The cumulative sampling time that sampling at the sample point was ended and is the begin time plus the sampling time per point.
Note on H23: Orifice pressure setting required for sampling isokinetically.
Note on I23: Orifice pressure actually sampled or reached, inches water
Note on L23: Filter inlet temperature or box temperature.
Note on M23: Filter outlet temperature only.
Note on O23: Dry gas meter inlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.
Note on P23: Dry gas meter outlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.

## Data Sheet 7
Plant: | [13 empty] | Optional Fields
Data required in ¶ 8.5 of EPA Method 5 and presented in Figure 5-3 of Method 5.
Sampling Location: | [5 empty] | Method: |  | Checks |  |  | Pre | Mid | Post
Run No.: |  |  |  | Testing Operator: |  |  |  | Vacuum (in Hg):
Run Date: |  |  |  | Job Number: |  |  |  | Leak Check Total Volume (ft3):
Equipment Identification | [7 empty] | Leak Rate (cfm):
Dry Gas Meter ID: |  |  |  | Calibration |  |  |  | Pitot (Yes/No):
Reagent Box ID: |  |  |  | Meter Box Correction Factor, Y: |  |  |  | Nozzle (Yes/No):
Unbilical Line ID: |  |  |  | Meter Box pressure differential, ΔH@ (in H2O): |  |  |  | Stack Thermocouple (Yes/No):
Stack Thermocouple ID: |  |  |  | Pitot tube coefficient, Cp:
Tedlar Bag IDs: |  |  |  | Nozzle diameter, Dn (in):
Orsat Pump ID: | [7 empty] | Moisture in: |  | Initial | Final | Micro-manometer Used (Y/N)?
Probe / Pitot ID: |  |  |  | Ambient Conditions |  |  |  | Impinger 1: |  |  |  | Sensitivity (in. H2O)
Nozzle ID: |  |  |  | Pbar (in. Hg): |  |  |  | Impinger 2:
|  |  |  | Temperature (°F): |  |  |  | Impinger 3:
Filters |  |  |  | Pg (in. H2O): |  |  |  | Impinger 4:
Filter No. 1 ID: |  |  |  | Concentrations |  |  |  | Impinger 5:
Filter No. 2 ID: |  |  |  | O2 Concentration (%): |  |  |  | Impinger 6:
Filter No. 3 ID: |  |  |  | CO2 Concentration (%): |  |  |  | Silica Gel. (gm): |  |  |  | Impinger Water Weight (g), Vlc: |  |  | 0
Traverse Point No. | Sampling Time (min) |  | Clock Time 24 hr | Dry Gas Meter Reading (Vm,ft3) | Velocity Head (ΔP in H2O) | Orfice Pressure Differential (ΔH, in H20) |  | Stack Temperature °F | Probe Temperature °F | Filter Temperature °F |  | Final exit Temperature °F | Dry Gas Meter Temperature °F |  | Pump Vacuum (in. Hg) | Sample Rate
| Begin | End |  |  |  | Desired | Actual |  |  | First | Second |  | Inlet | Outlet
Note on B2: Enter the name of the facility.
Note on B4: Enter the sampling location name, such as inlet, stack, ESP inlet, scrubber outlet, etc.
Note on H4: The method to be used to determine the analyte emissions is input here.
Note on B5: Enter the run number.
Note on F5: Enter the name of the operator performing the run.
Note on J5: Record the vacuum at which the Pre and Post sampling train leak checks were performed.
Note on B6: Enter the date that the run was performed.
Note on J6: Enter the value representing the total volume of all leak checks.
Note on B7: Enter an appropriate identification value for each piece of equipment.
Note on J7: Record the Pre and Post test sampling train leak check rates. For Method 5, the Post-test leak rate must be less than or equal to 0.02 acfm.
Note on B8: The Dry Gas Meter is a specialized flow meter, used to measure the volume of the extracted stack gas. THe Gas Meter measures a defined volume, regardless of the pressurized quantity or quality of the gas flowing through the meter. Temperature and pressure compensation must be made to measure the standard volume of gas moving through the meter.The ID is needed
Note on J8: Record the Pre and Post-test leak check results, as applicable.
Note on B9: Reagent Box, optional. A reagent box is a container that is sealable (for custody purposes) and is used to transport multiple recovered samples.
Note on F9: This is the dry gas meter correction coefficient, gamma, of an isokinetic sampling train meterbox (such as used for Method 5 sampling). The value should be between 0.95 and 1.05.
Note on J9: Nozzle inspections for dents, nicks, etc.
Note on B10: Sample gas transport line from the sample box to the meterbox, optional. The umbilical usually consists of bundled tubing, thermocouple, electrical lines, etc., used to control the probe and sample box filter temperatures.
Note on F10: This is the orifice coefficient of an isokinetic sampling train meterbox (such as used for Method 5 sampling).
Note on J10: Record the Pre and Post-test results of the thermocouple check, as applicable.
Note on B11: The ID of the thermocouple device used for monitoring the stack gas temperature. The ID is necessary for calibration documentation purposes
Note on F11: This is the pitot tube coefficient and is usually 0.84 for S-type pitots and 0.99 for standard pitots.
Note on B12: The ID of a Tedlar bag, if used. Optional.
Note on F12: This is the nozzle diameter, inches.
Note on B13: The ID of the pump used for filling a Tedlar bag, for example. Optional.
Note on K13: Select the units of measure for the impinger moisture readings.
Note on N13: Was a micro-manometer used for flow rate measurement?
Note on J14: Enter the initial and final moisture readings for impinger 1.
Note on N14: Enter sensitivity of micro-manometer if used.
Note on B15: The nozzle ID.
Note on F15: Input the barometric pressure of the sampling location. If the pressure is at sea level, remember to adjust the pressure for the elevation of the location above sea level. A 0.1 inches Hg decrease for every 100 feet of elevation is typically used.
Note on J15: Enter the initial and final moisture readings for impinger 2.
Note on J16: Enter the initial and final moisture readings for impinger 3.
Note on F17: Record the static pressure, inches of water, of the sampling location.
Note on J17: Enter the initial and final moisture readings for impinger 4.
Note on B18: For particulate sampling, the filters used have a unique identification. Record the number here.
Note on J18: Enter the initial and final moisture readings for impinger 5.
Note on B19: If two filters are used, record the second ID here.
Note on F19: Record the oxygen percent of the gas stream tested.
Note on J19: Enter the initial and final moisture readings for impinger 6.
Note on B20: If three filters are used, record the third filter ID here.
Note on F20: Record the carbon dioxide percent of the gas stream tested.
Note on J20: Enter the initial and final moisture readings for silica gel.
Note on N20: The volume or mass of the reagents increase in the impingers of an isokinetic sampling train is input here. This value is calculated from the impinger and silica gel weights entered above or can be entered directly.
Note on B22: The sampling point label, such as A1, A-1, D-2, etc.
Note on E22: This is the actual clock time at the start of sampling at a point. Necessary for non-isokinetic methods, see instructions.
Note on F22: This is the Dry Gas Meter volume reading at the beginning of the sampling at a point. This means that the final volume reading is recorded in a row without a Point label and no other recorded point data. Sometimes the sampling data is recorded at the end of sampling at a point which would require that the first volume reading is recorded without any other sampling data. Necessary for non-isokinetic methods, see instructions.
Note on G22: The velocity pressure (delta p) is expressed in inches of water.
Note on J22: Stack temperature is the temperature of the effluent gas at the sampling point and is expressed as degrees Fahrenheit.
Note on K22: Temperature of the sampling probe, degrees F. Necessary for non-isokinetic methods, see instructions.
Note on L22: Temperature of the filter box or compartment.
Note on N22: Enter the final impinger or condenser exit temperature. Necessary for non-isokinetic methods, see instructions.
Note on Q22: Vacuum of the sampling pump, inches mercury. Necessary for non-isokinetic methods, see instructions.
Note on R22: Sample flow rate.
Note on C23: The cumulative sampling time that sampling at the sample point was started, in minutes. Port changes DO NOT reset the time to zero (0).
Note on D23: The cumulative sampling time that sampling at the sample point was ended and is the begin time plus the sampling time per point.
Note on H23: Orifice pressure setting required for sampling isokinetically.
Note on I23: Orifice pressure actually sampled or reached, inches water
Note on L23: Filter inlet temperature or box temperature.
Note on M23: Filter outlet temperature only.
Note on O23: Dry gas meter inlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.
Note on P23: Dry gas meter outlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.

## Data Sheet 8
Plant: | [13 empty] | Optional Fields
Data required in ¶ 8.5 of EPA Method 5 and presented in Figure 5-3 of Method 5.
Sampling Location: | [5 empty] | Method: |  | Checks |  |  | Pre | Mid | Post
Run No.: |  |  |  | Testing Operator: |  |  |  | Vacuum (in Hg):
Run Date: |  |  |  | Job Number: |  |  |  | Leak Check Total Volume (ft3):
Equipment Identification | [7 empty] | Leak Rate (cfm):
Dry Gas Meter ID: |  |  |  | Calibration |  |  |  | Pitot (Yes/No):
Reagent Box ID: |  |  |  | Meter Box Correction Factor, Y: |  |  |  | Nozzle (Yes/No):
Unbilical Line ID: |  |  |  | Meter Box pressure differential, ΔH@ (in H2O): |  |  |  | Stack Thermocouple (Yes/No):
Stack Thermocouple ID: |  |  |  | Pitot tube coefficient, Cp:
Tedlar Bag IDs: |  |  |  | Nozzle diameter, Dn (in):
Orsat Pump ID: | [7 empty] | Moisture in: |  | Initial | Final | Micro-manometer Used (Y/N)?
Probe / Pitot ID: |  |  |  | Ambient Conditions |  |  |  | Impinger 1: |  |  |  | Sensitivity (in. H2O)
Nozzle ID: |  |  |  | Pbar (in. Hg): |  |  |  | Impinger 2:
|  |  |  | Temperature (°F): |  |  |  | Impinger 3:
Filters |  |  |  | Pg (in. H2O): |  |  |  | Impinger 4:
Filter No. 1 ID: |  |  |  | Concentrations |  |  |  | Impinger 5:
Filter No. 2 ID: |  |  |  | O2 Concentration (%): |  |  |  | Impinger 6:
Filter No. 3 ID: |  |  |  | CO2 Concentration (%): |  |  |  | Silica Gel. (gm): |  |  |  | Impinger Water Weight (g), Vlc: |  |  | 0
Traverse Point No. | Sampling Time (min) |  | Clock Time 24 hr | Dry Gas Meter Reading (Vm,ft3) | Velocity Head (ΔP in H2O) | Orfice Pressure Differential (ΔH, in H20) |  | Stack Temperature °F | Probe Temperature °F | Filter Temperature °F |  | Final exit Temperature °F | Dry Gas Meter Temperature °F |  | Pump Vacuum (in. Hg) | Sample Rate
| Begin | End |  |  |  | Desired | Actual |  |  | First | Second |  | Inlet | Outlet
Note on B2: Enter the name of the facility.
Note on B4: Enter the sampling location name, such as inlet, stack, ESP inlet, scrubber outlet, etc.
Note on H4: The method to be used to determine the analyte emissions is input here.
Note on B5: Enter the run number.
Note on F5: Enter the name of the operator performing the run.
Note on J5: Record the vacuum at which the Pre and Post sampling train leak checks were performed.
Note on B6: Enter the date that the run was performed.
Note on J6: Enter the value representing the total volume of all leak checks.
Note on B7: Enter an appropriate identification value for each piece of equipment.
Note on J7: Record the Pre and Post test sampling train leak check rates. For Method 5, the Post-test leak rate must be less than or equal to 0.02 acfm.
Note on B8: The Dry Gas Meter is a specialized flow meter, used to measure the volume of the extracted stack gas. THe Gas Meter measures a defined volume, regardless of the pressurized quantity or quality of the gas flowing through the meter. Temperature and pressure compensation must be made to measure the standard volume of gas moving through the meter.The ID is needed
Note on J8: Record the Pre and Post-test leak check results, as applicable.
Note on B9: Reagent Box, optional. A reagent box is a container that is sealable (for custody purposes) and is used to transport multiple recovered samples.
Note on F9: This is the dry gas meter correction coefficient, gamma, of an isokinetic sampling train meterbox (such as used for Method 5 sampling). The value should be between 0.95 and 1.05.
Note on J9: Nozzle inspections for dents, nicks, etc.
Note on B10: Sample gas transport line from the sample box to the meterbox, optional. The umbilical usually consists of bundled tubing, thermocouple, electrical lines, etc., used to control the probe and sample box filter temperatures.
Note on F10: This is the orifice coefficient of an isokinetic sampling train meterbox (such as used for Method 5 sampling).
Note on J10: Record the Pre and Post-test results of the thermocouple check, as applicable.
Note on B11: The ID of the thermocouple device used for monitoring the stack gas temperature. The ID is necessary for calibration documentation purposes
Note on F11: This is the pitot tube coefficient and is usually 0.84 for S-type pitots and 0.99 for standard pitots.
Note on B12: The ID of a Tedlar bag, if used. Optional.
Note on F12: This is the nozzle diameter, inches.
Note on B13: The ID of the pump used for filling a Tedlar bag, for example. Optional.
Note on K13: Select the units of measure for the impinger moisture readings.
Note on N13: Was a micro-manometer used for flow rate measurement?
Note on J14: Enter the initial and final moisture readings for impinger 1.
Note on N14: Enter sensitivity of micro-manometer if used.
Note on B15: The nozzle ID.
Note on F15: Input the barometric pressure of the sampling location. If the pressure is at sea level, remember to adjust the pressure for the elevation of the location above sea level. A 0.1 inches Hg decrease for every 100 feet of elevation is typically used.
Note on J15: Enter the initial and final moisture readings for impinger 2.
Note on J16: Enter the initial and final moisture readings for impinger 3.
Note on F17: Record the static pressure, inches of water, of the sampling location.
Note on J17: Enter the initial and final moisture readings for impinger 4.
Note on B18: For particulate sampling, the filters used have a unique identification. Record the number here.
Note on J18: Enter the initial and final moisture readings for impinger 5.
Note on B19: If two filters are used, record the second ID here.
Note on F19: Record the oxygen percent of the gas stream tested.
Note on J19: Enter the initial and final moisture readings for impinger 6.
Note on B20: If three filters are used, record the third filter ID here.
Note on F20: Record the carbon dioxide percent of the gas stream tested.
Note on J20: Enter the initial and final moisture readings for silica gel.
Note on N20: The volume or mass of the reagents increase in the impingers of an isokinetic sampling train is input here. This value is calculated from the impinger and silica gel weights entered above or can be entered directly.
Note on B22: The sampling point label, such as A1, A-1, D-2, etc.
Note on E22: This is the actual clock time at the start of sampling at a point. Necessary for non-isokinetic methods, see instructions.
Note on F22: This is the Dry Gas Meter volume reading at the beginning of the sampling at a point. This means that the final volume reading is recorded in a row without a Point label and no other recorded point data. Sometimes the sampling data is recorded at the end of sampling at a point which would require that the first volume reading is recorded without any other sampling data. Necessary for non-isokinetic methods, see instructions.
Note on G22: The velocity pressure (delta p) is expressed in inches of water.
Note on J22: Stack temperature is the temperature of the effluent gas at the sampling point and is expressed as degrees Fahrenheit.
Note on K22: Temperature of the sampling probe, degrees F. Necessary for non-isokinetic methods, see instructions.
Note on L22: Temperature of the filter box or compartment.
Note on N22: Enter the final impinger or condenser exit temperature. Necessary for non-isokinetic methods, see instructions.
Note on Q22: Vacuum of the sampling pump, inches mercury. Necessary for non-isokinetic methods, see instructions.
Note on R22: Sample flow rate.
Note on C23: The cumulative sampling time that sampling at the sample point was started, in minutes. Port changes DO NOT reset the time to zero (0).
Note on D23: The cumulative sampling time that sampling at the sample point was ended and is the begin time plus the sampling time per point.
Note on H23: Orifice pressure setting required for sampling isokinetically.
Note on I23: Orifice pressure actually sampled or reached, inches water
Note on L23: Filter inlet temperature or box temperature.
Note on M23: Filter outlet temperature only.
Note on O23: Dry gas meter inlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.
Note on P23: Dry gas meter outlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.

## Data Sheet 9
Plant: | [13 empty] | Optional Fields
Data required in ¶ 8.5 of EPA Method 5 and presented in Figure 5-3 of Method 5.
Sampling Location: | [5 empty] | Method: |  | Checks |  |  | Pre | Mid | Post
Run No.: |  |  |  | Testing Operator: |  |  |  | Vacuum (in Hg):
Run Date: |  |  |  | Job Number: |  |  |  | Leak Check Total Volume (ft3):
Equipment Identification | [7 empty] | Leak Rate (cfm):
Dry Gas Meter ID: |  |  |  | Calibration |  |  |  | Pitot (Yes/No):
Reagent Box ID: |  |  |  | Meter Box Correction Factor, Y: |  |  |  | Nozzle (Yes/No):
Unbilical Line ID: |  |  |  | Meter Box pressure differential, ΔH@ (in H2O): |  |  |  | Stack Thermocouple (Yes/No):
Stack Thermocouple ID: |  |  |  | Pitot tube coefficient, Cp:
Tedlar Bag IDs: |  |  |  | Nozzle diameter, Dn (in):
Orsat Pump ID: | [7 empty] | Moisture in: |  | Initial | Final | Micro-manometer Used (Y/N)?
Probe / Pitot ID: |  |  |  | Ambient Conditions |  |  |  | Impinger 1: |  |  |  | Sensitivity (in. H2O)
Nozzle ID: |  |  |  | Pbar (in. Hg): |  |  |  | Impinger 2:
|  |  |  | Temperature (°F): |  |  |  | Impinger 3:
Filters |  |  |  | Pg (in. H2O): |  |  |  | Impinger 4:
Filter No. 1 ID: |  |  |  | Concentrations |  |  |  | Impinger 5:
Filter No. 2 ID: |  |  |  | O2 Concentration (%): |  |  |  | Impinger 6:
Filter No. 3 ID: |  |  |  | CO2 Concentration (%): |  |  |  | Silica Gel. (gm): |  |  |  | Impinger Water Weight (g), Vlc: |  |  | 0
Traverse Point No. | Sampling Time (min) |  | Clock Time 24 hr | Dry Gas Meter Reading (Vm,ft3) | Velocity Head (ΔP in H2O) | Orfice Pressure Differential (ΔH, in H20) |  | Stack Temperature °F | Probe Temperature °F | Filter Temperature °F |  | Final exit Temperature °F | Dry Gas Meter Temperature °F |  | Pump Vacuum (in. Hg) | Sample Rate
| Begin | End |  |  |  | Desired | Actual |  |  | First | Second |  | Inlet | Outlet
Note on B2: Enter the name of the facility.
Note on B4: Enter the sampling location name, such as inlet, stack, ESP inlet, scrubber outlet, etc.
Note on H4: The method to be used to determine the analyte emissions is input here.
Note on B5: Enter the run number.
Note on F5: Enter the name of the operator performing the run.
Note on J5: Record the vacuum at which the Pre and Post sampling train leak checks were performed.
Note on B6: Enter the date that the run was performed.
Note on J6: Enter the value representing the total volume of all leak checks.
Note on B7: Enter an appropriate identification value for each piece of equipment.
Note on J7: Record the Pre and Post test sampling train leak check rates. For Method 5, the Post-test leak rate must be less than or equal to 0.02 acfm.
Note on B8: The Dry Gas Meter is a specialized flow meter, used to measure the volume of the extracted stack gas. THe Gas Meter measures a defined volume, regardless of the pressurized quantity or quality of the gas flowing through the meter. Temperature and pressure compensation must be made to measure the standard volume of gas moving through the meter.The ID is needed
Note on J8: Record the Pre and Post-test leak check results, as applicable.
Note on B9: Reagent Box, optional. A reagent box is a container that is sealable (for custody purposes) and is used to transport multiple recovered samples.
Note on F9: This is the dry gas meter correction coefficient, gamma, of an isokinetic sampling train meterbox (such as used for Method 5 sampling). The value should be between 0.95 and 1.05.
Note on J9: Nozzle inspections for dents, nicks, etc.
Note on B10: Sample gas transport line from the sample box to the meterbox, optional. The umbilical usually consists of bundled tubing, thermocouple, electrical lines, etc., used to control the probe and sample box filter temperatures.
Note on F10: This is the orifice coefficient of an isokinetic sampling train meterbox (such as used for Method 5 sampling).
Note on J10: Record the Pre and Post-test results of the thermocouple check, as applicable.
Note on B11: The ID of the thermocouple device used for monitoring the stack gas temperature. The ID is necessary for calibration documentation purposes
Note on F11: This is the pitot tube coefficient and is usually 0.84 for S-type pitots and 0.99 for standard pitots.
Note on B12: The ID of a Tedlar bag, if used. Optional.
Note on F12: This is the nozzle diameter, inches.
Note on B13: The ID of the pump used for filling a Tedlar bag, for example. Optional.
Note on K13: Select the units of measure for the impinger moisture readings.
Note on N13: Was a micro-manometer used for flow rate measurement?
Note on J14: Enter the initial and final moisture readings for impinger 1.
Note on N14: Enter sensitivity of micro-manometer if used.
Note on B15: The nozzle ID.
Note on F15: Input the barometric pressure of the sampling location. If the pressure is at sea level, remember to adjust the pressure for the elevation of the location above sea level. A 0.1 inches Hg decrease for every 100 feet of elevation is typically used.
Note on J15: Enter the initial and final moisture readings for impinger 2.
Note on J16: Enter the initial and final moisture readings for impinger 3.
Note on F17: Record the static pressure, inches of water, of the sampling location.
Note on J17: Enter the initial and final moisture readings for impinger 4.
Note on B18: For particulate sampling, the filters used have a unique identification. Record the number here.
Note on J18: Enter the initial and final moisture readings for impinger 5.
Note on B19: If two filters are used, record the second ID here.
Note on F19: Record the oxygen percent of the gas stream tested.
Note on J19: Enter the initial and final moisture readings for impinger 6.
Note on B20: If three filters are used, record the third filter ID here.
Note on F20: Record the carbon dioxide percent of the gas stream tested.
Note on J20: Enter the initial and final moisture readings for silica gel.
Note on N20: The volume or mass of the reagents increase in the impingers of an isokinetic sampling train is input here. This value is calculated from the impinger and silica gel weights entered above or can be entered directly.
Note on B22: The sampling point label, such as A1, A-1, D-2, etc.
Note on E22: This is the actual clock time at the start of sampling at a point. Necessary for non-isokinetic methods, see instructions.
Note on F22: This is the Dry Gas Meter volume reading at the beginning of the sampling at a point. This means that the final volume reading is recorded in a row without a Point label and no other recorded point data. Sometimes the sampling data is recorded at the end of sampling at a point which would require that the first volume reading is recorded without any other sampling data. Necessary for non-isokinetic methods, see instructions.
Note on G22: The velocity pressure (delta p) is expressed in inches of water.
Note on J22: Stack temperature is the temperature of the effluent gas at the sampling point and is expressed as degrees Fahrenheit.
Note on K22: Temperature of the sampling probe, degrees F. Necessary for non-isokinetic methods, see instructions.
Note on L22: Temperature of the filter box or compartment.
Note on N22: Enter the final impinger or condenser exit temperature. Necessary for non-isokinetic methods, see instructions.
Note on Q22: Vacuum of the sampling pump, inches mercury. Necessary for non-isokinetic methods, see instructions.
Note on R22: Sample flow rate.
Note on C23: The cumulative sampling time that sampling at the sample point was started, in minutes. Port changes DO NOT reset the time to zero (0).
Note on D23: The cumulative sampling time that sampling at the sample point was ended and is the begin time plus the sampling time per point.
Note on H23: Orifice pressure setting required for sampling isokinetically.
Note on I23: Orifice pressure actually sampled or reached, inches water
Note on L23: Filter inlet temperature or box temperature.
Note on M23: Filter outlet temperature only.
Note on O23: Dry gas meter inlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.
Note on P23: Dry gas meter outlet gas temperature, expressed as degrees Fahrenheit. Necessary for non-isokinetic methods, see instructions.

## HeaderExport
RunNumber | RunDate | JobNumber | Method | PbHDR | PgHDR | FilterNum1 | FilterNum2 | FilterNum3 | ReagBox | Umbical | StackTC | TedlarBag | OrasatPump | SampBox | MeterBox | TCReadOut | Pitot | Nozzle | VlcHDR | YHDR | DH@HDR | CpHDR | DnHDR | PercCO2HDR | PercO2HDR | VacCkPre | VacCkPost | InitDGM | FinalDGM | LeakRtPre | LeakRtPost | PitotCkPre | PitotCkMid | PitotCkPost | NozzleCkPre | NozzleCkMid | NozzleCkPost | StackTCCKPre | StackTCCKMid | StackTCCKPost | Personnel | Location | Impinit1 | Impinit2 | Impinit3 | Impinit4 | Impinit5 | Impinit6 | Silicainit | Impfin1 | Impfin2 | Impfin3 | Impfin4 | Impfin5 | Impfin6 | Silicafin | Manused | Mansense
1 | 41855 | BB019012.0000 | PM/Metals | 30.05 | -0.085 | QtzP1245-7 | 0 | 0 | ES246 | 1 | TypeK-456-9 | 0 | 0 | ES246 | ES245 |  | 128 | SetES2987-.5 | 237.7 | 0.99 | 1.63 | 0.83 | 0.502 | 7.4915 | 13.2167 | 15 | 7 | 0 | 0 | 0.001 | 0.001 | Yes |  | Yes | Yes |  | Yes | Yes |  | Yes | TWB |  | 541.3 | 767.8 | 797.9 | 650.8 | 771.1 | 779.9 | 913.3 | 581.2 | 882.4 | 843.7 | 660.4 | 776.3 | 0 | 933.9 | No | 0
2 | 41855 | BB019012.0000 | PM/Metals | 30.02 | -0.085 | QtzP1245-8 | 0 | 0 | ES246 | 1 | TypeK-456-9 | 0 | 0 | ES246 | ES245 |  | 128 | SetES2987-.5 | 261.2 | 0.99 | 1.63 | 0.83 | 0.502 | 7.5031 | 12.2625 | 15 | 7 | 0 | 0 | 0.001 | 0.001 | Yes |  | Yes | Yes |  | Yes | Yes |  | Yes | TWB |  | 541.4 | 688.8 | 804.5 | 651.4 | 728.5 | 769.4 | 909 | 593.1 | 820.1 | 847.9 | 659.8 | 734.2 | 0 | 927.5 | No | 0
3 | 41855 | BB019012.0000 | PM/Metals | 30.02 | -0.08 | QtzP1245-7 | 0 | 0 | ES246 | 1 | TypeK-456-9 | 0 | 0 | ES246 | ES245 |  | 128 | SetES2987-.5 | 257.9 | 0.99 | 1.68 | 0.83 | 0.502 | 7.3964 | 12.689 | 15 | 7 | 0 | 0 | 0.001 | 0.001 | Yes |  | Yes | Yes |  | Yes | Yes |  | Yes | TWB |  | 542.6 | 770.9 | 717.8 | 641 | 770.4 | 772.7 | 924.5 | 607.1 | 896.3 | 756.1 | 648 | 774.4 | 0 | 942.9 | No | 0
0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 0 |  | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 0 |  | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 0 |  | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 0 |  | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 0 |  | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 0 |  | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0